Control checklist for biosafety level 2 laboratory (BSL 2)

(Adapted from annex 2 of the Directive concerning the internal procedure for projects subject to notification/authorization according to the ContainO, RO and OPTM Ordinances »

| BSL 2 laboratory and OHS-PR Biosafety control |  |
| --- | --- |
| OHS-PR Biosafety person |  |
| Date of the audit |  |
| BSL 2 room number |  |
| Name of the Unit |  |
| BSO |  |

| WORK ORGANISATION |  |
| --- | --- |
| BSL 2 Training courses ☐ yes ☐ no ☐ partially |  |
| Biosafety officer (BSO) training ☐ Continuing education in biosafety done recently ☐ No recent biosafety training |  |
| Other visitors: ☐ yes ☐ no ☐ partially |  |
| Written work instructions provided to users either electronically or on paper or displayed in the laboratory. Which format? ☐ yes ☐ no ☐ partially |  |
| Organism inventory Which format ☐ yes ☐ no ☐ partially |  |
| Organisms – information for every user and access to the organism sheets (if there are) ☐ yes ☐ no ☐ partially |  |
| Management of organism’s storage How? ☐ storage in the lab ☐ storage out of the lab : - Storage plan, - Biohazard sign, - User list, - « keep-out » sign |  |
| Unit-specific biosafety concept ☐ yes ☐ no ☐ partially |  |

| INFRASTRUCTURES |  |
| --- | --- |
| Access control via CAMIPRO ☐ yes ☐ no N° Demande de travaux : |  |
| « Biohazard » sign: ☐ yes ☐ no |  |
| « Keep-out » sign : ☐ yes ☐ no |  |
| Door safety sheet : - Up to date - Representative examples of organisms mentioned ☐ yes ☐ no |  |
| Air pressure on premises: Not authorised to be in positive pressure in relation to adjoining corridor ☐ Positive pressure ☐ Atmospheric pressure ☐ Negative pressure |  |
| Sealed windows : NB: ContainO and ATCC regulation. If windows cannot be sealed, they must be closed while work is being carried out in P2 zone ☐ yes ☐ no |  |
| Autoclave in the building with maintenance procedure ☐ yes ☐ no |  |
| Outside the SV faculty: Autoclave in the building with filter to avoid potentially contaminated air being discharged into the room / water. (2 autoclaves in CH: LEAGO et EML + 1 in GR Tamar Kohn (LEV) + 2 in MXD PBL et LP) ☐ yes ☐ no |  |
| Outside the SV faculty: Autoclave with ISO 17665 tests performed each 5 years ☐ yes ☐ no |  |
| Floor that can be easily cleaned with a disinfectant ☐ yes ☐ no |  |
| Bench resistant to chemicals: Water, acids, bases solvents, disinfectants, decontaminants ☐ yes ☐ no |  |
| Order and cleanliness ☐ yes ☐ no |  |
| Floor drain present in the BSL 2 laboratory ☐ yes ☐ no |  |
| Items that can absorb and retain aerosols present in the laboratory (such as a CD player, etc.) ☐ yes ☐ no |  |

| EQUIPMENT & SAFETY EQUIPMENT |  |
| --- | --- |
| Fix phone in the room ☐ yes ☐ no N° Demande d’intervention : |  |
| Biosafety cabinet - with maintenance procedure (BSC Class II for human pathogen manipulation) ☐ BSC class I ☐ BSC class II ☐ BSC class III ☐ None Check completed on: |  |
| Aspiration system (for biosafety cabinets) such as the VACUSAFE Filters are replaced regularly and the type of hydrophobic filter used is appropriate for the application (0.2 µm filters when handling viruses or viral vectors) ☐ yes ☐ no |  |
| Liquid waste treatment How? ☐ yes ☐ no |  |
| Rotors and « buckets » with safety lids All buckets or rotors for centrifugation must be used with their safety lids. If the tubes were closed during centrifugation and no problems were experienced during the process, it is not necessary to transport the buckets and rotors under a hood to open them. ☐ yes ☐ no |  |
| Pipets and pipet-tips with aerosol protection ☐ yes ☐ no |  |
| Solid waste management How? ☐ yes ☐ no ☐ partially |  |
| Waste transport to the autoclave room ☐ adapted ☐ non-adapted |  |
| Eye-wash station ☐ yes ☐ no |  |
| Decontamination and cleaning equipment Which? ☐ yes ☐ no WARNING: Biocidal is not effective against non-enveloped viruses |  |
| Cleaning management: - organized (what, when, by who) - tracked (log sheet, logbook, or other) ☐ yes ☐ no ☐ partially |  |
| Transport of samples between labs in the building or on campus ☐ adapted ☐ non-adapted |  |

| PERSONAL PROTECTIVE EQUIPMENT |  |
| --- | --- |
| Cotton lab coat or surgical gown with long sleeves and tight cuffs and procedure to clean and autoclave them ☐ yes ☐ no ☐ partially |  |
| Respiratory protection NB: For activities during which use of a BSC is not possible, it may be requested that a mask is worn, e.g. stereotaxic injections. ☐ yes ☐ no Reason: |  |
| Safety glasses or face mask ☐ yes ☐ no Raison: |  |
| Gloves (EN 374/2 et EN 374/3) ☐ yes ☐ no |  |
| Hand disinfectant: Which? ☐ yes ☐ no |  |
| Cover shoes or sticky mat at the entrance ☐ yes ☐ no |  |