## ControlPanel
Scanning Mobility CCN Analysis Post-Processor Spreadsheet | [6 empty] | ver. 7.3 |  | Contact: Athanasios Nenes (athanasios.nenes@gatech.edu)
DataSource |  |  | Data |  |  |  | Fitting the Sigmoid |  |  | Data ranges | Sigmoid | Asymptote
Counts2ConcConv for CPC |  |  | 0.2 |  |  |  | B | #REF! |  | minDp | 20 | 45
CCN data shift |  |  | 8 |  |  |  | D | 35.5802 |  | max Dp | 120 | 120
ScanColumn |  |  | O | P | N |  | c | -19.1965
Min CCN counts threshold |  |  | 4 |  |  |  | Dp50 | 35.5802 | nm | sum of error^2 |  | #REF! | Solution found, optimality and constraints satisfied
SMPS timeseries duration |  |  | 135 |  |  |  | Dp50(wet) | #N/A | um
Additional time in CCN time series |  |  | 60 |  |  |  | Dp50+20(wet) | #N/A | um
Timeseries in DataSource Sheet: |  |  | 109 | 246 | 443
Copy + paste the following into "saved data" sheet: |  |  |  |  | SS | Total CCN | Total CN |  | Max. T | Dp50 | Dp50 | Dp50+20 | B | D | c
DataSource | Meas Date | Start Time | End Time | Shift | (%) | Conc | Conc | CCN/CN | Error (%) | (dry) | (wet) | (wet) | (sigmd) | (sigmd) | (sigmd)
Data!O | #REF! | #REF! | #REF! | 8 | #REF! | #REF! | #REF! | #REF! | #REF! | 35.5802 | #N/A | #N/A | #REF! | 35.5802 | -19.1965
Figure A: SMPS and CCN concentrations over the scan time (used for alignment) | [7 empty] | Figure B: CCN/CN vs. Dp (used for charge correction, function fitting)
Graph Inset Bounds: xmin = |  | 115 | ymin = | 500 |  |  | 1
| xmax = | 130 | ymax = | 1000
Figure C: Droplet Sizes as a function of dry particle size | [5 empty] | Figure D: % Deviation of DT from setpoint
+ 1 s shift
- 1 s shift
+ 1 column
- 1 column
Perform Charge Correction
Fit Sigmoid
Update Inset +1 s shift
Initialize Charge Correction
Store In 'SavedData'
Log Scale +1 s shift

## Data
| Date | 2010-02-03 | 2010-02-03 | 2010-02-03 | 2010-02-03 | 2010-02-03 | 2010-02-03 | 2010-02-03 | 2010-02-03 | 2010-02-03 | 2010-02-03 | 2010-02-03 | 2010-02-03 | 2010-02-03 | 2010-02-03 | 2010-02-03 | 2010-02-03
| Start time | 9:46 | 9:49 | 9:51 | 9:53 | 9:55 | 9:58 | 10:00 | 10:02 | 10:04 | 10:09 | 10:11 | 10:13 | 10:16 | 10:20 | 10:22 | 10:25
| End time | 9:49 | 9:51 | 9:53 | 9:55 | 9:58 | 10:00 | 10:02 | 10:04 | 10:07 | 10:11 | 10:13 | 10:16 | 10:18 | 10:22 | 10:25 | N/A
| Nominal SSat | 0.2 | 0.2 | 0.2 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | 0.8 | 0.8 | 0.8
| Nominal Temp Grad | 5.1 | 5.1 | 5.1 | 8.33 | 8.33 | 8.33 | 8.33 | 8.33 | 11.57 | 11.57 | 11.57 | 11.57 | 11.57 | 14.8 | 14.8 | 14.8
| Max % Temp Err | 26.275 | 26.274 | 26.275 | 60.864 | 24.37 | 24.61 | 25.09 | 27.731 | 49.697 | 23.941 | 24.114 | 24.114 | 23.941 | 23.243 | 23.514 | 23.514
| Dp | dNdlogDp_for_each_Scan
| 3.46 | 4144080 | 4131530 | 4245060 | 4126730 | 4309050 | 4620450 | 4338140 | 3911500 | 4510070 | 3932810 | 4015920 | 3906380 | 4240300 | 4310490 | 4244200 | 4494030
| 3.59 | 1914190 | 1990300 | 1834730 | 1940670 | 1903250 | 2037730 | 1888330 | 2081340 | 2026950 | 1728310 | 1824750 | 1759080 | 1840750 | 1821730 | 1963510 | 1987610
| 3.72 | 1039960 | 970654 | 922759 | 1013330 | 1018080 | 1245080 | 1094440 | 1103640 | 1328790 | 856014 | 1130120 | 856330 | 1060370 | 1076480 | 1142970 | 1010980
| 3.85 | 624390 | 640272 | 579507 | 608583 | 726022 | 696855 | 665345 | 734853 | 821357 | 656334 | 525256 | 577802 | 481094 | 635348 | 741646 | 636874
| 4 | 365097 | 407009 | 469879 | 468010 | 453207 | 405611 | 473864 | 479074 | 499731 | 418599 | 398261 | 369599 | 420258 | 410943 | 454510 | 469668
| 4.14 | 306578 | 277006 | 271094 | 240343 | 312289 | 288570 | 319565 | 369967 | 362376 | 276899 | 266912 | 304197 | 273195 | 310252 | 310831 | 351404
| 4.29 | 213512 | 224565 | 212549 | 202739 | 216639 | 266975 | 257324 | 256389 | 271194 | 219489 | 193308 | 207791 | 179839 | 224131 | 231122 | 252275
| 4.45 | 142658 | 146678 | 150776 | 168386 | 158404 | 213870 | 166215 | 196768 | 217428 | 174468 | 140539 | 117699 | 146231 | 151482 | 160004 | 214358
| 4.61 | 118117 | 110004 | 106861 | 143134 | 139855 | 134796 | 131197 | 150702 | 139306 | 139355 | 109368 | 118391 | 95654.2031 | 135434 | 145961 | 131925
| 4.78 | 75906.2969 | 103647 | 103077 | 99656.6016 | 97786.7969 | 104098 | 102351 | 121316 | 99400.6016 | 83744.2031 | 85926 | 81297.5 | 98550.7969 | 97752.8984 | 98167.5 | 130316
| 4.96 | 65855.5 | 72233.6016 | 85752.2969 | 69684 | 85537.7969 | 103859 | 90685.6016 | 97954.7969 | 76311.3984 | 75133.7031 | 60863.1016 | 67622.3984 | 65287.3008 | 83095 | 90984.2031 | 70491
| 5.14 | 54894.1016 | 62563.5 | 66322.2031 | 54864.6016 | 51885.6992 | 69121.2969 | 62337.3008 | 61709.5 | 71823 | 53627.8008 | 60348.3008 | 56857.8984 | 51617.3984 | 59978.1016 | 71311.3984 | 71922.2031
| 5.33 | 45981.8984 | 48321.5 | 48153.8984 | 69108.1016 | 52138.3008 | 65378.5 | 53099.1016 | 54321.5 | 52254.5 | 44853.3984 | 47047.8008 | 47546.1016 | 50332.6992 | 48153.8984 | 46614.3008 | 51027.6992
| 5.52 | 37614.1016 | 51212.8008 | 38753.1992 | 48262.1016 | 46044.6992 | 46670.1992 | 41712.8008 | 42013.3008 | 44452.6992 | 29024.5996 | 36629.8008 | 35347.3008 | 30780.1992 | 46190.5 | 41493.1016 | 40271.3008
| 5.73 | 40672.8984 | 35976.3008 | 34647.8984 | 41070.1016 | 39205.1016 | 42398.3984 | 45933.3984 | 42200.1016 | 35818.8008 | 29117.5996 | 35225.8008 | 29843.8008 | 37817.3984 | 38766.6992 | 33995.6016 | 34585.6992
| 5.94 | 30294.3008 | 30444.8008 | 39999 | 32847.1016 | 34090.5 | 28856.6992 | 34090.3008 | 32007 | 35748.1016 | 24011.8008 | 28042.5996 | 27932.5 | 32212.5996 | 31328.6992 | 30609.9004 | 37486.6016
| 6.15 | 25595.3008 | 19787.6992 | 19027.1992 | 19274.5996 | 29884.4004 | 17918.5 | 23991.8008 | 34737.3984 | 28051 | 27587.8008 | 15120.4004 | 21668.5996 | 22549.9004 | 30860.6992 | 33620.3984 | 24210.0996
| 6.38 | 23410 | 19632 | 20000.3008 | 22757.3008 | 24479.4004 | 24344.8008 | 19397.9004 | 26952.9004 | 27280.8008 | 20190.0996 | 18422.6992 | 19114.9004 | 19910.5996 | 25364.9004 | 19807 | 28535.5
| 6.61 | 22534.6992 | 20852.9004 | 22424.3008 | 23358.8008 | 18397 | 22474.1992 | 19566.1992 | 23896.6992 | 25365.9004 | 21781.3008 | 15789.9004 | 18476.8008 | 19883.9004 | 18664.8008 | 19174.3008 | 16474.3008
| 6.85 | 16945.1992 | 16029.5 | 18050.5996 | 19809.9004 | 18345.8008 | 22420.1992 | 17919.0996 | 21584.1992 | 20219.3008 | 20007 | 13862 | 18190.6992 | 19458.6992 | 20823.5996 | 19458.6992 | 18861.8008
| 7.1 | 18275.0996 | 18541.3008 | 16476.8008 | 14001.0996 | 16921.0996 | 16973.0996 | 18728 | 18646.8008 | 17707.9004 | 14299.7998 | 14863.7002 | 12010.7998 | 12062.7002 | 16260.9004 | 18295.5996 | 17092.0996
| 7.37 | 15295.7998 | 16538.1992 | 17325.5 | 15339.0996 | 14823.9004 | 16367.2002 | 13689.9004 | 16218 | 17132.8008 | 16372.0996 | 17195.5 | 13131.4004 | 11975.7002 | 19201 | 15746.0996 | 21026.0996
| 7.64 | 14179.2998 | 14579.7002 | 14164.7002 | 16219.0996 | 17026.4004 | 19545.8008 | 17063 | 16172 | 15666.7998 | 17437.4004 | 14274 | 14891.9004 | 16444.5 | 14252.4004 | 15590.0996 | 14350.2002
| 7.91 | 15041.7002 | 15002.5996 | 15348.0996 | 17540.0996 | 17568.5996 | 18481.1992 | 16560 | 14104.7998 | 16687.4004 | 13096.0996 | 12769 | 14818.2998 | 14495.2998 | 16781.5996 | 15701.2998 | 16332.4004
| 8.2 | 14983.5996 | 14057.5 | 12906 | 16148.7998 | 13253.7998 | 17148.0996 | 15243.9004 | 15293.2002 | 14932.5996 | 16875.0996 | 20272.8008 | 15423.2998 | 14455.5 | 17705.1992 | 17028.1992 | 18753
| 8.51 | 17217.5 | 17212.6992 | 19705 | 16771.5 | 15049.9004 | 18277.6992 | 17696.5 | 16225.0996 | 18699.6992 | 15694.2002 | 18030.4004 | 17072.4004 | 16521.0996 | 19248.1992 | 17489.9004 | 20347.6992
| 8.82 | 19237.5 | 15564.7998 | 21047.5 | 19957.9004 | 21835.6992 | 20489.6992 | 17672.6992 | 18793.9004 | 19297.1992 | 14207.7998 | 16631.9004 | 17985.0996 | 16308 | 19924.5 | 16019.5996 | 22436.5
| 9.14 | 23857.4004 | 18396.1992 | 22285.5 | 22742 | 21786.6992 | 22122 | 17828.4004 | 20723.3008 | 19610.6992 | 17909.5 | 22498.8008 | 19058.1992 | 18319.0996 | 20799 | 22586.5996 | 20867.9004
| 9.47 | 22448.1992 | 23525.8008 | 22000.6992 | 23360.1992 | 20682 | 21914.5 | 24124.6992 | 21625.3008 | 22413.3008 | 22437.8008 | 23683.5 | 28955.3008 | 22778.8008 | 23828.5996 | 22069.4004 | 25216
| 9.82 | 26448.1992 | 28163.4004 | 22354.6992 | 25447.5 | 23546.1992 | 25871.6992 | 22313.9004 | 28016.0996 | 27053.9004 | 25835.5 | 27789.9004 | 27437.4004 | 24619.8008 | 26583.4004 | 28373 | 24451.1992
| 10.2 | 30969.0996 | 33184.3008 | 28704.0996 | 30217.0996 | 26902.6992 | 32516.5996 | 29774.9004 | 31129.0996 | 30427.4004 | 27641.0996 | 26068.6992 | 28157.9004 | 31354.6992 | 33316.6016 | 32376.4004 | 32459.5
| 10.6 | 33163.3984 | 34657.6992 | 31812.9004 | 32714.5996 | 33502.6992 | 32141.1992 | 28300.3008 | 37709.1016 | 34285.6016 | 33549.3008 | 33561.6992 | 31245 | 30403.4004 | 33895.6016 | 28638.1992 | 36076.6992
| 10.9 | 36798.5 | 37873.1016 | 37882.8008 | 38808.6016 | 37651.1992 | 38756.5 | 33043 | 39121.1992 | 41262.1016 | 38729.5 | 36379.8984 | 34703.6992 | 35874.6016 | 36216 | 41426.1992 | 37876.8984
| 11.3 | 42079.3984 | 43147.3984 | 43361.3984 | 44710.8008 | 40805.3984 | 48091.6016 | 44952.3008 | 44456.8008 | 45325.3984 | 40125.6992 | 42930.6016 | 40439.5 | 45312.3984 | 44675.8008 | 40751.8984 | 43350.6992
| 11.8 | 51668.3984 | 48742.3008 | 46941.5 | 50537.6016 | 52667.8008 | 47463.6016 | 50303.8008 | 45846 | 51326.3984 | 50139.1992 | 50296 | 48364 | 49407.3008 | 48922.1016 | 50450 | 55157.3008
| 12.2 | 52882.1016 | 56330.5 | 54405.6016 | 57266.8008 | 49622.3984 | 52931.6016 | 55390.6992 | 56551 | 58366.3008 | 53209.3008 | 54172.3984 | 56204 | 57063.3008 | 54739.1992 | 57986.8984 | 56333.3008
| 12.6 | 63428 | 64542.8008 | 63113.3008 | 68278.2031 | 67556.1016 | 63915.3984 | 59281.3984 | 65088.6992 | 67344.7969 | 61565.1992 | 63152.1992 | 62121.3008 | 57807.1016 | 62860 | 66003.7969 | 69590.2969
| 13.1 | 67238 | 74208.8984 | 69610.7969 | 68512.5 | 68195.2969 | 77929.5 | 70344.7969 | 71493.7969 | 68858.7969 | 70240.6016 | 70647.5 | 69239.1016 | 73112.5 | 66011.7031 | 68792.5 | 74795.1016
| 13.6 | 80932.1016 | 84152.2969 | 81078.2031 | 85691.6016 | 75718.2031 | 79648.3984 | 82746.6016 | 76603.8984 | 83092 | 80353.7031 | 76204.3984 | 77607.7969 | 82408.5 | 77727.3984 | 83063 | 86056.5
| 14.1 | 86059.1016 | 88595.5 | 90947.5 | 89571.8984 | 82148.6016 | 92968 | 87580.2969 | 86565 | 92242.7969 | 81398 | 91391.5 | 87398.2031 | 86716 | 83996.7969 | 93172.1016 | 92850.8984
| 14.6 | 94579.1016 | 102777 | 103759 | 95694.5 | 101074 | 98452.2969 | 94099.7969 | 101446 | 98694.2969 | 97615.5 | 101468 | 100730 | 90064.3984 | 97431 | 97455.1016 | 101636
| 15.1 | 108677 | 108283 | 105140 | 114487 | 108471 | 110358 | 111118 | 108645 | 116249 | 107837 | 107872 | 113106 | 106613 | 106680 | 110859 | 113848
| 15.7 | 116477 | 122772 | 124908 | 129883 | 116347 | 122846 | 115372 | 116778 | 125200 | 117270 | 121940 | 121360 | 119384 | 113967 | 121333 | 131619
| 16.3 | 131301 | 137975 | 134501 | 135376 | 130737 | 141650 | 132802 | 137079 | 135535 | 130269 | 133044 | 131118 | 131875 | 134158 | 133973 | 138779
| 16.8 | 140739 | 149295 | 139857 | 144354 | 138367 | 148638 | 146869 | 146225 | 150674 | 146872 | 140614 | 143887 | 134586 | 137944 | 143415 | 149367
| 17.5 | 158996 | 158038 | 151297 | 165179 | 147410 | 161234 | 156695 | 159722 | 164035 | 159229 | 157952 | 157608 | 158476 | 160416 | 157452 | 168232
| 18.1 | 171428 | 178300 | 170191 | 176106 | 166106 | 172322 | 171718 | 176726 | 173517 | 171033 | 172896 | 171537 | 173620 | 164668 | 170626 | 175264
| 18.8 | 181669 | 187433 | 182371 | 183865 | 182039 | 187588 | 184527 | 189679 | 195246 | 179407 | 183395 | 184402 | 182512 | 181931 | 185315 | 188505
| 19.5 | 201597 | 197861 | 205916 | 200041 | 193685 | 202316 | 195928 | 194841 | 196915 | 197584 | 202105 | 198356 | 195819 | 196140 | 195635 | 197317
| 20.2 | 208321 | 216098 | 207896 | 219850 | 212223 | 215715 | 205124 | 206934 | 213614 | 208533 | 206804 | 214697 | 212036 | 207904 | 213248 | 217003
| 20.9 | 229146 | 227898 | 223273 | 228089 | 218477 | 220986 | 225832 | 226886 | 229828 | 225392 | 223356 | 223107 | 224734 | 229212 | 228438 | 221971
| 21.7 | 235454 | 232270 | 237591 | 240315 | 239593 | 238745 | 240387 | 241673 | 240635 | 238131 | 239744 | 239450 | 242902 | 239351 | 238475 | 241901
| 22.5 | 250190 | 249753 | 250591 | 247783 | 251523 | 252783 | 252766 | 252026 | 257216 | 255072 | 257435 | 248612 | 248869 | 257481 | 257425 | 254727
| 23.3 | 258769 | 269591 | 266051 | 267304 | 258296 | 253111 | 262555 | 265208 | 268628 | 261998 | 275563 | 267208 | 264482 | 264656 | 274118 | 266841
| 24.1 | 280580 | 286919 | 284609 | 280509 | 278170 | 279367 | 286787 | 276784 | 281569 | 283002 | 277598 | 273067 | 274845 | 277363 | 280192 | 278157
| 25 | 300829 | 294536 | 294556 | 290747 | 294805 | 288110 | 291073 | 289832 | 299645 | 294027 | 290872 | 293132 | 288939 | 295233 | 288653 | 289359
| 25.9 | 296862 | 302110 | 307290 | 309098 | 310898 | 301258 | 312471 | 308361 | 300076 | 308831 | 305366 | 301288 | 296180 | 304665 | 304705 | 301043
| 26.9 | 315169 | 310895 | 314785 | 313414 | 321511 | 311611 | 318313 | 308124 | 312320 | 322880 | 315441 | 311517 | 314244 | 317481 | 316872 | 305790
| 27.9 | 319861 | 316787 | 324765 | 314693 | 322304 | 309898 | 330468 | 323832 | 316112 | 331102 | 324909 | 324338 | 322756 | 322851 | 317613 | 319698
| 28.9 | 330512 | 329322 | 336081 | 326382 | 339243 | 327579 | 334797 | 327401 | 323713 | 335692 | 330772 | 334362 | 327609 | 332976 | 321634 | 326140
| 30 | 332535 | 330766 | 341189 | 331644 | 334001 | 334686 | 329131 | 336150 | 333053 | 349236 | 336007 | 332947 | 337312 | 341576 | 337741 | 324717
| 31.1 | 344683 | 329196 | 343417 | 341692 | 342901 | 330667 | 348067 | 349106 | 340047 | 351489 | 341627 | 344187 | 342850 | 347996 | 346375 | 340330
| 32.2 | 343770 | 346076 | 353401 | 335107 | 351794 | 339801 | 355297 | 355995 | 338443 | 354547 | 334704 | 354525 | 354597 | 358005 | 348090 | 341881
| 33.4 | 347295 | 339695 | 353023 | 346636 | 351220 | 339412 | 350467 | 353504 | 341230 | 364271 | 345981 | 349487 | 356411 | 348257 | 350843 | 333841
| 34.6 | 359642 | 335987 | 352065 | 342137 | 353942 | 345066 | 360114 | 358254 | 339009 | 371314 | 348489 | 358070 | 351468 | 359529 | 353758 | 342240
| 35.9 | 357976 | 343998 | 361837 | 336950 | 367633 | 349118 | 373844 | 349123 | 340998 | 365960 | 350710 | 361464 | 359891 | 364320 | 350243 | 346477
| 37.2 | 364970 | 345452 | 367198 | 345189 | 356091 | 337878 | 366500 | 346149 | 345400 | 359496 | 344857 | 361674 | 356380 | 359711 | 350116 | 336960
| 38.5 | 361152 | 330143 | 356051 | 345940 | 359716 | 331803 | 360201 | 347145 | 339837 | 360495 | 339330 | 359528 | 354221 | 352800 | 353549 | 331292
| 40 | 350910 | 331591 | 357226 | 343466 | 360814 | 345298 | 365755 | 341596 | 337645 | 359223 | 341526 | 354638 | 359611 | 355381 | 350120 | 333741
| 41.4 | 356987 | 326721 | 350583 | 334743 | 353244 | 332444 | 356152 | 355319 | 326667 | 360870 | 331443 | 346190 | 350992 | 353424 | 351374 | 326862
| 42.9 | 358411 | 328940 | 354153 | 335233 | 351316 | 323660 | 354927 | 335999 | 316326 | 352126 | 335191 | 349896 | 344553 | 350226 | 338818 | 322012
| 44.5 | 335759 | 319125 | 344367 | 320915 | 344759 | 322699 | 342945 | 328936 | 324301 | 347292 | 325897 | 338420 | 338257 | 342469 | 335859 | 313236
| 46.1 | 332705 | 311115 | 337177 | 311494 | 341283 | 312302 | 334449 | 322699 | 309151 | 344764 | 316242 | 348498 | 339269 | 327252 | 339247 | 310114
| 47.8 | 332231 | 296731 | 332566 | 310292 | 325074 | 293103 | 334384 | 323058 | 302457 | 329222 | 305643 | 337423 | 330949 | 329941 | 318122 | 290097
| 49.6 | 325445 | 298216 | 319914 | 300772 | 325458 | 294091 | 323854 | 307139 | 298089 | 326045 | 295340 | 320107 | 320189 | 314077 | 318338 | 292609
| 51.4 | 309295 | 284930 | 314857 | 298264 | 311239 | 291454 | 314216 | 300966 | 287165 | 316375 | 289776 | 321584 | 316023 | 307943 | 314542 | 284373
| 53.3 | 304920 | 276431 | 304664 | 287260 | 308221 | 273988 | 308286 | 291238 | 276407 | 308211 | 282400 | 308579 | 300821 | 300024 | 302414 | 278168
| 55.2 | 291277 | 264282 | 293675 | 278799 | 301501 | 264093 | 297184 | 288530 | 271365 | 296242 | 277327 | 300580 | 296534 | 287558 | 296460 | 271802
| 57.3 | 276855 | 259435 | 281798 | 269397 | 285771 | 260956 | 288995 | 274073 | 254334 | 287272 | 264341 | 285116 | 288618 | 281236 | 279668 | 248944
| 59.4 | 272784 | 245935 | 274110 | 251226 | 277035 | 249945 | 277139 | 263573 | 244647 | 274309 | 250695 | 273835 | 272017 | 266792 | 270959 | 240674
| 61.5 | 257693 | 237249 | 255876 | 249832 | 264489 | 238181 | 257471 | 256408 | 241141 | 264064 | 240275 | 264577 | 267202 | 252116 | 253291 | 238298
| 63.8 | 246198 | 226647 | 248142 | 243089 | 250065 | 221121 | 253961 | 243260 | 225905 | 248578 | 232084 | 244268 | 248355 | 245073 | 243389 | 224140
| 66.1 | 236689 | 213695 | 231602 | 221569 | 234339 | 214552 | 236294 | 229529 | 207686 | 236353 | 219404 | 234935 | 237541 | 229511 | 231553 | 213237
| 68.5 | 220233 | 205069 | 224168 | 213005 | 225165 | 200762 | 218918 | 207825 | 206514 | 220576 | 202643 | 225549 | 222983 | 215173 | 220037 | 201846
| 71 | 207003 | 192275 | 209828 | 198171 | 210618 | 189255 | 207311 | 198282 | 190764 | 212651 | 196597 | 205208 | 207262 | 201248 | 209990 | 196098
| 73.7 | 192811 | 176157 | 194226 | 186637 | 194632 | 179861 | 198331 | 193612 | 182493 | 195655 | 180955 | 195993 | 200251 | 187908 | 192545 | 181989
| 76.4 | 178160 | 168397 | 177790 | 171142 | 182138 | 168288 | 181269 | 176770 | 168575 | 183396 | 171254 | 181923 | 179290 | 176460 | 182758 | 165876
| 79.1 | 165834 | 158740 | 168383 | 163059 | 166042 | 159598 | 169975 | 167086 | 152678 | 170491 | 158811 | 166981 | 170847 | 162051 | 167442 | 158198
| 82 | 155131 | 146218 | 157334 | 150138 | 157419 | 142404 | 153036 | 154148 | 140833 | 157797 | 147043 | 154088 | 155201 | 150279 | 158341 | 146732
| 85.1 | 140558 | 137853 | 145903 | 143237 | 142233 | 134935 | 141300 | 142607 | 134827 | 143740 | 132862 | 145492 | 143417 | 141573 | 145672 | 132012
| 88.2 | 132251 | 123171 | 134529 | 133488 | 131367 | 123025 | 130721 | 128343 | 124433 | 132183 | 123427 | 135810 | 134760 | 129302 | 132510 | 123724
| 91.4 | 120600 | 114369 | 118977 | 119156 | 120840 | 112179 | 119732 | 118827 | 113712 | 119071 | 116397 | 121267 | 120024 | 115556 | 121201 | 114395
| 94.7 | 110746 | 106281 | 109531 | 107097 | 110526 | 101823 | 107967 | 105965 | 106304 | 110463 | 104542 | 112089 | 112427 | 106934 | 111221 | 105849
| 98.2 | 101692 | 97522.7031 | 103985 | 99741.2031 | 103480 | 97244.2969 | 98655.2969 | 98066.2969 | 97419.2031 | 104259 | 96271.2969 | 102657 | 102202 | 100808 | 101927 | 98305.2969
| 102 | 94478.2031 | 90492.2969 | 91880.2031 | 94023.3984 | 94120.7969 | 87048.8984 | 92343 | 93034.2031 | 88188.3984 | 92653.3984 | 88734 | 96795.8984 | 97514.2969 | 91245.1016 | 96097.3984 | 88823.2031
| 106 | 85308.2969 | 85143 | 89494.5 | 83754.2969 | 88543 | 82934.2031 | 85004.5 | 85432.7969 | 85342.6016 | 89191 | 84150.6016 | 86041.3984 | 89356.5 | 85143.2031 | 87004 | 83188.3984
| 109 | 80727.6016 | 75306.2969 | 79678.8984 | 79733.6016 | 80858.1016 | 74925.6016 | 79484.6016 | 79290.2969 | 76417.7969 | 79254 | 75040.3984 | 79042 | 80378.7969 | 78841.2031 | 79204.3984 | 74972.2031
| 113 | 74250.3984 | 69410.6016 | 75990.2969 | 71741.2969 | 74552 | 70299.1016 | 71492.8984 | 72498 | 71432.2031 | 75160.1016 | 70968.6016 | 75005.1016 | 74811.2031 | 71147.1016 | 75846.2969 | 70485.2969
| 118 | 70252 | 65981.3984 | 70825.8984 | 68815.7969 | 70882.2969 | 65416.8008 | 68224.6016 | 69629.6016 | 65492.5 | 70181.7969 | 66208.7031 | 70197.7031 | 71303.5 | 67919.3984 | 71707.6016 | 63840.8984
Dp | Scantime | Counts_for_each_SMPS_Scan
3.1349 | 1 | 891 | 970 | 927 | 915 | 899 | 938 | 933 | 916 | 921 | 888 | 931 | 887 | 966 | 959 | 947 | 939
3.2262 | 2 | 623 | 616 | 544 | 569 | 611 | 598 | 556 | 554 | 601 | 589 | 562 | 603 | 636 | 605 | 647 | 606
3.3201 | 3 | 401 | 375 | 400 | 412 | 396 | 444 | 462 | 382 | 383 | 401 | 394 | 391 | 417 | 420 | 405 | 431
3.4168 | 4 | 287 | 287 | 283 | 273 | 313 | 320 | 321 | 282 | 324 | 295 | 286 | 273 | 297 | 287 | 279 | 294
3.5163 | 5 | 237 | 214 | 226 | 244 | 222 | 250 | 210 | 220 | 236 | 202 | 226 | 194 | 234 | 232 | 224 | 255
3.6187 | 6 | 172 | 192 | 170 | 178 | 175 | 199 | 191 | 202 | 205 | 158 | 172 | 171 | 153 | 163 | 206 | 174
3.7241 | 7 | 154 | 147 | 136 | 151 | 150 | 180 | 157 | 160 | 193 | 122 | 164 | 122 | 170 | 159 | 157 | 149
3.8326 | 8 | 118 | 134 | 117 | 134 | 143 | 141 | 136 | 143 | 179 | 130 | 111 | 126 | 98 | 139 | 161 | 134
3.9443 | 9 | 113 | 113 | 140 | 115 | 122 | 123 | 136 | 138 | 119 | 115 | 110 | 97 | 108 | 111 | 128 | 122
4.0592 | 10 | 95 | 89 | 92 | 103 | 127 | 93 | 101 | 135 | 139 | 110 | 94 | 101 | 92 | 101 | 107 | 115
4.1776 | 11 | 94 | 101 | 91 | 79 | 96 | 112 | 114 | 112 | 113 | 84 | 85 | 94 | 104 | 103 | 94 | 113
4.2993 | 12 | 97 | 90 | 86 | 80 | 87 | 101 | 111 | 104 | 113 | 93 | 79 | 84 | 62 | 94 | 98 | 109
4.4247 | 13 | 72 | 80 | 75 | 84 | 87 | 107 | 79 | 97 | 108 | 82 | 71 | 65 | 75 | 70 | 82 | 114
4.5537 | 14 | 72 | 68 | 61 | 95 | 77 | 98 | 73 | 98 | 94 | 96 | 73 | 65 | 54 | 88 | 93 | 85
4.6864 | 15 | 63 | 66 | 82 | 72 | 72 | 71 | 82 | 89 | 91 | 66 | 56 | 69 | 79 | 80 | 73 | 84
4.8231 | 16 | 52 | 73 | 66 | 74 | 75 | 78 | 68 | 87 | 58 | 67 | 70 | 55 | 64 | 67 | 72 | 87
4.9638 | 17 | 61 | 60 | 73 | 56 | 74 | 88 | 82 | 79 | 64 | 65 | 48 | 61 | 52 | 64 | 80 | 60
5.1086 | 18 | 46 | 61 | 63 | 56 | 55 | 70 | 64 | 61 | 66 | 51 | 58 | 53 | 44 | 62 | 67 | 70
5.2577 | 19 | 53 | 55 | 63 | 72 | 54 | 79 | 69 | 66 | 75 | 45 | 56 | 61 | 71 | 57 | 64 | 69
5.4111 | 20 | 67 | 66 | 54 | 70 | 69 | 61 | 60 | 63 | 55 | 53 | 58 | 46 | 41 | 57 | 50 | 55
5.569 | 21 | 39 | 63 | 50 | 66 | 53 | 68 | 47 | 53 | 59 | 38 | 43 | 52 | 42 | 59 | 57 | 51
5.7315 | 22 | 65 | 53 | 52 | 60 | 63 | 59 | 72 | 67 | 54 | 50 | 54 | 42 | 61 | 60 | 46 | 47
5.8989 | 23 | 50 | 51 | 71 | 56 | 56 | 56 | 59 | 51 | 52 | 31 | 42 | 49 | 54 | 55 | 58 | 67
6.0711 | 24 | 54 | 47 | 35 | 46 | 53 | 31 | 50 | 67 | 63 | 63 | 43 | 39 | 43 | 52 | 59 | 47
6.2485 | 25 | 45 | 35 | 49 | 39 | 57 | 50 | 44 | 68 | 53 | 43 | 26 | 38 | 49 | 63 | 51 | 59
6.431 | 26 | 54 | 47 | 45 | 52 | 53 | 45 | 43 | 50 | 64 | 43 | 46 | 45 | 38 | 48 | 48 | 57
6.619 | 27 | 55 | 51 | 50 | 56 | 47 | 56 | 51 | 56 | 58 | 55 | 34 | 48 | 50 | 46 | 45 | 40
6.8124 | 28 | 36 | 44 | 52 | 51 | 43 | 61 | 43 | 68 | 51 | 51 | 36 | 38 | 53 | 52 | 51 | 44
7.0115 | 29 | 54 | 58 | 49 | 47 | 43 | 52 | 57 | 42 | 54 | 40 | 36 | 49 | 37 | 56 | 56 | 46
7.2166 | 30 | 56 | 44 | 45 | 43 | 61 | 49 | 45 | 55 | 51 | 49 | 57 | 39 | 37 | 48 | 51 | 64
7.4275 | 31 | 48 | 55 | 59 | 48 | 41 | 60 | 53 | 50 | 52 | 54 | 56 | 42 | 47 | 67 | 50 | 58
7.6449 | 32 | 50 | 47 | 51 | 59 | 63 | 65 | 53 | 60 | 58 | 61 | 47 | 51 | 55 | 46 | 55 | 56
7.8686 | 33 | 55 | 53 | 51 | 67 | 63 | 75 | 63 | 63 | 74 | 43 | 45 | 53 | 54 | 64 | 56 | 61
8.0988 | 34 | 65 | 59 | 62 | 60 | 61 | 65 | 63 | 53 | 61 | 64 | 74 | 61 | 61 | 62 | 72 | 72
8.3358 | 35 | 68 | 74 | 63 | 78 | 59 | 71 | 71 | 66 | 63 | 77 | 80 | 68 | 60 | 88 | 72 | 90
8.5799 | 36 | 75 | 66 | 90 | 76 | 74 | 84 | 83 | 76 | 87 | 73 | 82 | 86 | 77 | 88 | 72 | 88
8.8313 | 37 | 93 | 82 | 104 | 93 | 108 | 101 | 83 | 90 | 93 | 66 | 85 | 83 | 80 | 92 | 82 | 108
9.09 | 38 | 117 | 84 | 109 | 118 | 109 | 124 | 91 | 110 | 95 | 94 | 110 | 90 | 93 | 111 | 111 | 104
9.3565 | 39 | 127 | 121 | 123 | 123 | 112 | 116 | 113 | 119 | 113 | 106 | 123 | 153 | 105 | 117 | 121 | 138
9.6307 | 40 | 146 | 144 | 131 | 142 | 118 | 138 | 131 | 126 | 149 | 137 | 150 | 158 | 139 | 140 | 147 | 131
9.9131 | 41 | 140 | 192 | 145 | 158 | 149 | 157 | 146 | 174 | 166 | 157 | 166 | 157 | 160 | 178 | 178 | 153
10.2039 | 42 | 218 | 204 | 182 | 197 | 179 | 208 | 194 | 203 | 194 | 180 | 164 | 186 | 206 | 209 | 201 | 217
10.5033 | 43 | 207 | 235 | 200 | 216 | 210 | 206 | 189 | 238 | 239 | 228 | 224 | 217 | 188 | 235 | 185 | 243
10.8116 | 44 | 270 | 240 | 266 | 266 | 262 | 256 | 236 | 273 | 257 | 265 | 254 | 227 | 246 | 242 | 263 | 264
11.1292 | 45 | 270 | 299 | 281 | 294 | 286 | 312 | 296 | 313 | 318 | 280 | 286 | 270 | 301 | 281 | 328 | 302
11.4561 | 46 | 353 | 377 | 352 | 361 | 358 | 384 | 342 | 348 | 383 | 329 | 345 | 337 | 357 | 369 | 330 | 344
11.7927 | 47 | 425 | 380 | 388 | 414 | 421 | 395 | 406 | 379 | 408 | 411 | 413 | 394 | 411 | 400 | 403 | 454
12.1395 | 48 | 449 | 475 | 448 | 472 | 428 | 449 | 466 | 471 | 492 | 451 | 452 | 486 | 498 | 463 | 482 | 486
12.4965 | 49 | 519 | 546 | 537 | 602 | 526 | 533 | 513 | 579 | 576 | 547 | 543 | 494 | 490 | 534 | 585 | 574
12.8643 | 50 | 635 | 648 | 607 | 600 | 652 | 645 | 601 | 621 | 641 | 578 | 601 | 642 | 586 | 602 | 604 | 669
13.2429 | 51 | 700 | 747 | 725 | 732 | 684 | 793 | 712 | 715 | 695 | 715 | 722 | 688 | 750 | 649 | 722 | 740
13.633 | 52 | 805 | 860 | 828 | 876 | 773 | 797 | 853 | 769 | 847 | 825 | 782 | 796 | 847 | 815 | 839 | 897
14.0346 | 53 | 924 | 938 | 944 | 934 | 857 | 975 | 918 | 902 | 957 | 836 | 950 | 917 | 911 | 847 | 966 | 985
14.4482 | 54 | 988 | 1062 | 1108 | 1000 | 1044 | 1066 | 1016 | 1062 | 1119 | 1051 | 1107 | 1065 | 945 | 1068 | 1062 | 1071
14.8745 | 55 | 1162 | 1224 | 1210 | 1215 | 1205 | 1208 | 1177 | 1218 | 1208 | 1212 | 1212 | 1216 | 1166 | 1164 | 1210 | 1218
15.3134 | 56 | 1348 | 1335 | 1284 | 1444 | 1298 | 1318 | 1316 | 1343 | 1408 | 1260 | 1304 | 1385 | 1314 | 1308 | 1351 | 1460
15.7655 | 57 | 1430 | 1538 | 1592 | 1634 | 1484 | 1577 | 1461 | 1453 | 1548 | 1480 | 1519 | 1526 | 1479 | 1419 | 1520 | 1639
16.2312 | 58 | 1687 | 1761 | 1722 | 1733 | 1689 | 1800 | 1682 | 1748 | 1746 | 1650 | 1718 | 1653 | 1692 | 1708 | 1737 | 1762
16.7108 | 59 | 1835 | 1929 | 1833 | 1835 | 1791 | 1976 | 1963 | 1878 | 1945 | 1887 | 1770 | 1857 | 1790 | 1818 | 1798 | 1975
17.2052 | 60 | 2061 | 2156 | 2046 | 2230 | 1962 | 2089 | 2077 | 2233 | 2213 | 2166 | 2120 | 2123 | 2021 | 2123 | 2122 | 2183
17.7142 | 61 | 2424 | 2374 | 2245 | 2416 | 2235 | 2399 | 2341 | 2310 | 2405 | 2353 | 2382 | 2326 | 2337 | 2285 | 2361 | 2430
18.2387 | 62 | 2537 | 2692 | 2570 | 2626 | 2534 | 2614 | 2610 | 2677 | 2624 | 2564 | 2556 | 2599 | 2621 | 2531 | 2551 | 2662
18.7792 | 63 | 2813 | 2867 | 2798 | 2834 | 2786 | 2889 | 2835 | 2913 | 2979 | 2760 | 2874 | 2849 | 2848 | 2759 | 2840 | 2912
19.336 | 64 | 3103 | 3113 | 3243 | 3137 | 3100 | 3207 | 3083 | 3062 | 3151 | 3103 | 3185 | 3068 | 3069 | 3094 | 3115 | 3169
19.9099 | 65 | 3409 | 3539 | 3355 | 3484 | 3348 | 3444 | 3311 | 3397 | 3434 | 3402 | 3335 | 3466 | 3409 | 3380 | 3426 | 3371
20.5009 | 66 | 3784 | 3733 | 3682 | 3903 | 3620 | 3722 | 3592 | 3669 | 3730 | 3631 | 3599 | 3777 | 3680 | 3704 | 3702 | 3776
21.1104 | 67 | 4002 | 4027 | 4005 | 4043 | 3994 | 3975 | 4087 | 4024 | 4109 | 4021 | 4047 | 3982 | 4073 | 4040 | 4107 | 4020
21.738 | 68 | 4329 | 4240 | 4358 | 4364 | 4371 | 4372 | 4407 | 4428 | 4397 | 4385 | 4407 | 4378 | 4432 | 4410 | 4359 | 4398
22.385 | 69 | 4664 | 4616 | 4698 | 4666 | 4711 | 4753 | 4652 | 4766 | 4852 | 4724 | 4817 | 4647 | 4662 | 4820 | 4835 | 4753
23.0518 | 70 | 5038 | 5124 | 5039 | 5044 | 4963 | 4911 | 5070 | 4987 | 5058 | 5059 | 5194 | 5019 | 5012 | 5038 | 5241 | 5046
23.7391 | 71 | 5286 | 5615 | 5554 | 5430 | 5388 | 5282 | 5566 | 5546 | 5548 | 5503 | 5599 | 5469 | 5400 | 5495 | 5494 | 5558
24.4475 | 72 | 6001 | 6051 | 6044 | 5968 | 5819 | 5914 | 6000 | 5729 | 5942 | 5949 | 5847 | 5861 | 5818 | 5872 | 5942 | 5829
25.178 | 73 | 6466 | 6284 | 6197 | 6227 | 6412 | 6172 | 6219 | 6285 | 6434 | 6298 | 6237 | 6226 | 6187 | 6333 | 6155 | 6205
25.9305 | 74 | 6531 | 6646 | 6745 | 6739 | 6753 | 6616 | 6877 | 6682 | 6573 | 6785 | 6708 | 6596 | 6499 | 6637 | 6641 | 6644
26.7069 | 75 | 6939 | 6874 | 7145 | 7114 | 7249 | 6857 | 7095 | 6974 | 7039 | 7188 | 7070 | 6945 | 6934 | 7128 | 7127 | 6699
27.5067 | 76 | 7479 | 7349 | 7434 | 7288 | 7378 | 7296 | 7471 | 7365 | 7144 | 7548 | 7455 | 7503 | 7522 | 7388 | 7346 | 7486
28.3317 | 77 | 7737 | 7725 | 7865 | 7621 | 8026 | 7590 | 8171 | 7912 | 7742 | 8108 | 7733 | 7878 | 7751 | 7913 | 7677 | 7688
29.1827 | 78 | 8125 | 8185 | 8369 | 8096 | 8275 | 8116 | 8156 | 8007 | 8048 | 8318 | 8228 | 8175 | 8107 | 8232 | 7918 | 7963
30.0599 | 79 | 8488 | 8375 | 8632 | 8432 | 8512 | 8525 | 8360 | 8568 | 8489 | 8877 | 8602 | 8483 | 8618 | 8697 | 8680 | 8306
30.9646 | 80 | 8995 | 8556 | 8906 | 8866 | 8868 | 8656 | 8994 | 9059 | 8833 | 9185 | 8870 | 8912 | 8898 | 9022 | 9065 | 8841
31.8977 | 81 | 9083 | 9091 | 9331 | 9022 | 9419 | 8948 | 9418 | 9419 | 9129 | 9330 | 8974 | 9452 | 9404 | 9549 | 9266 | 9188
32.8598 | 82 | 9534 | 9414 | 9708 | 9300 | 9556 | 9318 | 9706 | 9688 | 9140 | 9907 | 9305 | 9492 | 9771 | 9624 | 9546 | 9125
33.8526 | 83 | 9952 | 9472 | 9963 | 9725 | 9884 | 9628 | 9938 | 10018 | 9733 | 10401 | 9713 | 9988 | 9996 | 9947 | 9941 | 9481
34.8772 | 84 | 10291 | 9785 | 10218 | 9839 | 10310 | 9848 | 10446 | 10400 | 9753 | 10661 | 10199 | 10442 | 10173 | 10333 | 10118 | 9938
35.9336 | 85 | 10619 | 10190 | 10674 | 9961 | 10912 | 10406 | 11124 | 10273 | 10062 | 10772 | 10340 | 10682 | 10596 | 10799 | 10355 | 10234
37.0243 | 86 | 10974 | 10419 | 11104 | 10340 | 10804 | 10274 | 11096 | 10429 | 10427 | 10918 | 10372 | 10812 | 10682 | 10915 | 10648 | 10248
38.1491 | 87 | 11275 | 10374 | 11233 | 10825 | 11118 | 10338 | 11219 | 10914 | 10674 | 11211 | 10722 | 11342 | 11124 | 11162 | 10916 | 10264
39.3104 | 88 | 11333 | 10504 | 11222 | 10972 | 11424 | 10682 | 11533 | 10846 | 10653 | 11419 | 10709 | 11257 | 11362 | 11115 | 11255 | 10614
40.5094 | 89 | 11502 | 10830 | 11599 | 11098 | 11710 | 11271 | 11808 | 11309 | 11007 | 11668 | 11014 | 11514 | 11663 | 11426 | 11329 | 10777
41.7467 | 90 | 11890 | 10841 | 11615 | 11211 | 11786 | 10994 | 11936 | 11788 | 10775 | 12061 | 11126 | 11570 | 11627 | 11956 | 11710 | 10900
43.024 | 91 | 12218 | 11187 | 12153 | 11383 | 11935 | 10987 | 12028 | 11378 | 10742 | 11956 | 11387 | 11876 | 11745 | 11842 | 11490 | 10940
44.3426 | 92 | 11773 | 11229 | 12052 | 11226 | 12103 | 11358 | 11996 | 11395 | 11234 | 12144 | 11341 | 11799 | 11743 | 12010 | 11764 | 11043
45.7049 | 93 | 11838 | 11224 | 12193 | 11187 | 12120 | 11247 | 11971 | 11745 | 11269 | 12378 | 11432 | 12309 | 12212 | 11872 | 11921 | 11062
47.1116 | 94 | 12190 | 10947 | 12139 | 11308 | 12199 | 10972 | 12202 | 11772 | 11094 | 12274 | 11257 | 12601 | 12239 | 11852 | 12116 | 10914
48.5645 | 95 | 12383 | 11158 | 12209 | 11512 | 12088 | 10830 | 12335 | 11711 | 11295 | 12072 | 11330 | 12309 | 12108 | 12070 | 11814 | 10768
50.0655 | 96 | 12091 | 11194 | 12008 | 11395 | 12252 | 11258 | 12258 | 11658 | 11150 | 12389 | 11169 | 12055 | 12115 | 11930 | 12158 | 11075
51.6159 | 97 | 11913 | 10920 | 12178 | 11418 | 11944 | 11097 | 12054 | 11560 | 11062 | 12165 | 11119 | 12479 | 12156 | 11809 | 11999 | 10932
53.2189 | 98 | 12011 | 10898 | 11991 | 11311 | 12105 | 10821 | 12168 | 11391 | 10869 | 12160 | 11173 | 12120 | 11875 | 11850 | 11906 | 10961
54.8753 | 99 | 11775 | 10489 | 11831 | 11202 | 12054 | 10520 | 11951 | 11688 | 11006 | 11927 | 10971 | 12078 | 11877 | 11556 | 11818 | 10843
56.587 | 100 | 11486 | 10666 | 11652 | 11097 | 11970 | 10746 | 11758 | 11254 | 10480 | 11844 | 10988 | 11774 | 11761 | 11559 | 11714 | 10595
58.3577 | 101 | 11261 | 10546 | 11460 | 10710 | 11489 | 10545 | 11805 | 11006 | 10233 | 11489 | 10707 | 11548 | 11605 | 11245 | 11297 | 9975
60.1877 | 102 | 11195 | 10056 | 11253 | 10583 | 11483 | 10303 | 11296 | 11110 | 10280 | 11335 | 10315 | 11379 | 11391 | 11138 | 11137 | 10154
62.0814 | 103 | 10960 | 10075 | 10842 | 10600 | 11247 | 10085 | 10986 | 10788 | 10196 | 11143 | 10224 | 11106 | 11236 | 10600 | 10681 | 9966
64.0398 | 104 | 10672 | 9883 | 10670 | 10514 | 10755 | 9510 | 10979 | 10540 | 9692 | 10803 | 9973 | 10574 | 10736 | 10570 | 10565 | 9711
66.0656 | 105 | 10457 | 9449 | 10263 | 9794 | 10362 | 9474 | 10500 | 10143 | 9208 | 10422 | 9740 | 10455 | 10546 | 10136 | 10235 | 9464
68.1632 | 106 | 9934 | 9274 | 10149 | 9499 | 10195 | 9078 | 9898 | 9401 | 9262 | 10084 | 9155 | 10065 | 10018 | 9830 | 9903 | 9039
70.3337 | 107 | 9492 | 8829 | 9656 | 9330 | 9698 | 8830 | 9446 | 9211 | 8918 | 9712 | 9026 | 9693 | 9697 | 9257 | 9698 | 8933
72.5803 | 108 | 9338 | 8409 | 9305 | 8765 | 9282 | 8467 | 9314 | 8942 | 8615 | 9313 | 8712 | 9127 | 9290 | 8873 | 9249 | 8797
74.9075 | 109 | 8554 | 8115 | 8674 | 8370 | 8826 | 8092 | 9080 | 8765 | 8082 | 8984 | 8170 | 8863 | 8943 | 8579 | 8758 | 8145
77.3179 | 110 | 8159 | 7851 | 8222 | 7988 | 8371 | 7783 | 8268 | 8194 | 7773 | 8351 | 7881 | 8281 | 8332 | 8027 | 8312 | 7679
79.8159 | 111 | 7952 | 7382 | 7871 | 7541 | 7835 | 7462 | 7987 | 7814 | 7174 | 8009 | 7502 | 7906 | 7964 | 7701 | 7969 | 7432
82.404 | 112 | 7416 | 7027 | 7592 | 7317 | 7542 | 6888 | 7324 | 7409 | 6759 | 7578 | 7054 | 7412 | 7481 | 7162 | 7601 | 7026
85.0868 | 113 | 6896 | 6767 | 7161 | 7031 | 6985 | 6604 | 6966 | 7013 | 6611 | 7037 | 6491 | 7150 | 7030 | 6982 | 7142 | 6479
87.8686 | 114 | 6590 | 6189 | 6786 | 6638 | 6490 | 6137 | 6468 | 6370 | 6229 | 6590 | 6161 | 6782 | 6683 | 6473 | 6588 | 6165
90.7537 | 115 | 6160 | 5782 | 6071 | 6175 | 6221 | 5708 | 6128 | 6082 | 5802 | 6134 | 5881 | 6289 | 6224 | 5922 | 6206 | 5779
93.747 | 116 | 5791 | 5473 | 5728 | 5480 | 5811 | 5336 | 5689 | 5531 | 5495 | 5624 | 5485 | 5703 | 5825 | 5462 | 5786 | 5533
96.8535 | 117 | 5270 | 5210 | 5399 | 5280 | 5300 | 5021 | 5162 | 5137 | 5115 | 5544 | 5063 | 5467 | 5280 | 5353 | 5382 | 5131
100.0783 | 118 | 5147 | 4890 | 5012 | 5037 | 5119 | 4818 | 4928 | 4947 | 4842 | 4993 | 4803 | 5146 | 5186 | 4866 | 5094 | 4914
103.4272 | 119 | 4632 | 4531 | 4672 | 4607 | 4791 | 4362 | 4622 | 4608 | 4491 | 4711 | 4517 | 4720 | 4934 | 4619 | 4806 | 4438
106.906 | 120 | 4418 | 4258 | 4571 | 4341 | 4492 | 4288 | 4378 | 4410 | 4294 | 4539 | 4227 | 4391 | 4519 | 4387 | 4398 | 4228
110.5211 | 121 | 4171 | 3903 | 4120 | 4137 | 4204 | 3820 | 4090 | 4068 | 3956 | 4124 | 3875 | 4167 | 4138 | 4078 | 4204 | 3857
114.2782 | 122 | 3914 | 3706 | 4019 | 3740 | 3918 | 3684 | 3785 | 3911 | 3750 | 3954 | 3677 | 3950 | 4002 | 3683 | 3930 | 3744
118.1863 | 123 | 3739 | 3468 | 3765 | 3673 | 3758 | 3471 | 3603 | 3703 | 3499 | 3800 | 3588 | 3729 | 3752 | 3665 | 3864 | 3399
122.2526 | 124 | 3657 | 3216 | 3599 | 3505 | 3611 | 3290 | 3492 | 3385 | 3320 | 3590 | 3421 | 3349 | 3608 | 3446 | 3546 | 3367
126.4836 | 125 | 3898 | 3648 | 3827 | 3665 | 3871 | 3641 | 3785 | 3638 | 3418 | 3880 | 3712 | 3806 | 3867 | 3814 | 3926 | 3553
130.8886 | 126 | 5299 | 4858 | 5367 | 5074 | 5191 | 4901 | 5062 | 5002 | 4909 | 5143 | 4904 | 5177 | 5262 | 5219 | 5345 | 4920
135.4767 | 127 | 7884 | 6770 | 7747 | 7458 | 7705 | 6922 | 7565 | 7116 | 7081 | 7679 | 7219 | 7430 | 7649 | 7372 | 8052 | 7171
140.2573 | 128 | 9790 | 8847 | 9838 | 9483 | 10056 | 8972 | 9460 | 9297 | 9225 | 9836 | 9149 | 9725 | 10115 | 9190 | 9960 | 9095
145.2401 | 129 | 10521 | 9731 | 10466 | 10417 | 10753 | 9846 | 10350 | 10509 | 9812 | 10825 | 9832 | 10643 | 10764 | 9874 | 10620 | 9811
150.4375 | 130 | 9993 | 9329 | 9990 | 9855 | 10229 | 9552 | 9748 | 9667 | 9425 | 10220 | 9495 | 9955 | 9915 | 9637 | 10046 | 9552
155.8611 | 131 | 8231 | 8151 | 8486 | 8345 | 8279 | 7880 | 8312 | 8095 | 7955 | 8272 | 8118 | 8258 | 8387 | 8126 | 8331 | 7937
161.5214 | 132 | 6207 | 5911 | 6137 | 6124 | 5974 | 6014 | 5999 | 5965 | 5813 | 6095 | 6034 | 6315 | 6171 | 5839 | 6362 | 5787
167.4319 | 133 | 3943 | 3847 | 3961 | 3872 | 4057 | 4016 | 3959 | 3952 | 4038 | 3957 | 3994 | 3921 | 3957 | 3934 | 3930 | 3941
173.6091 | 134 | 2425 | 2392 | 2442 | 2304 | 2391 | 2506 | 2516 | 2550 | 2411 | 2496 | 2419 | 2410 | 2378 | 2349 | 2225 | 2438
180.065 | 135 | 1428 | 1460 | 1477 | 1446 | 1522 | 1434 | 1473 | 1499 | 1451 | 1512 | 1516 | 1502 | 1522 | 1458 | 1505 | 1432
| Scantime | CCNCounts_for_each_CCNC_Scan
| 1 | 1089.57 | 1012.44 | 1043.68 | 1084.6999 | 1071.84 | 1218.39 | 1158.71 | 1248.26 | 1279.11 | 1533.0601 | 1478.15 | 1402.99 | 1444.73 | 1704.65 | 1587.28 | 1678.8101
| 2 | 657.28 | 579.49 | 599.92 | 576.81 | 1538.3101 | 1634.17 | 1633 | 1676.52 | 1592.33 | 1870.22 | 1892.18 | 1641.14 | 1797.98 | 2007.16 | 1791.4 | 1975.71
| 3 | 337.38 | 307.07 | 292.48 | 323.03 | 1971.85 | 1868.79 | 1491.65 | 1621.26 | 1861.58 | 2070.0901 | 1971.86 | 1807.43 | 1998.71 | 1931.27 | 1814.46 | 1834.9
| 4 | 164.52 | 160.71 | 158.6 | 180.47 | 1954.01 | 1152.8001 | 780.23 | 798.03 | 1691.96 | 1457.08 | 1373.17 | 1158.28 | 1195.53 | 1542.26 | 1306.13 | 1167.77
| 5 | 83.14 | 83.09 | 67.7 | 70.02 | 1257.6899 | 477.35 | 407.73 | 394.25 | 788.04 | 725.08 | 599.38 | 487.68 | 511.35 | 894.15 | 633.96 | 544.18
| 6 | 38.4 | 40.43 | 27.72 | 33.16 | 493.22 | 211.16 | 189.15 | 155.75 | 359.21 | 270.33 | 230.11 | 203.44 | 224.09 | 333.13 | 231 | 228.81
| 7 | 17.02 | 18.84 | 15.29 | 9.89 | 178.48 | 76.08 | 65.12 | 73.79 | 112.56 | 104.47 | 99.19 | 79.96 | 83.07 | 119.93 | 73.3 | 63.37
| 8 | 6.23 | 10.75 | 5.37 | 7.15 | 80.77 | 35.81 | 31.93 | 32.96 | 38.48 | 37.45 | 33.06 | 46.33 | 29.49 | 31.2 | 30.48 | 34.03
| 9 | 6.25 | 6.26 | 4.48 | 1.8 | 19.73 | 21.41 | 5.34 | 8.9 | 26.69 | 15.16 | 16.76 | 13.36 | 16.96 | 13.52 | 19.63 | 16.17
| 10 | 3.58 | 5.38 | 5.38 | 4.47 | 9.76 | 13.45 | 5.35 | 9.02 | 11.56 | 8.86 | 7.1 | 5.35 | 8.02 | 7.16 | 7.18 | 9.87
| 11 | 2.67 | 0.89 | 4.48 | 5.28 | 8.07 | 3.56 | 4.45 | 4.44 | 10.7 | 8.88 | 2.66 | 5.34 | 8.04 | 9.82 | 6.24 | 7.14
| 12 | 0.89 | 0.9 | 0.9 | 1.79 | 8.93 | 10.67 | 2.67 | 2.68 | 2.68 | 4.43 | 4.43 | 5.34 | 6.22 | 5.37 | 3.57 | 1.78
| 13 | 0.89 | 2.68 | 1.78 | 2.68 | 5.37 | 6.24 | 8.01 | 4.51 | 2.67 | 3.56 | 1.77 | 3.55 | 8.04 | 3.55 | 5.37 | 2.67
| 14 | 2.67 | 2.68 | 2.69 | 3.6 | 3.59 | 1.78 | 3.56 | 4.46 | 0.89 | 1.78 | 0.89 | 1.78 | 1.8 | 1.78 | 0 | 3.57
| 15 | 1.79 | 2.68 | 0 | 3.58 | 2.68 | 1.78 | 1.78 | 1.78 | 3.57 | 1.77 | 2.67 | 2.67 | 1.79 | 1.77 | 0 | 2.67
| 16 | 2.67 | 0.9 | 0 | 0.89 | 4.48 | 2.68 | 1.76 | 0.89 | 1.78 | 3.56 | 0 | 2.65 | 1.8 | 3.57 | 1.79 | 2.67
| 17 | 0.89 | 0.89 | 1.83 | 0.89 | 2.68 | 1.78 | 0.89 | 0.89 | 0.89 | 2.68 | 0.89 | 0.89 | 1.78 | 1.79 | 2.68 | 0.89
| 18 | 0 | 0 | 0 | 1.8 | 2.69 | 1.78 | 0 | 0 | 0.89 | 1.78 | 2.66 | 2.66 | 1.79 | 1.79 | 2.68 | 0
| 19 | 0 | 0 | 0 | 1.78 | 1.79 | 1.79 | 0.89 | 0 | 2.67 | 2.68 | 0.89 | 0.89 | 2.67 | 4.46 | 0 | 0
| 20 | 1.78 | 0 | 2.68 | 0 | 0.9 | 1.78 | 0.89 | 1.78 | 0.9 | 2.63 | 4.43 | 0 | 0 | 0.89 | 2.69 | 1.78
| 21 | 0.89 | 0 | 0.89 | 0 | 1.79 | 2.68 | 1.78 | 0.89 | 0 | 3.55 | 1.78 | 1.78 | 0 | 3.56 | 0.9 | 2.67
| 22 | 0 | 0.89 | 0 | 0 | 0.89 | 3.56 | 0.89 | 0 | 2.66 | 0 | 2.66 | 0 | 0 | 1.79 | 0.89 | 0
| 23 | 1.78 | 0 | 0 | 0.89 | 1.79 | 0 | 1.78 | 0.89 | 0.89 | 0 | 0 | 0 | 1.79 | 0 | 0.89 | 0.89
| 24 | 1.79 | 1.79 | 0.89 | 0.89 | 0 | 0 | 2.67 | 0.89 | 0.88 | 0 | 0 | 0.89 | 0 | 2.7 | 1.78 | 0.9
| 25 | 1.78 | 1.79 | 0 | 2.68 | 0.89 | 2.66 | 2.66 | 0 | 0 | 0 | 0.89 | 0.89 | 1.79 | 0 | 0 | 0.89
| 26 | 0.88 | 0 | 0.9 | 0.89 | 0 | 0 | 0 | 0 | 2.68 | 0 | 0 | 0.89 | 0.89 | 1.81 | 0.9 | 0.89
| 27 | 0.89 | 0 | 1.78 | 0.89 | 0 | 0 | 0 | 1.78 | 0 | 0.89 | 0 | 0.89 | 1.79 | 0 | 0.89 | 0
| 28 | 2.68 | 0 | 1.79 | 1.79 | 0.9 | 0 | 2.67 | 0.89 | 0.89 | 3.55 | 1.78 | 1.78 | 2.63 | 1.79 | 1.79 | 0.89
| 29 | 2.68 | 0.89 | 2.69 | 0 | 0.89 | 0.89 | 0.88 | 1.78 | 0.89 | 1.78 | 0 | 2.67 | 0.89 | 0.9 | 0 | 0
| 30 | 1.79 | 0.89 | 0 | 0.89 | 0 | 0 | 0.89 | 0 | 0 | 0 | 0.89 | 0 | 0 | 0.89 | 0 | 0
| 31 | 0.89 | 1.79 | 0 | 0 | 0.9 | 0 | 0.89 | 0.89 | 2.69 | 0 | 1.76 | 0 | 0.89 | 1.79 | 0.89 | 0
| 32 | 0.89 | 1.78 | 0.89 | 0 | 0.89 | 1.78 | 0 | 0 | 0.89 | 1.77 | 0.89 | 0 | 0 | 0.89 | 0 | 0
| 33 | 0 | 0 | 0 | 0 | 0.89 | 0.89 | 0 | 0 | 0.89 | 0 | 0.89 | 0.89 | 0.9 | 0.89 | 0.9 | 1.78
| 34 | 0.89 | 0.9 | 0.9 | 0 | 0 | 0 | 0.89 | 0.89 | 0 | 0 | 0 | 0 | 3.57 | 0.89 | 0 | 0.89
| 35 | 0.89 | 0 | 0 | 0 | 0.89 | 0.89 | 0 | 0 | 0 | 0.89 | 0 | 0.89 | 0 | 1.78 | 0.9 | 0
| 36 | 0 | 0 | 0 | 0 | 0.89 | 0 | 0 | 0.89 | 0 | 1.77 | 0 | 3.56 | 1.8 | 1.78 | 0.89 | 0.89
| 37 | 0.9 | 0 | 0 | 0 | 0 | 0.89 | 0 | 1.78 | 0 | 0.89 | 0 | 0 | 0 | 0.9 | 0 | 0
| 38 | 0 | 0 | 0.9 | 0 | 0.89 | 0 | 0 | 0.89 | 1.75 | 0.89 | 2.67 | 1.79 | 0 | 0.89 | 0.9 | 3.58
| 39 | 0 | 0 | 0 | 0.9 | 0 | 0.89 | 1.78 | 0 | 0 | 1.78 | 0.89 | 0.89 | 0.89 | 3.58 | 1.79 | 0
| 40 | 0.9 | 0.89 | 0 | 0 | 0 | 0 | 0 | 0 | 0.89 | 0 | 0.89 | 0 | 0 | 0 | 0.89 | 0
| 41 | 0 | 0.89 | 0 | 0 | 0.9 | 0 | 0 | 0 | 0 | 0.89 | 0 | 0 | 0.89 | 0.9 | 1.78 | 0
| 42 | 0 | 0 | 0 | 0 | 0 | 0 | 0.89 | 1.78 | 0 | 0 | 0.89 | 0 | 0 | 0.89 | 0 | 0
| 43 | 0 | 0 | 0 | 0 | 0 | 0.89 | 0 | 0 | 0 | 1.78 | 1.78 | 0.89 | 0.9 | 0 | 0 | 0
| 44 | 0 | 0 | 0 | 0.9 | 0.89 | 2.68 | 0 | 0 | 0 | 0 | 0.88 | 0.89 | 0 | 2.7 | 0 | 0
| 45 | 0 | 0.89 | 0.89 | 0.9 | 0 | 0.89 | 1.78 | 0 | 0 | 0 | 0 | 0.89 | 0.89 | 0.89 | 0 | 0
| 46 | 0 | 0.89 | 0.9 | 0.89 | 0.9 | 0.89 | 0.89 | 0 | 0 | 0 | 1.77 | 0.89 | 0 | 0.89 | 0 | 1.8
| 47 | 0.89 | 0 | 0 | 0.91 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.9 | 0
| 48 | 0 | 0 | 0 | 0.9 | 0 | 0.89 | 0.89 | 0 | 0.89 | 0 | 0 | 0 | 1.79 | 0.89 | 0 | 1.79
| 49 | 0 | 0.9 | 0 | 0.9 | 2.68 | 0 | 0 | 1.79 | 0 | 0 | 1.78 | 1.79 | 0 | 0.89 | 0.89 | 0.9
| 50 | 0 | 0 | 0.9 | 0 | 1.79 | 0.89 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.89 | 0.89
| 51 | 0 | 0 | 0 | 0 | 0 | 0 | 0.89 | 0.89 | 0 | 0.89 | 0 | 0 | 0 | 0.9 | 0.89 | 0
| 52 | 0 | 0 | 0 | 0 | 0 | 3.59 | 0.89 | 0 | 0 | 0 | 1.77 | 0 | 1.78 | 0 | 0 | 0.89
| 53 | 0 | 0 | 0.9 | 0 | 0.9 | 0.89 | 0 | 0 | 0 | 0.89 | 0.88 | 0.9 | 0 | 0.9 | 1.79 | 0.89
| 54 | 0 | 0.89 | 0 | 0 | 0.9 | 0 | 0 | 0.9 | 1.78 | 0.89 | 0 | 2.68 | 0 | 0.89 | 0.89 | 0.89
| 55 | 0 | 0 | 1.79 | 0.9 | 0 | 0.9 | 0.89 | 0 | 0.89 | 0 | 0.88 | 0 | 0.9 | 0 | 0 | 1.79
| 56 | 0.89 | 0 | 0 | 0 | 0 | 0 | 0.88 | 0 | 0.89 | 0 | 1.78 | 0 | 0.89 | 0 | 0.9 | 1.79
| 57 | 0 | 0.89 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.89 | 0 | 0 | 0 | 0 | 0.9 | 2.68
| 58 | 0.89 | 0 | 0 | 0 | 0 | 0 | 0.89 | 1.78 | 0.89 | 0 | 0 | 0 | 1.8 | 0 | 0 | 0
| 59 | 0 | 0 | 0 | 0 | 0.89 | 1.79 | 0 | 1.78 | 0 | 0 | 0.89 | 1.79 | 0.9 | 0 | 0 | 0.9
| 60 | 0 | 0.89 | 1.81 | 0.9 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1.79 | 0 | 0
| 61 | 0 | 0 | 0.89 | 0.89 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.89 | 0 | 1.78 | 0.89 | 0
| 62 | 0 | 0 | 0 | 0 | 0 | 0.89 | 1.77 | 0 | 0 | 1.79 | 0.89 | 0 | 0 | 0 | 0 | 0
| 63 | 1.79 | 0 | 0 | 0 | 0 | 0.89 | 0 | 0.89 | 0 | 0 | 0.89 | 0 | 0 | 0 | 2.68 | 0
| 64 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.89 | 0 | 0.9 | 0 | 0 | 0 | 0 | 0.89 | 0.89
| 65 | 0 | 0 | 0 | 0 | 0.9 | 0 | 0 | 0.89 | 0 | 0 | 0.89 | 0.9 | 0.9 | 0.89 | 0 | 0
| 66 | 0.9 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.89 | 0.89 | 0 | 0.89 | 1.79 | 0 | 0 | 0
| 67 | 0 | 0 | 0 | 0.9 | 0.89 | 0.89 | 1.79 | 1.78 | 0 | 0.89 | 0 | 1.78 | 0 | 0 | 0 | 0
| 68 | 0 | 0.89 | 0 | 0 | 0.9 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.89 | 0 | 0 | 0.89
| 69 | 1.78 | 0 | 0 | 0 | 2.69 | 0 | 0 | 0 | 0 | 0.89 | 0 | 0 | 1.78 | 0.89 | 2.69 | 1.78
| 70 | 0 | 0.9 | 0 | 0.9 | 0 | 0 | 0.89 | 0 | 0 | 0 | 0.89 | 0 | 0.89 | 1.78 | 1.79 | 1.78
| 71 | 1.79 | 0.89 | 0 | 0 | 0.89 | 0.89 | 0 | 0 | 0 | 0 | 0 | 0 | 0.89 | 4.47 | 6.22 | 6.24
| 72 | 0 | 0 | 0.9 | 0 | 0 | 0.89 | 0 | 0.89 | 0.89 | 0 | 0 | 0 | 0 | 4.47 | 2.67 | 3.57
| 73 | 0 | 0 | 0.89 | 0.9 | 0 | 0 | 0.89 | 0 | 0.89 | 0 | 0.89 | 0 | 0 | 3.56 | 5.35 | 6.24
| 74 | 0 | 0 | 0 | 0 | 0 | 1.79 | 0 | 1.77 | 0 | 3.58 | 0.89 | 0 | 0 | 4.48 | 7.14 | 7.13
| 75 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2.69 | 0 | 0.89 | 1.79 | 11.58 | 12.54 | 6.25
| 76 | 0.89 | 0 | 0 | 1.79 | 0 | 0 | 0.89 | 1.78 | 0 | 4.45 | 4.44 | 2.67 | 0.9 | 8.05 | 11.64 | 15.19
| 77 | 0 | 0 | 0.89 | 0 | 0.9 | 1.78 | 0.89 | 0 | 1.78 | 8.9 | 0.89 | 4.4 | 9.82 | 10.72 | 13.38 | 8.02
| 78 | 0 | 0.9 | 0 | 0 | 1.78 | 0.89 | 0 | 0 | 0.89 | 7.98 | 20.39 | 8.02 | 7.17 | 33.23 | 12.44 | 10.73
| 79 | 0.89 | 0.9 | 0 | 0.89 | 0.9 | 0 | 1.78 | 0 | 0.9 | 14.19 | 8 | 11.5 | 15.21 | 43.77 | 17.88 | 15.2
| 80 | 0 | 0.89 | 0 | 0 | 1.79 | 0 | 0 | 0.89 | 1.78 | 14.22 | 12.34 | 15.16 | 23.21 | 147.24 | 71.61 | 70.61
| 81 | 0 | 0.89 | 0 | 0 | 3.58 | 0 | 0 | 0 | 0 | 22.13 | 18.71 | 17.84 | 13.24 | 269.62 | 152.35 | 139.47
| 82 | 0.89 | 0 | 0 | 0 | 5.37 | 0.89 | 0 | 1.78 | 8.02 | 23.18 | 23.06 | 21.47 | 17.02 | 454.57 | 299.91 | 279.95
| 83 | 0 | 0 | 0 | 0 | 8.15 | 3.57 | 1.78 | 1.78 | 4.43 | 24.94 | 28.42 | 29.44 | 24.06 | 805.27 | 657.08 | 624.89
| 84 | 0 | 0 | 0 | 0 | 16.11 | 1.77 | 2.67 | 0.88 | 7.99 | 29.33 | 27.5 | 32.19 | 18.74 | 930.11 | 772.36 | 674.08
| 85 | 0 | 0 | 0.89 | 0 | 17.83 | 16.03 | 5.35 | 8.87 | 18.51 | 59.9 | 35.68 | 39.25 | 28.7 | 1062.74 | 987.55 | 912.01
| 86 | 0 | 0 | 0.89 | 0 | 27.6 | 14.21 | 20.51 | 14.21 | 22.29 | 75.18 | 48.94 | 40.34 | 40.4 | 1195.73 | 1127.59 | 1116.78
| 87 | 0 | 0 | 0.9 | 0 | 35.8 | 25.91 | 25.12 | 29.4 | 28.77 | 162.29 | 112.41 | 101.04 | 102.58 | 1330.28 | 1271.32 | 1169.54
| 88 | 0 | 0.9 | 0.89 | 0.89 | 38.38 | 36.11 | 23.21 | 29.31 | 34.69 | 384.42 | 316.17 | 307.21 | 336.57 | 1390.22 | 1393.72 | 1309.24
| 89 | 0 | 0 | 0 | 0 | 37.6 | 38.38 | 39.27 | 40.86 | 34.87 | 611.07 | 492.82 | 378.68 | 374.98 | 1436.18 | 1469.74 | 1352.57
| 90 | 0 | 0.9 | 0 | 0 | 38.44 | 42.16 | 54.16 | 48.11 | 42.8 | 958.96 | 839.96 | 768.69 | 732.53 | 1512.62 | 1553.53 | 1549.87
| 91 | 0 | 0.89 | 0.89 | 0 | 59.26 | 36.34 | 63.55 | 38.17 | 57.93 | 1202.76 | 1049.1 | 1091.98 | 1040.57 | 1587.76 | 1594.76 | 1468.18
| 92 | 0 | 0.89 | 1.8 | 0 | 90.41 | 45.66 | 51.44 | 58.91 | 86.51 | 1480.77 | 1271.83 | 1306.66 | 1348.17 | 1614.58 | 1613.5 | 1570.08
| 93 | 1.79 | 0 | 0 | 1.79 | 162.21 | 70.53 | 47.88 | 54.39 | 152.88 | 1638.4399 | 1476.61 | 1474.22 | 1557.84 | 1711.36 | 1694.57 | 1605.93
| 94 | 0.89 | 0 | 0 | 0 | 314.61 | 69.71 | 70.05 | 74.79 | 264.92 | 1654.43 | 1658.1999 | 1645.32 | 1640.62 | 1831.4399 | 1830.43 | 1688.32
| 95 | 1.78 | 0.9 | 1.79 | 2.69 | 548.56 | 156.43 | 118.61 | 251.88 | 476.74 | 1796.63 | 1671.29 | 1815.18 | 1730.86 | 1822.13 | 1834.13 | 1724.8001
| 96 | 4.46 | 5.35 | 6.33 | 1.79 | 767.16 | 189.04 | 239.72 | 606.78 | 957.84 | 1771.25 | 1728.6999 | 1802.0501 | 1809.01 | 1888.04 | 1834.1 | 1679.1899
| 97 | 5.39 | 12.55 | 13.45 | 8.94 | 1092.4499 | 357.5 | 432.87 | 1403.9499 | 1201.02 | 1913.0501 | 1864.29 | 1935 | 1974.75 | 1887.83 | 1901.79 | 1741.39
| 98 | 9.86 | 11.65 | 17.92 | 21.54 | 1513.93 | 908.71 | 607.71 | 1670.01 | 1454.97 | 1945.38 | 1832.53 | 1936.73 | 1971.26 | 1966.5501 | 1979.0501 | 1811.33
| 99 | 31.31 | 17 | 26.95 | 27.79 | 1878.42 | 1123.67 | 878.14 | 1953.6999 | 1598.73 | 2028.46 | 1907.61 | 1961.75 | 1934.03 | 2010.51 | 1939.12 | 1865.5501
| 100 | 35.89 | 48.33 | 56.42 | 43.08 | 1963.5601 | 1398.29 | 1502.58 | 1933.79 | 1785.02 | 2026.89 | 1929.03 | 2056.05 | 2101.29 | 2021.63 | 2059.9199 | 1860.18
| 101 | 57.64 | 57.41 | 59.18 | 57.19 | 2057 | 1525.68 | 1706.1899 | 2047.51 | 1868.14 | 2049.79 | 1901.67 | 2042.3101 | 2156.1101 | 2075.8899 | 2085.9199 | 1904.28
| 102 | 69.64 | 69.62 | 66.56 | 64.47 | 2062.8501 | 1845.62 | 1882.18 | 2067.9399 | 1841.78 | 2045.43 | 1948.37 | 2101.04 | 2005.57 | 2031.9 | 2011.6 | 1779.93
| 103 | 78.45 | 73.44 | 77.79 | 93.46 | 2115.45 | 1902.9 | 1974.76 | 2077.8501 | 1846.66 | 2118.3301 | 1913.35 | 2037.8001 | 2107.6499 | 2090.1899 | 2080.4199 | 1855.5
| 104 | 89.84 | 78.03 | 89.95 | 108.69 | 2116.1201 | 1812.5601 | 2090.3301 | 2057.4099 | 1823.71 | 2149.3601 | 1975.32 | 2201.05 | 2042.85 | 2038.46 | 2079.8601 | 1930.4499
| 105 | 97.07 | 92.38 | 102.46 | 89.7 | 2097.8701 | 1861.74 | 2057.4099 | 2071.1001 | 1975.6999 | 2199.02 | 1982.86 | 2205.53 | 2132.8601 | 2008.68 | 2071.8501 | 1986.73
| 106 | 113.89 | 108.54 | 97.21 | 117.37 | 2081.1101 | 1914.66 | 2215.6001 | 1994.32 | 1940.35 | 2177.27 | 1912.72 | 2110.3999 | 2078.46 | 2089.55 | 2020.18 | 1840.64
| 107 | 102.67 | 116.4 | 135.68 | 115.47 | 2133.27 | 1974.8001 | 2190.6799 | 2124.8201 | 1848.34 | 2089.76 | 1983.25 | 2047.41 | 2112.46 | 1995.13 | 2061.3701 | 1896.8101
| 108 | 118.79 | 146.51 | 135.85 | 203.91 | 2167.6699 | 1955.49 | 2105.26 | 2045.3101 | 1876.1999 | 2067.3101 | 1886.04 | 2024.84 | 2045.59 | 1990.66 | 1998.67 | 1848.16
| 109 | 226.22 | 246.94 | 184.82 | 373.52 | 2075.55 | 1933.74 | 2028.02 | 2099.8601 | 1975.6899 | 2163.72 | 1897.98 | 2045.72 | 2028.9499 | 2013.09 | 1914.21 | 1877.86
| 110 | 448.72 | 485.44 | 360.89 | 309.81 | 2132.54 | 1975.99 | 2065.1101 | 2014.3101 | 1835.78 | 2081.53 | 1940.01 | 2086.5601 | 2085.52 | 1956.4 | 2002.86 | 1780.71
| 111 | 436.01 | 614.8 | 570.13 | 431.87 | 2007.43 | 1884.42 | 2068.04 | 1962.88 | 1774.53 | 2036.25 | 1859.47 | 2056.46 | 1951.32 | 1858.12 | 1958.1999 | 1785.96
| 112 | 647.56 | 812.98 | 725.02 | 619.9 | 1994.65 | 1861.29 | 2017.6899 | 1985.4399 | 1830.54 | 1938.62 | 1814.07 | 1973.9399 | 2000.97 | 1905.4499 | 1857.71 | 1757.98
| 113 | 707.24 | 889.3 | 868.09 | 785.49 | 1976.1 | 1746.6 | 1997.85 | 1829.59 | 1701.96 | 1953.4 | 1685.17 | 1874.02 | 1848.72 | 1743.57 | 1827.42 | 1711.73
| 114 | 942.43 | 1010.23 | 915.22 | 1003.85 | 1883.04 | 1712.74 | 1850.83 | 1865.72 | 1715.79 | 1905.35 | 1633.52 | 1895.51 | 1875.63 | 1713.1 | 1754.9399 | 1627.1999
| 115 | 1248.73 | 1167.03 | 1148.14 | 1198.11 | 1931.96 | 1654.12 | 1828.03 | 1817.03 | 1582.86 | 1841.1899 | 1546.76 | 1835.36 | 1754.32 | 1692.5601 | 1762.4399 | 1496.84
| 116 | 1234.86 | 1345.5501 | 1309.64 | 1373.23 | 1837.4 | 1598.98 | 1739.1899 | 1655.43 | 1575.01 | 1748.64 | 1609.35 | 1762.5501 | 1727.01 | 1645.4499 | 1652.48 | 1515.9499
| 117 | 1293.41 | 1373.39 | 1396.54 | 1428.74 | 1704.28 | 1476.24 | 1749.66 | 1579.9399 | 1540.99 | 1703.9499 | 1437.63 | 1587.08 | 1695.21 | 1541.5601 | 1582.9499 | 1406.63
| 118 | 1310.62 | 1251.99 | 1311.62 | 1364.61 | 1670.91 | 1482.52 | 1662.38 | 1557.08 | 1523.0601 | 1492.26 | 1408.8001 | 1539.67 | 1537.62 | 1493.93 | 1490.49 | 1379.5601
| 119 | 1331.5501 | 1304.86 | 1375.14 | 1319.62 | 1528.17 | 1425.43 | 1517.5 | 1479.29 | 1354.07 | 1462.83 | 1377.46 | 1398.93 | 1456.12 | 1373.0601 | 1356.08 | 1299.49
| 120 | 1285.78 | 1213.52 | 1332.57 | 1284.49 | 1538.26 | 1349.62 | 1402.1999 | 1457.49 | 1323.57 | 1442.91 | 1275.63 | 1396.6899 | 1411.58 | 1331.75 | 1361.22 | 1234.64
| 121 | 1241.26 | 1205.62 | 1263.01 | 1181.53 | 1317.48 | 1253.41 | 1318.1 | 1305.84 | 1365.5601 | 1327.86 | 1288.89 | 1326.86 | 1379.64 | 1211.23 | 1290.33 | 1199.99
| 122 | 1254.68 | 1111.47 | 1201.14 | 1174.03 | 1228.83 | 1144.28 | 1174.36 | 1219.36 | 1192.1 | 1178.62 | 1130.62 | 1277.28 | 1234.02 | 1168.65 | 1200 | 1101.79
| 123 | 1137.03 | 1042.86 | 1161.3001 | 1073.8001 | 1186.0601 | 1125.97 | 1211.9399 | 1192.64 | 1176.87 | 1174.72 | 1112.09 | 1121.3001 | 1157.68 | 1072.96 | 1165.42 | 1012.9
| 124 | 1059.42 | 1065.17 | 1062.9 | 1044.04 | 1148.49 | 1005.02 | 1093.39 | 1054.98 | 1084.24 | 1076.15 | 924.59 | 1080.86 | 1030.9399 | 1021.5 | 1021.46 | 958.79
| 125 | 982.39 | 957.24 | 1049.14 | 1046.9499 | 1099.6999 | 994.36 | 999.35 | 1061.36 | 1016.26 | 999.87 | 999.98 | 1029.1 | 981.79 | 989.1 | 985.04 | 941.84
| 126 | 926.31 | 901.65 | 941.23 | 928.32 | 947.85 | 885.16 | 927.93 | 968.19 | 902.67 | 921.82 | 884.7 | 947.76 | 925.02 | 890.12 | 884.88 | 804.47
| 127 | 934.52 | 880.11 | 865.67 | 858.14 | 971.48 | 859.83 | 912.9 | 865.67 | 846.97 | 931.94 | 832.59 | 913.87 | 858.51 | 828.62 | 872.27 | 749.77
| 128 | 832.2 | 803.61 | 829.32 | 793.29 | 871.15 | 796.41 | 792.4 | 859.89 | 764.22 | 895.79 | 767.71 | 827.45 | 847 | 764.85 | 803.26 | 761.07
| 129 | 798.8 | 758.27 | 797.78 | 807.23 | 833.21 | 802.48 | 755.18 | 793.18 | 739.08 | 791.01 | 766.55 | 773.22 | 804.61 | 697.28 | 798.64 | 736.01
| 130 | 727.25 | 673.46 | 772.22 | 727.93 | 752.37 | 718 | 760.11 | 742.46 | 741.33 | 785.51 | 689.7 | 738.02 | 754.42 | 688.6 | 756.33 | 674.94
| 131 | 674.32 | 663.82 | 800.26 | 676.56 | 730.51 | 664.52 | 664.06 | 694.3 | 637.82 | 700.02 | 666.12 | 710.33 | 686.35 | 640.17 | 676.25 | 653.73
| 132 | 666.65 | 666.69 | 644.91 | 663.19 | 686.86 | 621.74 | 685.73 | 651.43 | 611.46 | 652.02 | 620.04 | 658.87 | 681.38 | 649.35 | 649.25 | 626.64
| 133 | 607.53 | 582.54 | 612.1 | 621.35 | 679.28 | 591.1 | 580.96 | 620.84 | 630.55 | 660.55 | 610.75 | 620.35 | 670.83 | 599.92 | 677.08 | 608.86
| 134 | 667.57 | 715.17 | 677.85 | 713.77 | 643.21 | 621.8 | 608.85 | 625.56 | 627.12 | 751.83 | 669.79 | 720.39 | 747.62 | 802.66 | 823.56 | 850.9
| 135 | 900.68 | 865.3 | 935.42 | 1071.84 | 791.13 | 830.41 | 857.92 | 867.12 | 823.91 | 1042.07 | 1003.67 | 1046.28 | 1098.9 | 1170.62 | 1300.38 | 1223.66
| 136 | 1012.44 | 1043.68 | 1084.6999 | 1538.3101 | 1218.39 | 1158.71 | 1248.26 | 1279.11 | 1206.38 | 1478.15 | 1402.99 | 1444.73 | 1519.29 | 1587.28 | 1678.8101 | 1506.72
| 137 | 579.49 | 599.92 | 576.81 | 1971.85 | 1634.17 | 1633 | 1676.52 | 1592.33 | 1531.47 | 1892.18 | 1641.14 | 1797.98 | 1843.1999 | 1791.4 | 1975.71 | 1787.13
| 138 | 307.07 | 292.48 | 323.03 | 1954.01 | 1868.79 | 1491.65 | 1621.26 | 1861.58 | 1474.3101 | 1971.86 | 1807.43 | 1998.71 | 1919.92 | 1814.46 | 1834.9 | 1689.9399
| 139 | 160.71 | 158.6 | 180.47 | 1257.6899 | 1152.8001 | 780.23 | 798.03 | 1691.96 | 763.76 | 1373.17 | 1158.28 | 1195.53 | 1143.8001 | 1306.13 | 1167.77 | 1107.49
| 140 | 83.09 | 67.7 | 70.02 | 493.22 | 477.35 | 407.73 | 394.25 | 788.04 | 369.76 | 599.38 | 487.68 | 511.35 | 469 | 633.96 | 544.18 | 506.67
| 141 | 40.43 | 27.72 | 33.16 | 178.48 | 211.16 | 189.15 | 155.75 | 359.21 | 148.42 | 230.11 | 203.44 | 224.09 | 203.04 | 231 | 228.81 | 170.99
| 142 | 18.84 | 15.29 | 9.89 | 80.77 | 76.08 | 65.12 | 73.79 | 112.56 | 60.5 | 99.19 | 79.96 | 83.07 | 79.02 | 73.3 | 63.37 | 72.57
| 143 | 10.75 | 5.37 | 7.15 | 19.73 | 35.81 | 31.93 | 32.96 | 38.48 | 28.52 | 33.06 | 46.33 | 29.49 | 28.59 | 30.48 | 34.03 | 30.35
| 144 | 6.26 | 4.48 | 1.8 | 9.76 | 21.41 | 5.34 | 8.9 | 26.69 | 13.36 | 16.76 | 13.36 | 16.96 | 9.86 | 19.63 | 16.17 | 13.35
| 145 | 5.38 | 5.38 | 4.47 | 8.07 | 13.45 | 5.35 | 9.02 | 11.56 | 9.8 | 7.1 | 5.35 | 8.02 | 6.27 | 7.18 | 9.87 | 3.56
| 146 | 0.89 | 4.48 | 5.28 | 8.93 | 3.56 | 4.45 | 4.44 | 10.7 | 3.56 | 2.66 | 5.34 | 8.04 | 5.31 | 6.24 | 7.14 | 3.58
| 147 | 0.9 | 0.9 | 1.79 | 5.37 | 10.67 | 2.67 | 2.68 | 2.68 | 5.32 | 4.43 | 5.34 | 6.22 | 6.24 | 3.57 | 1.78 | 3.57
| 148 | 2.68 | 1.78 | 2.68 | 3.59 | 6.24 | 8.01 | 4.51 | 2.67 | 3.55 | 1.77 | 3.55 | 8.04 | 5.38 | 5.37 | 2.67 | 2.68
| 149 | 2.68 | 2.69 | 3.6 | 2.68 | 1.78 | 3.56 | 4.46 | 0.89 | 4.46 | 0.89 | 1.78 | 1.8 | 3.58 | 0 | 3.57 | 1.79
| 150 | 2.68 | 0 | 3.58 | 4.48 | 1.78 | 1.78 | 1.78 | 3.57 | 1.78 | 2.67 | 2.67 | 1.79 | 0.9 | 0 | 2.67 | 1.78
| 151 | 0.9 | 0 | 0.89 | 2.68 | 2.68 | 1.76 | 0.89 | 1.78 | 0 | 0 | 2.65 | 1.8 | 2.69 | 1.79 | 2.67 | 3.58
| 152 | 0.89 | 1.83 | 0.89 | 2.69 | 1.78 | 0.89 | 0.89 | 0.89 | 0.89 | 0.89 | 0.89 | 1.78 | 2.68 | 2.68 | 0.89 | 1.79
| 153 | 0 | 0 | 1.8 | 1.79 | 1.78 | 0 | 0 | 0.89 | 0.89 | 2.66 | 2.66 | 1.79 | 2.69 | 2.68 | 0 | 0.89
| 154 | 0 | 0 | 1.78 | 0.9 | 1.79 | 0.89 | 0 | 2.67 | 0 | 0.89 | 0.89 | 2.67 | 3.57 | 0 | 0 | 0.9
| 155 | 0 | 2.68 | 0 | 1.79 | 1.78 | 0.89 | 1.78 | 0.9 | 2.67 | 4.43 | 0 | 0 | 0.89 | 2.69 | 1.78 | 0
| 156 | 0 | 0.89 | 0 | 0.89 | 2.68 | 1.78 | 0.89 | 0 | 0 | 1.78 | 1.78 | 0 | 2.68 | 0.9 | 2.67 | 0.89
| 157 | 0.89 | 0 | 0 | 1.79 | 3.56 | 0.89 | 0 | 2.66 | 1.79 | 2.66 | 0 | 0 | 1.79 | 0.89 | 0 | 0.89
| 158 | 0 | 0 | 0.89 | 0 | 0 | 1.78 | 0.89 | 0.89 | 2.66 | 0 | 0 | 1.79 | 0.9 | 0.89 | 0.89 | 0.89
| 159 | 1.79 | 0.89 | 0.89 | 0.89 | 0 | 2.67 | 0.89 | 0.88 | 0.89 | 0 | 0.89 | 0 | 0.9 | 1.78 | 0.9 | 1.78
| 160 | 1.79 | 0 | 2.68 | 0 | 2.66 | 2.66 | 0 | 0 | 1.78 | 0.89 | 0.89 | 1.79 | 0.89 | 0 | 0.89 | 0.89
| 161 | 0 | 0.9 | 0.89 | 0 | 0 | 0 | 0 | 2.68 | 0.89 | 0 | 0.89 | 0.89 | 0 | 0.9 | 0.89 | 1.77
| 162 | 0 | 1.78 | 0.89 | 0.9 | 0 | 0 | 1.78 | 0 | 1.78 | 0 | 0.89 | 1.79 | 2.68 | 0.89 | 0 | 0
| 163 | 0 | 1.79 | 1.79 | 0.89 | 0 | 2.67 | 0.89 | 0.89 | 0 | 1.78 | 1.78 | 2.63 | 0 | 1.79 | 0.89 | 1.79
| 164 | 0.89 | 2.69 | 0 | 0 | 0.89 | 0.88 | 1.78 | 0.89 | 0 | 0 | 2.67 | 0.89 | 0.91 | 0 | 0 | 1.78
| 165 | 0.89 | 0 | 0.89 | 0.9 | 0 | 0.89 | 0 | 0 | 0 | 0.89 | 0 | 0 | 0 | 0 | 0 | 1.78
| 166 | 1.79 | 0 | 0 | 0.89 | 0 | 0.89 | 0.89 | 2.69 | 0.89 | 1.76 | 0 | 0.89 | 0 | 0.89 | 0 | 1.79
| 167 | 1.78 | 0.89 | 0 | 0.89 | 1.78 | 0 | 0 | 0.89 | 0.88 | 0.89 | 0 | 0 | 0 | 0 | 0 | 1.78
| 168 | 0 | 0 | 0 | 0 | 0.89 | 0 | 0 | 0.89 | 0.89 | 0.89 | 0.89 | 0.9 | 2.69 | 0.9 | 1.78 | 0
| 169 | 0.9 | 0.9 | 0 | 0.89 | 0 | 0.89 | 0.89 | 0 | 0 | 0 | 0 | 3.57 | 0.89 | 0 | 0.89 | 0.89
| 170 | 0 | 0 | 0 | 0.89 | 0.89 | 0 | 0 | 0 | 0 | 0 | 0.89 | 0 | 0 | 0.9 | 0 | 0.89
| 171 | 0 | 0 | 0 | 0 | 0 | 0 | 0.89 | 0 | 1.78 | 0 | 3.56 | 1.8 | 1.78 | 0.89 | 0.89 | 0
| 172 | 0 | 0 | 0 | 0.89 | 0.89 | 0 | 1.78 | 0 | 0.89 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 173 | 0 | 0.9 | 0 | 0 | 0 | 0 | 0.89 | 1.75 | 0 | 2.67 | 1.79 | 0 | 0.89 | 0.9 | 3.58 | 0.89
| 174 | 0 | 0 | 0.9 | 0 | 0.89 | 1.78 | 0 | 0 | 0 | 0.89 | 0.89 | 0.89 | 0 | 1.79 | 0 | 1.78
| 175 | 0.89 | 0 | 0 | 0.9 | 0 | 0 | 0 | 0.89 | 0.89 | 0.89 | 0 | 0 | 0 | 0.89 | 0 | 0.89
| 176 | 0.89 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.89 | 0 | 0 | 0.89 | 0.89 | 1.78 | 0 | 0
| 177 | 0 | 0 | 0 | 0 | 0 | 0.89 | 1.78 | 0 | 0.88 | 0.89 | 0 | 0 | 0.89 | 0 | 0 | 0
| 178 | 0 | 0 | 0 | 0.89 | 0.89 | 0 | 0 | 0 | 0 | 1.78 | 0.89 | 0.9 | 0.89 | 0 | 0 | 0.89
| 179 | 0 | 0 | 0.9 | 0 | 2.68 | 0 | 0 | 0 | 0 | 0.88 | 0.89 | 0 | 0 | 0 | 0 | 0.9
| 180 | 0.89 | 0.89 | 0.9 | 0.9 | 0.89 | 1.78 | 0 | 0 | 0.89 | 0 | 0.89 | 0.89 | 1.79 | 0 | 0 | 0
| 181 | 0.89 | 0.9 | 0.89 | 0 | 0.89 | 0.89 | 0 | 0 | 0.89 | 1.77 | 0.89 | 0 | 0.9 | 0 | 1.8 | 0
| 182 | 0 | 0 | 0.91 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.89 | 0.9 | 0 | 0.89
| 183 | 0 | 0 | 0.9 | 2.68 | 0.89 | 0.89 | 0 | 0.89 | 0.89 | 0 | 0 | 1.79 | 0.89 | 0 | 1.79 | 0.89
| 184 | 0.9 | 0 | 0.9 | 1.79 | 0 | 0 | 1.79 | 0 | 0.89 | 1.78 | 1.79 | 0 | 0.89 | 0.89 | 0.9 | 0
| 185 | 0 | 0.9 | 0 | 0 | 0.89 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.89 | 0.89 | 0
| 186 | 0 | 0 | 0 | 0 | 0 | 0.89 | 0.89 | 0 | 0 | 0 | 0 | 0 | 0 | 0.89 | 0 | 0
| 187 | 0 | 0 | 0 | 0.9 | 3.59 | 0.89 | 0 | 0 | 0 | 1.77 | 0 | 1.78 | 0.91 | 0 | 0.89 | 0.89
| 188 | 0 | 0.9 | 0 | 0.9 | 0.89 | 0 | 0 | 0 | 1.78 | 0.88 | 0.9 | 0 | 0 | 1.79 | 0.89 | 0
| 189 | 0.89 | 0 | 0 | 0 | 0 | 0 | 0.9 | 1.78 | 0 | 0 | 2.68 | 0 | 0.9 | 0.89 | 0.89 | 0.89
| 190 | 0 | 1.79 | 0.9 | 0 | 0.9 | 0.89 | 0 | 0.89 | 0.89 | 0.88 | 0 | 0.9 | 0.89 | 0 | 1.79 | 0.89
| 191 | 0 | 0 | 0 | 0 | 0 | 0.88 | 0 | 0.89 | 0 | 1.78 | 0 | 0.89 | 0.89 | 0.9 | 1.79 | 0.89
| 192 | 0.89 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.9 | 2.68 | 1.79
| 193 | 0 | 0 | 0 | 0.89 | 0 | 0.89 | 1.78 | 0.89 | 0 | 0 | 0 | 1.8 | 0.89 | 0 | 0 | 0
| 194 | 0 | 0 | 0 | 0 | 1.79 | 0 | 1.78 | 0 | 0.89 | 0.89 | 1.79 | 0.9 | 0 | 0 | 0.9 | 0
| 195 | 0.89 | 1.81 | 0.9 | 0 | 0 | 0 | 0 | 0 | 0.89 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Scantime | AvgSiz_for_each_CCNC_Scan
| 1 | 1.2678 | 1.2695 | 1.2752 | 1.2812 | 1.6714 | 2.1567 | 2.2322 | 2.2538 | 1.951 | 3.4022 | 3.5565 | 3.6097 | 3.652 | 4.3736 | 4.7171 | 4.7109
| 2 | 1.2819 | 1.2779 | 1.2829 | 1.2767 | 1.6363 | 2.0449 | 2.1246 | 2.131 | 1.8218 | 3.2345 | 3.3713 | 3.3937 | 3.4389 | 4.3427 | 4.587 | 4.6571
| 3 | 1.2378 | 1.2297 | 1.2853 | 1.2923 | 1.577 | 2.0244 | 2.2129 | 2.1928 | 1.7264 | 3.2126 | 3.3987 | 3.436 | 3.4764 | 4.3005 | 4.5482 | 4.666
| 4 | 1.2097 | 1.2688 | 1.3367 | 1.2356 | 1.5142 | 2.1556 | 2.3224 | 2.2511 | 1.7503 | 3.4144 | 3.6155 | 3.6126 | 3.7201 | 4.3035 | 4.6972 | 4.7882
| 5 | 1.2936 | 1.3528 | 1.3669 | 1.3311 | 1.5669 | 2.1691 | 2.3322 | 2.3707 | 1.7921 | 3.528 | 3.7654 | 3.8031 | 3.756 | 4.5296 | 4.8269 | 4.9676
| 6 | 1.3487 | 1.4821 | 1.4118 | 1.4773 | 1.6759 | 2.2559 | 2.3185 | 2.4127 | 1.8244 | 3.5748 | 3.9392 | 3.8944 | 4.0995 | 4.5672 | 5.061 | 5.0246
| 7 | 1.6786 | 1.4062 | 1.375 | 1.4688 | 1.6764 | 2.3875 | 2.4028 | 2.5203 | 2.0291 | 3.9762 | 4.0068 | 4.0385 | 4.2045 | 4.8357 | 5 | 5.1053
| 8 | 1.4821 | 1.5 | 1.375 | 1.75 | 1.7273 | 2.2917 | 2.4167 | 2.6 | 1.9667 | 4.2794 | 4.2237 | 4.3833 | 4.25 | 5.0833 | 5.2955 | 5.1389
| 9 | 1.75 | 1.5833 | 1.5833 | 1.55 | 1.8864 | 2.2833 | 2.75 | 2.55 | 1.9423 | 4.25 | 4.125 | 4.25 | 4.3611 | 4.8125 | 5.4375 | 5.1136
| 10 | 1.75 | 2.25 | 1.275 | 1.375 | 1.75 | 2.75 | 2.75 | 2.65 | 2.2917 | 4.1 | 4.4167 | 4.3333 | 4.4722 | 5.0682 | 5.3929 | 5.4375
| 11 | 1.25 | 1.25 | 1.25 | 1.5 | 1.8 | 2.625 | 2.4167 | 2.4167 | 1.9167 | 4.25 | 4.75 | 4.25 | 4.3929 | 4.9167 | 5.125 | 5.75
| 12 | 1.75 | 1.5833 | 1.5 | 1.75 | 1.8333 | 2.8929 | 2.5833 | 2.95 | 2.75 | 4.375 | 4 | 4.25 | 4.4722 | 5.5 | 5.25 | 5.4167
| 13 | 1.5833 | 1.5833 | 1.5833 | 1.625 | 2.25 | 2.5 | 2.875 | 2.65 | 1.75 | 4.75 | 5.25 | 4.25 | 4 | 4.25 | 0 | 5.75
| 14 | 1.5 | 1.4167 | 0 | 1.625 | 1.9167 | 3 | 2.5 | 2.5 | 2.625 | 4.5 | 4.9167 | 4.75 | 4 | 5.5 | 0 | 5.5833
| 15 | 1.2917 | 2.25 | 0 | 1.75 | 2.05 | 2.4167 | 3.5 | 2.75 | 2 | 4.25 | 0 | 4.4167 | 4.75 | 5.125 | 5 | 5.25
| 16 | 1.75 | 1.75 | 1.3125 | 1.75 | 2.75 | 3 | 2.25 | 2.75 | 2.25 | 4.4167 | 4.25 | 4.75 | 4 | 5.25 | 5.4167 | 5.75
| 17 | 0 | 0 | 0 | 1.75 | 1.75 | 2.25 | 0 | 0 | 1.75 | 4.5 | 4.75 | 4.4167 | 4.75 | 5.25 | 5.25 | 0
| 18 | 0 | 0 | 0 | 1.75 | 2.25 | 2.25 | 2.75 | 0 | 2.25 | 4.75 | 4.25 | 5.25 | 4.75 | 5.25 | 0 | 0
| 19 | 2 | 0 | 1.5833 | 0 | 2.25 | 2.5 | 2.75 | 2.75 | 1.75 | 4.0833 | 4.35 | 0 | 0 | 5.75 | 5.25 | 5.75
| 20 | 1.75 | 0 | 1.25 | 0 | 2.25 | 2.9167 | 2.5 | 3.25 | 0 | 4.5 | 5 | 4.75 | 0 | 5.25 | 5.25 | 5.4167
| 21 | 0 | 1.75 | 0 | 0 | 2.75 | 2.75 | 1.75 | 0 | 2.9167 | 0 | 4.75 | 0 | 0 | 5.25 | 5.75 | 0
| 22 | 1.3125 | 0 | 0 | 1.75 | 2 | 0 | 3 | 2.75 | 3.75 | 0 | 0 | 0 | 4.5 | 0 | 5.75 | 5.25
| 23 | 1.75 | 1.25 | 1.25 | 1.75 | 0 | 0 | 2.75 | 3.25 | 2.25 | 0 | 0 | 4.75 | 0 | 5.25 | 5.75 | 5.25
| 24 | 1.75 | 1.5 | 0 | 1.4167 | 2.75 | 2.75 | 2.4167 | 0 | 0 | 0 | 4.25 | 3.25 | 4.75 | 0 | 0 | 5.25
| 25 | 1.75 | 0 | 0.875 | 1.75 | 0 | 0 | 0 | 0 | 2.4167 | 0 | 0 | 5.25 | 4.25 | 5.5 | 5.75 | 5.25
| 26 | 1.75 | 0 | 1.75 | 1.25 | 0 | 0 | 0 | 3 | 0 | 4.25 | 0 | 4.25 | 4 | 0 | 4.75 | 0
| 27 | 1.5833 | 0 | 1.5 | 1.75 | 2.75 | 0 | 2.9167 | 3.25 | 2.75 | 4.25 | 4.75 | 4.25 | 4.5833 | 5.25 | 5 | 5.25
| 28 | 1.75 | 0.875 | 1.75 | 0 | 1.75 | 2.75 | 2.25 | 3 | 2.25 | 4.5 | 0 | 4.4167 | 4.75 | 5.25 | 0 | 0
| 29 | 1.3125 | 1.25 | 0 | 1.75 | 0 | 0 | 2.25 | 0 | 0 | 0 | 4.75 | 0 | 0 | 5.75 | 0 | 0
| 30 | 1.75 | 1.75 | 0 | 0 | 1.75 | 0 | 2.75 | 3.25 | 2.5833 | 0 | 4.75 | 0 | 5.25 | 5.5 | 5.75 | 0
| 31 | 1.25 | 1.5 | 1.75 | 0 | 2.25 | 3 | 0 | 0 | 2.75 | 3.75 | 4.75 | 0 | 0 | 5.25 | 0 | 0
| 32 | 0 | 0 | 0 | 0 | 2.25 | 2.75 | 0 | 0 | 3.25 | 0 | 4.75 | 4.75 | 4.25 | 5.75 | 5.25 | 5.75
| 33 | 1.75 | 1.75 | 1.75 | 0 | 0 | 0 | 2.25 | 2.75 | 0 | 0 | 0 | 0 | 4.75 | 5.25 | 0 | 5.75
| 34 | 1.75 | 0 | 0 | 0 | 2.75 | 2.75 | 0 | 0 | 0 | 4.25 | 0 | 3.75 | 0 | 4.75 | 5.25 | 0
| 35 | 0 | 0 | 0 | 0 | 2.25 | 0 | 0 | 2.75 | 0 | 4.5 | 0 | 4.5 | 4.25 | 5.25 | 5.75 | 5.75
| 36 | 1.25 | 0 | 0 | 0 | 0 | 2.75 | 0 | 3 | 0 | 4.25 | 0 | 0 | 0 | 4.25 | 0 | 0
| 37 | 0 | 0 | 1.75 | 0 | 1.75 | 0 | 0 | 3.25 | 2.75 | 3.75 | 4.4167 | 4.75 | 0 | 5.25 | 5.75 | 5.75
| 38 | 0 | 0 | 0 | 1.25 | 0 | 2.75 | 3.75 | 0 | 0 | 4.25 | 4.25 | 4.75 | 3.75 | 5.125 | 5.25 | 0
| 39 | 1.25 | 1.25 | 0 | 0 | 0 | 0 | 0 | 0 | 2.25 | 0 | 4.75 | 0 | 0 | 0 | 5.75 | 0
| 40 | 0 | 1.75 | 0 | 0 | 2.75 | 0 | 0 | 0 | 0 | 4.75 | 0 | 0 | 3.25 | 5.25 | 5 | 0
| 41 | 0 | 0 | 0 | 0 | 0 | 0 | 3.75 | 2.75 | 0 | 0 | 5.25 | 0 | 0 | 6.25 | 0 | 0
| 42 | 0 | 0 | 0 | 0 | 0 | 2.75 | 0 | 0 | 0 | 4.25 | 4.75 | 4.25 | 4.25 | 0 | 0 | 0
| 43 | 0 | 0 | 0 | 0.875 | 1.75 | 3.25 | 0 | 0 | 0 | 0 | 4.75 | 4.25 | 0 | 5.0833 | 0 | 0
| 44 | 0 | 1.25 | 1.75 | 0.875 | 0 | 2.75 | 3 | 0 | 0 | 0 | 0 | 5.25 | 5.25 | 5.25 | 0 | 0
| 45 | 0 | 1.75 | 1.25 | 1.75 | 2.25 | 3.25 | 3.75 | 0 | 0 | 0 | 5 | 4.75 | 0 | 5.25 | 0 | 6
| 46 | 2.75 | 0 | 0 | 1.25 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4.75 | 0
| 47 | 0 | 0 | 0 | 1.25 | 0 | 2.75 | 2.25 | 0 | 1.75 | 0 | 0 | 0 | 4.25 | 5.25 | 0 | 4.5
| 48 | 0 | 0.875 | 0 | 1.25 | 2.0833 | 0 | 0 | 2.5 | 0 | 0 | 4.25 | 4.75 | 0 | 5.25 | 5.25 | 4.75
| 49 | 0 | 0 | 0.875 | 0 | 2.5 | 2.25 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5.75 | 5.25
| 50 | 0 | 0 | 0 | 0 | 0 | 0 | 2.75 | 2.75 | 0 | 3.75 | 0 | 0 | 0 | 5.25 | 5.75 | 0
| 51 | 0 | 0 | 0 | 0 | 0 | 2.75 | 3.25 | 0 | 0 | 0 | 4.25 | 0 | 4.5 | 0 | 0 | 5.75
| 52 | 0 | 0 | 1.75 | 0 | 2.75 | 2.75 | 0 | 0 | 0 | 3.75 | 3.75 | 4.25 | 0 | 5.25 | 5.25 | 4.75
| 53 | 0 | 1.75 | 0 | 0 | 1.75 | 0 | 0 | 2.75 | 2.75 | 3.25 | 0 | 4.5833 | 0 | 5.25 | 5.75 | 5.25
| 54 | 0 | 0 | 1.75 | 0.875 | 0 | 2.75 | 2.75 | 0 | 2.75 | 0 | 4.75 | 0 | 4.75 | 0 | 0 | 5.25
| 55 | 1.75 | 0 | 0 | 0 | 0 | 0 | 2.75 | 0 | 2.25 | 0 | 4 | 0 | 3.75 | 0 | 5.75 | 5.25
| 56 | 0 | 1.25 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5.25 | 0 | 0 | 0 | 0 | 5.75 | 5.25
| 57 | 0.875 | 0 | 0 | 0 | 0 | 0 | 2.75 | 3 | 1.25 | 0 | 0 | 0 | 5.25 | 0 | 0 | 0
| 58 | 0 | 0 | 0 | 0 | 1.75 | 3 | 0 | 2.75 | 0 | 0 | 4.75 | 5 | 4.75 | 0 | 0 | 5.25
| 59 | 0 | 1.25 | 1.75 | 1.75 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5 | 0 | 0
| 60 | 0 | 0 | 1.25 | 2.25 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4.25 | 0 | 5 | 5.25 | 0
| 61 | 0 | 0 | 0 | 0 | 0 | 2.75 | 2.5 | 0 | 0 | 4.75 | 4.25 | 0 | 0 | 0 | 0 | 0
| 62 | 1.3125 | 0 | 0 | 0 | 0 | 2.25 | 0 | 3.25 | 0 | 0 | 4.75 | 0 | 0 | 0 | 5.25 | 0
| 63 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.75 | 0 | 4.25 | 0 | 0 | 0 | 0 | 5.75 | 4.75
| 64 | 0 | 0 | 0 | 0 | 2.75 | 0 | 0 | 4.25 | 0 | 0 | 4.25 | 4.25 | 5.25 | 5.25 | 0 | 0
| 65 | 0.875 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2.75 | 4.25 | 0 | 5.25 | 4.25 | 0 | 0 | 0
| 66 | 0 | 0 | 0 | 1.25 | 2.75 | 2.75 | 3 | 2.75 | 0 | 4.75 | 0 | 4.5 | 0 | 0 | 0 | 0
| 67 | 0 | 1.75 | 0 | 0 | 2.25 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4.25 | 0 | 0 | 5.25
| 68 | 1.75 | 0 | 0 | 0 | 2.25 | 0 | 0 | 0 | 0 | 4.25 | 0 | 0 | 4.75 | 4.75 | 4.5833 | 5
| 69 | 0 | 1.75 | 0 | 1.25 | 0 | 0 | 2.75 | 0 | 0 | 0 | 3.25 | 0 | 4.75 | 4.75 | 5 | 4.5
| 70 | 2 | 1.75 | 0 | 0 | 1.75 | 2.75 | 0 | 0 | 0 | 0 | 0 | 0 | 5.25 | 4.85 | 4.75 | 5.0357
| 71 | 0 | 0 | 2.25 | 0 | 0 | 2.75 | 0 | 2.75 | 2.75 | 0 | 0 | 0 | 0 | 4.45 | 4.75 | 4.75
| 72 | 0 | 0 | 0.875 | 1.75 | 0 | 0 | 2.75 | 0 | 2.75 | 0 | 3.25 | 0 | 0 | 4.5 | 4.6667 | 4.6786
| 73 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 3 | 0 | 3.375 | 3.75 | 0 | 0 | 4.75 | 4.9375 | 5.0625
| 74 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3.9167 | 0 | 5.25 | 4 | 4.9808 | 4.7143 | 5.3929
| 75 | 1.75 | 0 | 0 | 1.3125 | 0 | 0 | 2.75 | 3.5 | 0 | 3.05 | 3.95 | 3.5833 | 4.25 | 4.8056 | 4.7885 | 5.1029
| 76 | 0 | 0 | 1.75 | 0 | 1.75 | 2.25 | 2.25 | 0 | 2.5 | 3.45 | 3.25 | 3.65 | 3.3409 | 5.0417 | 5.0833 | 5.1944
| 77 | 0 | 1.75 | 0 | 0 | 2.75 | 2.75 | 0 | 0 | 2.75 | 3.1944 | 3.4239 | 3.4722 | 3.5 | 4.5338 | 5.25 | 4.7917
| 78 | 1.75 | 1.25 | 0 | 1.25 | 1.75 | 0 | 2.25 | 0 | 3.25 | 3.7812 | 3.5278 | 3.5577 | 3.6029 | 4.5459 | 4.925 | 4.9265
| 79 | 0 | 2.75 | 0 | 0 | 1.75 | 0 | 0 | 1.75 | 1.25 | 3.5938 | 3.5 | 3.6324 | 3.5 | 4.4955 | 4.55 | 4.5854
| 80 | 0 | 1.75 | 0 | 0 | 2.375 | 0 | 0 | 0 | 0 | 3.57 | 3.7262 | 4 | 3.8833 | 4.3808 | 4.6324 | 4.6699
| 81 | 1.75 | 0 | 0 | 0 | 1.6667 | 2.75 | 0 | 2.5 | 1.3472 | 3.6923 | 3.7308 | 3.7292 | 3.7237 | 4.4587 | 4.4694 | 4.504
| 82 | 0 | 0 | 0 | 0 | 1.8611 | 2.125 | 2.5 | 2.5 | 1.55 | 3.8393 | 3.7344 | 3.8864 | 3.9167 | 4.3037 | 4.4094 | 4.4782
| 83 | 0 | 0 | 0 | 0 | 1.9722 | 2 | 1.9167 | 1.75 | 1.4861 | 3.75 | 3.8145 | 3.8056 | 3.8214 | 4.3704 | 4.4406 | 4.6306
| 84 | 0 | 0 | 1.25 | 0 | 2 | 1.9167 | 1.8333 | 2 | 1.5655 | 3.6679 | 3.7875 | 3.875 | 4.2031 | 4.3252 | 4.4427 | 4.5333
| 85 | 0 | 0 | 1.75 | 0 | 1.754 | 2.0938 | 2.0978 | 1.9375 | 1.58 | 3.4912 | 3.7227 | 3.8944 | 3.8833 | 4.34 | 4.4789 | 4.5221
| 86 | 0 | 0 | 1.25 | 0 | 2.0375 | 2.1121 | 1.9107 | 2.2045 | 1.5781 | 3.3402 | 3.4626 | 3.4712 | 3.4761 | 4.3931 | 4.4956 | 4.6104
| 87 | 0 | 1.75 | 1.75 | 1.75 | 2.0058 | 2.0125 | 2 | 2.1591 | 1.7276 | 3.2523 | 3.2754 | 3.2529 | 3.3182 | 4.4583 | 4.5002 | 4.6033
| 88 | 0 | 0 | 0 | 0 | 2.0119 | 2.0872 | 2.1591 | 2.1196 | 1.7756 | 3.2158 | 3.2446 | 3.3029 | 3.4179 | 4.5003 | 4.5638 | 4.6114
| 89 | 0 | 1.75 | 0 | 0 | 2.0058 | 2.1436 | 2.2746 | 2.1481 | 1.638 | 3.1876 | 3.2321 | 3.2221 | 3.3257 | 4.5405 | 4.5825 | 4.6357
| 90 | 0 | 0.875 | 2.75 | 0 | 2.0227 | 2.2256 | 2.2007 | 2.2733 | 1.7058 | 3.1643 | 3.2094 | 3.2291 | 3.2738 | 4.5611 | 4.601 | 4.7014
| 91 | 0 | 1.75 | 1.25 | 0 | 2.0074 | 2.3382 | 2.4569 | 2.2273 | 1.6495 | 3.1555 | 3.1893 | 3.2514 | 3.2702 | 4.5923 | 4.6085 | 4.6782
| 92 | 1.3125 | 0 | 0 | 1.3125 | 1.8695 | 2.1424 | 2.3519 | 2.2172 | 1.4797 | 3.1286 | 3.2356 | 3.2403 | 3.2379 | 4.6039 | 4.5909 | 4.6949
| 93 | 1.25 | 0 | 0 | 0 | 1.8269 | 2.1474 | 2.3892 | 2.0967 | 1.4962 | 3.2245 | 3.2379 | 3.2841 | 3.3124 | 4.5791 | 4.6135 | 4.7033
| 94 | 0.875 | 0.875 | 1.0625 | 1.4167 | 1.8043 | 2.19 | 2.2049 | 1.727 | 1.4476 | 3.2366 | 3.3041 | 3.2955 | 3.3089 | 4.5895 | 4.6351 | 4.7024
| 95 | 0.95 | 1.1458 | 1.3393 | 0.875 | 1.8158 | 2.1197 | 2.0046 | 1.6114 | 1.4505 | 3.2898 | 3.3205 | 3.3371 | 3.3547 | 4.602 | 4.6459 | 4.7168
| 96 | 1.0833 | 1.0982 | 1 | 1.025 | 1.802 | 2.0254 | 2.0521 | 1.5936 | 1.5234 | 3.2993 | 3.3148 | 3.3531 | 3.3849 | 4.6062 | 4.6255 | 4.7359
| 97 | 1.2273 | 1.1538 | 1.1125 | 1.0833 | 1.834 | 1.9583 | 2.0013 | 1.6301 | 1.5318 | 3.3041 | 3.3402 | 3.3866 | 3.4121 | 4.6047 | 4.6567 | 4.7268
| 98 | 1.1 | 1.0724 | 1.2417 | 1.1411 | 1.8099 | 1.9499 | 1.9776 | 1.6535 | 1.5678 | 3.3232 | 3.3653 | 3.4297 | 3.4743 | 4.5457 | 4.6499 | 4.693
| 99 | 1.1187 | 1.1181 | 1.125 | 1.151 | 1.8548 | 1.9281 | 1.9472 | 1.6916 | 1.5669 | 3.3627 | 3.3723 | 3.4161 | 3.4552 | 4.5789 | 4.6664 | 4.7259
| 100 | 1.1192 | 1.2129 | 1.0966 | 1.1855 | 1.8767 | 1.9437 | 1.966 | 1.7147 | 1.6037 | 3.3619 | 3.4109 | 3.4071 | 3.4394 | 4.5332 | 4.6376 | 4.7467
| 101 | 1.1715 | 1.2067 | 1.1672 | 1.2049 | 1.8944 | 1.9699 | 1.9792 | 1.7443 | 1.6391 | 3.4 | 3.4177 | 3.423 | 3.482 | 4.5597 | 4.651 | 4.8225
| 102 | 1.1918 | 1.1433 | 1.1744 | 1.1623 | 1.9142 | 1.9975 | 1.9764 | 1.7786 | 1.6665 | 3.3925 | 3.4432 | 3.4212 | 3.4927 | 4.579 | 4.6702 | 4.7403
| 103 | 1.19 | 1.2141 | 1.2075 | 1.2603 | 1.9623 | 2.0228 | 2.0245 | 1.8303 | 1.7203 | 3.4165 | 3.4466 | 3.4496 | 3.4985 | 4.597 | 4.6565 | 4.7539
| 104 | 1.1852 | 1.1796 | 1.2259 | 1.295 | 1.9815 | 2.0328 | 2.0342 | 1.7992 | 1.7657 | 3.4072 | 3.4984 | 3.4391 | 3.5251 | 4.5802 | 4.6444 | 4.7219
| 105 | 1.1949 | 1.2717 | 1.2708 | 1.3225 | 2.0051 | 2.0678 | 2.0705 | 1.7799 | 1.8033 | 3.4194 | 3.5765 | 3.4517 | 3.525 | 4.5723 | 4.663 | 4.7345
| 106 | 1.2804 | 1.2846 | 1.3038 | 1.2888 | 2.0021 | 2.0861 | 2.098 | 1.8049 | 1.8672 | 3.4453 | 3.5736 | 3.5099 | 3.5431 | 4.6295 | 4.6793 | 4.7401
| 107 | 1.344 | 1.2889 | 1.2418 | 1.2379 | 2.0149 | 2.1085 | 2.104 | 1.8995 | 1.9409 | 3.4327 | 3.6562 | 3.5431 | 3.6271 | 4.6301 | 4.6818 | 4.7362
| 108 | 1.1877 | 1.1848 | 1.2324 | 1.1527 | 1.9975 | 2.115 | 2.1136 | 1.9023 | 1.9962 | 3.4543 | 3.6339 | 3.5498 | 3.6546 | 4.6531 | 4.6918 | 4.7803
| 109 | 1.1335 | 1.1666 | 1.1483 | 1.1692 | 1.9928 | 2.0926 | 2.1013 | 1.9938 | 2.0278 | 3.4838 | 3.6332 | 3.5568 | 3.6187 | 4.6615 | 4.6764 | 4.8506
| 110 | 1.1583 | 1.1517 | 1.1449 | 1.1307 | 2.0276 | 2.1083 | 2.1186 | 2.0031 | 2.0735 | 3.5122 | 3.6418 | 3.5979 | 3.7073 | 4.6962 | 4.7616 | 4.8519
| 111 | 1.1136 | 1.1246 | 1.1173 | 1.1058 | 2.0669 | 2.1449 | 2.1825 | 1.9434 | 2.091 | 3.5707 | 3.6483 | 3.6347 | 3.7065 | 4.6747 | 4.8213 | 4.8583
| 112 | 1.1142 | 1.1401 | 1.11 | 1.1227 | 2.1254 | 2.2485 | 2.2745 | 1.8745 | 2.1009 | 3.5946 | 3.6572 | 3.654 | 3.7364 | 4.6986 | 4.8217 | 4.8743
| 113 | 1.1019 | 1.1344 | 1.1343 | 1.125 | 2.1462 | 2.2925 | 2.2922 | 1.8716 | 2.1076 | 3.6561 | 3.7023 | 3.6728 | 3.7007 | 4.7237 | 4.8011 | 4.8856
| 114 | 1.1238 | 1.1683 | 1.1273 | 1.1596 | 2.1234 | 2.2652 | 2.2815 | 1.8863 | 2.1035 | 3.6815 | 3.7391 | 3.6794 | 3.7363 | 4.7418 | 4.8123 | 4.93
| 115 | 1.1625 | 1.1904 | 1.1737 | 1.1991 | 2.0951 | 2.2492 | 2.2626 | 1.9486 | 2.1157 | 3.7025 | 3.7583 | 3.7241 | 3.7379 | 4.7968 | 4.8814 | 4.9296
| 116 | 1.1756 | 1.2134 | 1.2116 | 1.2323 | 2.0654 | 2.2491 | 2.3152 | 2.0087 | 2.0938 | 3.6795 | 3.8676 | 3.8444 | 3.7894 | 4.8274 | 4.882 | 4.9516
| 117 | 1.2032 | 1.2436 | 1.2374 | 1.2806 | 2.1054 | 2.2237 | 2.2856 | 2.0256 | 2.1367 | 3.7673 | 3.8741 | 3.9026 | 3.9215 | 4.8537 | 4.9201 | 4.9804
| 118 | 1.2378 | 1.238 | 1.2663 | 1.2839 | 2.1478 | 2.2249 | 2.2497 | 2.041 | 2.2595 | 3.7342 | 3.8915 | 3.9333 | 3.9925 | 4.9207 | 4.9421 | 4.9531
| 119 | 1.2632 | 1.2731 | 1.2851 | 1.2908 | 2.2529 | 2.21 | 2.3228 | 2.0748 | 2.2635 | 3.8415 | 3.8475 | 3.9075 | 3.9767 | 4.9009 | 4.935 | 5.019
| 120 | 1.3073 | 1.3067 | 1.3086 | 1.2701 | 2.3315 | 2.2397 | 2.4899 | 2.1216 | 2.3048 | 3.9292 | 3.8424 | 3.9107 | 3.9686 | 4.9099 | 4.9448 | 5.0209
| 121 | 1.3388 | 1.3326 | 1.3146 | 1.3127 | 2.2901 | 2.3386 | 2.5209 | 2.1129 | 2.2478 | 3.9209 | 3.8899 | 3.9202 | 3.9481 | 4.8813 | 4.9926 | 5.0753
| 122 | 1.357 | 1.3476 | 1.322 | 1.3353 | 2.2435 | 2.5158 | 2.5127 | 2.1246 | 2.2644 | 3.8848 | 3.9663 | 3.9201 | 3.9366 | 4.9426 | 5.0148 | 5.0224
| 123 | 1.3587 | 1.3615 | 1.3769 | 1.3841 | 2.2053 | 2.5142 | 2.5139 | 2.2172 | 2.406 | 3.9337 | 4.1336 | 4.0056 | 4.1178 | 5.0079 | 5.0535 | 5.0486
| 124 | 1.3829 | 1.3688 | 1.402 | 1.4312 | 2.2549 | 2.4577 | 2.4767 | 2.2773 | 2.5081 | 4.0267 | 4.0772 | 4.1164 | 4.17 | 5.0285 | 4.9689 | 5.045
| 125 | 1.4171 | 1.3595 | 1.4631 | 1.5086 | 2.4411 | 2.5321 | 2.475 | 2.3023 | 2.4941 | 4.1602 | 4.1087 | 4.1436 | 4.1647 | 5.0423 | 4.9594 | 5.1908
| 126 | 1.475 | 1.3722 | 1.4766 | 1.562 | 2.4268 | 2.636 | 2.51 | 2.2991 | 2.4913 | 4.1536 | 4.033 | 4.1547 | 4.1031 | 4.9701 | 5.0213 | 5.2237
| 127 | 1.5016 | 1.4089 | 1.5076 | 1.6062 | 2.5235 | 2.6873 | 2.6916 | 2.3363 | 2.5 | 4.1835 | 3.9948 | 4.0844 | 4.1208 | 4.9711 | 5.1606 | 5.2165
| 128 | 1.516 | 1.4416 | 1.5252 | 1.6537 | 2.5296 | 2.7295 | 2.6941 | 2.3228 | 2.4964 | 4.0937 | 4.0207 | 4.0493 | 4.0989 | 5.1333 | 5.1967 | 5.2
| 129 | 1.5408 | 1.4822 | 1.5045 | 1.6733 | 2.5027 | 2.7027 | 2.6924 | 2.3744 | 2.5 | 4.112 | 4.0984 | 4.0115 | 4.2381 | 5.1727 | 5.1584 | 5.1032
| 130 | 1.5269 | 1.5024 | 1.5113 | 1.71 | 2.4607 | 2.6822 | 2.631 | 2.4357 | 2.5691 | 4.0617 | 4.1441 | 4.1225 | 4.4434 | 5.1965 | 5.0504 | 5.1025
| 131 | 1.5206 | 1.5162 | 1.5181 | 1.7767 | 2.5153 | 2.6748 | 2.662 | 2.5 | 2.6631 | 4.0759 | 4.2228 | 4.29 | 4.4418 | 5.2245 | 5.0293 | 5.0418
| 132 | 1.5208 | 1.5448 | 1.5275 | 1.8157 | 2.5688 | 2.6753 | 2.7347 | 2.5075 | 2.6813 | 4.1501 | 4.1988 | 4.366 | 4.4021 | 5.1641 | 5.1238 | 5.0824
| 133 | 1.5232 | 1.5277 | 1.4925 | 1.8031 | 2.6245 | 2.6902 | 2.7263 | 2.5469 | 2.6916 | 4.0924 | 4.0913 | 4.2203 | 4.3319 | 5.0343 | 5.1457 | 5.1192
| 134 | 1.3795 | 1.4451 | 1.3903 | 1.7403 | 2.531 | 2.5643 | 2.6591 | 2.375 | 2.5577 | 4.031 | 3.9781 | 4.0492 | 4.0682 | 4.9605 | 5.0122 | 5.0204
| 135 | 1.2303 | 1.2706 | 1.2612 | 1.6714 | 2.325 | 2.4118 | 2.4508 | 2.1518 | 2.3138 | 3.8071 | 3.7907 | 3.8783 | 3.8969 | 4.7965 | 4.8476 | 4.9194
| 136 | 1.2695 | 1.2752 | 1.2812 | 1.6363 | 2.1567 | 2.2322 | 2.2538 | 1.951 | 2.1216 | 3.5565 | 3.6097 | 3.652 | 3.6578 | 4.7171 | 4.7109 | 4.7497
| 137 | 1.2779 | 1.2829 | 1.2767 | 1.577 | 2.0449 | 2.1246 | 2.131 | 1.8218 | 1.8874 | 3.3713 | 3.3937 | 3.4389 | 3.4703 | 4.587 | 4.6571 | 4.687
| 138 | 1.2297 | 1.2853 | 1.2923 | 1.5142 | 2.0244 | 2.2129 | 2.1928 | 1.7264 | 1.8798 | 3.3987 | 3.436 | 3.4764 | 3.5218 | 4.5482 | 4.666 | 4.6884
| 139 | 1.2688 | 1.3367 | 1.2356 | 1.5669 | 2.1556 | 2.3224 | 2.2511 | 1.7503 | 1.9684 | 3.6155 | 3.6126 | 3.7201 | 3.8033 | 4.6972 | 4.7882 | 4.8126
| 140 | 1.3528 | 1.3669 | 1.3311 | 1.6759 | 2.1691 | 2.3322 | 2.3707 | 1.7921 | 1.881 | 3.7654 | 3.8031 | 3.756 | 3.8473 | 4.8269 | 4.9676 | 4.9463
| 141 | 1.4821 | 1.4118 | 1.4773 | 1.6764 | 2.2559 | 2.3185 | 2.4127 | 1.8244 | 2.0772 | 3.9392 | 3.8944 | 4.0995 | 4.0398 | 5.061 | 5.0246 | 5.0833
| 142 | 1.4062 | 1.375 | 1.4688 | 1.7273 | 2.3875 | 2.4028 | 2.5203 | 2.0291 | 2.0508 | 4.0068 | 4.0385 | 4.2045 | 4.125 | 5 | 5.1053 | 5.3676
| 143 | 1.5 | 1.375 | 1.75 | 1.8864 | 2.2917 | 2.4167 | 2.6 | 1.9667 | 2.2167 | 4.2237 | 4.3833 | 4.25 | 4.1136 | 5.2955 | 5.1389 | 5.2167
| 144 | 1.5833 | 1.5833 | 1.55 | 1.75 | 2.2833 | 2.75 | 2.55 | 1.9423 | 3.0682 | 4.125 | 4.25 | 4.3611 | 4.1786 | 5.4375 | 5.1136 | 5.375
| 145 | 2.25 | 1.275 | 1.375 | 1.8 | 2.75 | 2.75 | 2.65 | 2.2917 | 3.5 | 4.4167 | 4.3333 | 4.4722 | 4.4167 | 5.3929 | 5.4375 | 5.5
| 146 | 1.25 | 1.25 | 1.5 | 1.8333 | 2.625 | 2.4167 | 2.4167 | 1.9167 | 2.8333 | 4.75 | 4.25 | 4.3929 | 4.4643 | 5.125 | 5.75 | 5.5
| 147 | 1.5833 | 1.5 | 1.75 | 2.25 | 2.8929 | 2.5833 | 2.95 | 2.75 | 2.125 | 4 | 4.25 | 4.4722 | 4.4167 | 5.25 | 5.4167 | 5.5833
| 148 | 1.5833 | 1.5833 | 1.625 | 1.9167 | 2.5 | 2.875 | 2.65 | 1.75 | 2.45 | 5.25 | 4.25 | 4 | 4.625 | 0 | 5.75 | 5.5
| 149 | 1.4167 | 0 | 1.625 | 2.05 | 3 | 2.5 | 2.5 | 2.625 | 2.75 | 4.9167 | 4.75 | 4 | 5.25 | 0 | 5.5833 | 5.25
| 150 | 2.25 | 0 | 1.75 | 2.75 | 2.4167 | 3.5 | 2.75 | 2 | 0 | 0 | 4.4167 | 4.75 | 4.25 | 5 | 5.25 | 5.375
| 151 | 1.75 | 1.3125 | 1.75 | 1.75 | 3 | 2.25 | 2.75 | 2.25 | 2.75 | 4.25 | 4.75 | 4 | 4.5833 | 5.4167 | 5.75 | 5.5
| 152 | 0 | 0 | 1.75 | 2.25 | 2.25 | 0 | 0 | 1.75 | 2.75 | 4.75 | 4.4167 | 4.75 | 4.5833 | 5.25 | 0 | 5.25
| 153 | 0 | 0 | 1.75 | 2.25 | 2.25 | 2.75 | 0 | 2.25 | 0 | 4.25 | 5.25 | 4.75 | 4.5 | 0 | 0 | 4.75
| 154 | 0 | 1.5833 | 0 | 2.25 | 2.5 | 2.75 | 2.75 | 1.75 | 2.75 | 4.35 | 0 | 0 | 4.75 | 5.25 | 5.75 | 0
| 155 | 0 | 1.25 | 0 | 2.75 | 2.9167 | 2.5 | 3.25 | 0 | 0 | 5 | 4.75 | 0 | 4.4167 | 5.25 | 5.4167 | 5.25
| 156 | 1.75 | 0 | 0 | 2 | 2.75 | 1.75 | 0 | 2.9167 | 3.25 | 4.75 | 0 | 0 | 4.75 | 5.75 | 0 | 5.25
| 157 | 0 | 0 | 1.75 | 0 | 0 | 3 | 2.75 | 3.75 | 4.0833 | 0 | 0 | 4.5 | 5.25 | 5.75 | 5.25 | 4.75
| 158 | 1.25 | 1.25 | 1.75 | 2.75 | 0 | 2.75 | 3.25 | 2.25 | 2.75 | 0 | 4.75 | 0 | 4.75 | 5.75 | 5.25 | 5.75
| 159 | 1.5 | 0 | 1.4167 | 0 | 2.75 | 2.4167 | 0 | 0 | 3.25 | 4.25 | 3.25 | 4.75 | 3.75 | 0 | 5.25 | 5.25
| 160 | 0 | 0.875 | 1.75 | 0 | 0 | 0 | 0 | 2.4167 | 1.75 | 0 | 5.25 | 4.25 | 0 | 5.75 | 5.25 | 5.25
| 161 | 0 | 1.75 | 1.25 | 2.75 | 0 | 0 | 3 | 0 | 2.5 | 0 | 4.25 | 4 | 4.75 | 4.75 | 0 | 0
| 162 | 0 | 1.5 | 1.75 | 1.75 | 0 | 2.9167 | 3.25 | 2.75 | 0 | 4.75 | 4.25 | 4.5833 | 0 | 5 | 5.25 | 5.5
| 163 | 0.875 | 1.75 | 0 | 0 | 2.75 | 2.25 | 3 | 2.25 | 0 | 0 | 4.4167 | 4.75 | 5.25 | 0 | 0 | 5.25
| 164 | 1.25 | 0 | 1.75 | 1.75 | 0 | 2.25 | 0 | 0 | 0 | 4.75 | 0 | 0 | 0 | 0 | 0 | 5.25
| 165 | 1.75 | 0 | 0 | 2.25 | 0 | 2.75 | 3.25 | 2.5833 | 3.25 | 4.75 | 0 | 5.25 | 0 | 5.75 | 0 | 5.25
| 166 | 1.5 | 1.75 | 0 | 2.25 | 3 | 0 | 0 | 2.75 | 2.25 | 4.75 | 0 | 0 | 0 | 0 | 0 | 5.75
| 167 | 0 | 0 | 0 | 0 | 2.75 | 0 | 0 | 3.25 | 2.75 | 4.75 | 4.75 | 4.25 | 4.0833 | 5.25 | 5.75 | 0
| 168 | 1.75 | 1.75 | 0 | 2.75 | 0 | 2.25 | 2.75 | 0 | 0 | 0 | 0 | 4.75 | 3.75 | 0 | 5.75 | 5.25
| 169 | 0 | 0 | 0 | 2.25 | 2.75 | 0 | 0 | 0 | 0 | 0 | 3.75 | 0 | 0 | 5.25 | 0 | 5.75
| 170 | 0 | 0 | 0 | 0 | 0 | 0 | 2.75 | 0 | 3.75 | 0 | 4.5 | 4.25 | 4.25 | 5.75 | 5.75 | 0
| 171 | 0 | 0 | 0 | 1.75 | 2.75 | 0 | 3 | 0 | 2.75 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 172 | 0 | 1.75 | 0 | 0 | 0 | 0 | 3.25 | 2.75 | 0 | 4.4167 | 4.75 | 0 | 4.25 | 5.75 | 5.75 | 5.25
| 173 | 0 | 0 | 1.25 | 0 | 2.75 | 3.75 | 0 | 0 | 0 | 4.25 | 4.75 | 3.75 | 0 | 5.25 | 0 | 5.5
| 174 | 1.25 | 0 | 0 | 2.75 | 0 | 0 | 0 | 2.25 | 1.75 | 4.75 | 0 | 0 | 0 | 5.75 | 0 | 5.25
| 175 | 1.75 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2.25 | 0 | 0 | 3.25 | 4.75 | 5 | 0 | 0
| 176 | 0 | 0 | 0 | 0 | 0 | 3.75 | 2.75 | 0 | 2.75 | 5.25 | 0 | 0 | 3.75 | 0 | 0 | 0
| 177 | 0 | 0 | 0 | 1.75 | 2.75 | 0 | 0 | 0 | 0 | 4.75 | 4.25 | 4.25 | 4.75 | 0 | 0 | 6.25
| 178 | 0 | 0 | 0.875 | 0 | 3.25 | 0 | 0 | 0 | 0 | 4.75 | 4.25 | 0 | 0 | 0 | 0 | 5.75
| 179 | 1.25 | 1.75 | 0.875 | 2.25 | 2.75 | 3 | 0 | 0 | 2.75 | 0 | 5.25 | 5.25 | 3.75 | 0 | 0 | 0
| 180 | 1.75 | 1.25 | 1.75 | 0 | 3.25 | 3.75 | 0 | 0 | 3.25 | 5 | 4.75 | 0 | 4.75 | 0 | 6 | 0
| 181 | 0 | 0 | 1.25 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4.75 | 4.75 | 0 | 5.25
| 182 | 0 | 0 | 1.25 | 2.0833 | 2.75 | 2.25 | 0 | 1.75 | 3.25 | 0 | 0 | 4.25 | 3.75 | 0 | 4.5 | 5.25
| 183 | 0.875 | 0 | 1.25 | 2.5 | 0 | 0 | 2.5 | 0 | 3.75 | 4.25 | 4.75 | 0 | 4.25 | 5.25 | 4.75 | 0
| 184 | 0 | 0.875 | 0 | 0 | 2.25 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5.75 | 5.25 | 0
| 185 | 0 | 0 | 0 | 0 | 0 | 2.75 | 2.75 | 0 | 0 | 0 | 0 | 0 | 0 | 5.75 | 0 | 0
| 186 | 0 | 0 | 0 | 2.75 | 2.75 | 3.25 | 0 | 0 | 0 | 4.25 | 0 | 4.5 | 4.75 | 0 | 5.75 | 5.25
| 187 | 0 | 1.75 | 0 | 1.75 | 2.75 | 0 | 0 | 0 | 3.25 | 3.75 | 4.25 | 0 | 0 | 5.25 | 4.75 | 0
| 188 | 1.75 | 0 | 0 | 0 | 0 | 0 | 2.75 | 2.75 | 0 | 0 | 4.5833 | 0 | 4.25 | 5.75 | 5.25 | 5.25
| 189 | 0 | 1.75 | 0.875 | 0 | 2.75 | 2.75 | 0 | 2.75 | 3.25 | 4.75 | 0 | 4.75 | 4.75 | 0 | 5.25 | 5.25
| 190 | 0 | 0 | 0 | 0 | 0 | 2.75 | 0 | 2.25 | 0 | 4 | 0 | 3.75 | 5.25 | 5.75 | 5.25 | 5.75
| 191 | 1.25 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5.75 | 5.25 | 5.75
| 192 | 0 | 0 | 0 | 1.75 | 0 | 2.75 | 3 | 1.25 | 0 | 0 | 0 | 5.25 | 3.75 | 0 | 0 | 0
| 193 | 0 | 0 | 0 | 0 | 3 | 0 | 2.75 | 0 | 3.75 | 4.75 | 5 | 4.75 | 0 | 0 | 5.25 | 0
| 194 | 1.25 | 1.75 | 1.75 | 0 | 0 | 0 | 0 | 0 | 3.25 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| 195 | 0 | 1.25 | 2.25 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4.25 | 0 | 0 | 5.25 | 0 | 0
| Scantime | T3-T2 % Error for each Scan
| 1 | -25.098 | -22.745 | -22.745 | -21.9608 | -21.4886 | -22.6891 | -22.9292 | -22.9291 | -23.4093 | -22.5583 | -22.3854 | -22.5583 | -22.7312 | -20.9459 | -22.2973 | -22.1621
| 2 | -23.1372 | -24.3137 | -22.745 | -23.1372 | -16.6866 | -22.6891 | -22.9291 | -22.449 | -22.449 | -22.3855 | -22.2126 | -22.5583 | -22.2126 | -22.1621 | -22.0271 | -22.1622
| 3 | -23.5294 | -24.7059 | -22.745 | -23.5294 | -13.8055 | -22.449 | -22.9291 | -22.689 | -22.449 | -22.5583 | -22.5583 | -23.4226 | -22.904 | -22.1621 | -22.4325 | -22.7027
| 4 | -23.9215 | -25.098 | -24.3137 | -23.1372 | -12.1249 | -22.689 | -23.4093 | -23.6495 | -22.9291 | -22.3855 | -22.7312 | -22.9041 | -22.3854 | -21.2162 | -22.7027 | -22.973
| 5 | -22.7451 | -23.9215 | -23.9215 | -23.5294 | -13.8056 | -22.6891 | -21.7287 | -23.1693 | -23.1692 | -23.2498 | -22.2126 | -22.0398 | -22.3854 | -21.4865 | -22.7027 | -22.4325
| 6 | -23.5294 | -22.745 | -21.9607 | -23.1372 | -13.5654 | -23.4094 | -22.689 | -22.449 | -23.1692 | -22.2126 | -22.2126 | -22.0397 | -22.3855 | -21.7567 | -21.4865 | -22.2973
| 7 | -23.9215 | -23.1373 | -24.7059 | -24.7059 | -15.7263 | -23.4094 | -23.6495 | -22.9291 | -23.4093 | -21.8669 | -22.5583 | -22.0397 | -23.0769 | -21.6216 | -22.2973 | -21.8919
| 8 | -25.098 | -23.5295 | -24.7059 | -22.353 | -18.3674 | -23.4094 | -22.4489 | -23.4094 | -21.7287 | -21.694 | -23.4227 | -22.9041 | -23.0769 | -21.3514 | -22.2973 | -22.4324
| 9 | -23.5295 | -24.7059 | -23.1372 | -23.9215 | -19.8079 | -22.6891 | -22.9291 | -22.449 | -22.449 | -22.9041 | -23.7684 | -23.4226 | -23.2498 | -21.6216 | -22.5676 | -22.7027
| 10 | -25.098 | -23.5294 | -22.3529 | -22.745 | -16.4466 | -22.9292 | -22.689 | -23.6495 | -22.2088 | -21.694 | -22.9041 | -23.0769 | -23.4226 | -21.8919 | -22.4325 | -22.1621
| 11 | -23.5295 | -22.353 | -23.1372 | -23.1372 | -15.9664 | -21.9688 | -22.2089 | -23.8896 | -22.449 | -22.5583 | -21.8669 | -22.3855 | -22.2126 | -22.027 | -22.2973 | -22.0271
| 12 | -24.3137 | -23.1372 | -23.5294 | -23.9215 | -16.9268 | -22.2089 | -22.9292 | -24.3698 | -22.2089 | -21.6941 | -22.7312 | -22.5583 | -22.5583 | -22.1622 | -22.2973 | -22.2973
| 13 | -26.2745 | -25.8824 | -23.5294 | -24.3137 | -18.6075 | -23.6495 | -22.9291 | -22.4489 | -21.7287 | -22.5583 | -22.7312 | -23.0769 | -23.0769 | -22.4325 | -22.5676 | -22.2973
| 14 | -25.098 | -24.7059 | -23.1372 | -22.7451 | -19.0876 | -23.1692 | -23.6495 | -23.1693 | -22.6891 | -23.0769 | -23.4226 | -22.3855 | -22.904 | -22.5676 | -22.973 | -23.1081
| 15 | -22.745 | -25.098 | -22.7451 | -24.7059 | -20.2881 | -22.2089 | -22.9291 | -22.6891 | -22.449 | -23.0769 | -22.904 | -22.7312 | -22.0398 | -21.4865 | -22.4325 | -22.2973
| 16 | -22.7451 | -24.3137 | -24.3137 | -22.745 | -16.4466 | -22.449 | -21.9687 | -21.7287 | -21.9687 | -22.9041 | -22.3854 | -23.2498 | -23.2498 | -21.6216 | -22.027 | -22.2973
| 17 | -25.098 | -23.1372 | -21.9608 | -23.9215 | -16.9268 | -22.689 | -22.9292 | -22.6891 | -21.7287 | -22.2126 | -22.9041 | -22.3855 | -22.9041 | -21.7567 | -21.7567 | -22.4325
| 18 | -23.5294 | -23.5294 | -26.2745 | -23.1372 | -18.8475 | -23.6495 | -23.6495 | -22.689 | -22.449 | -22.5583 | -22.5583 | -22.904 | -22.9041 | -21.8919 | -21.8919 | -22.5676
| 19 | -22.745 | -22.3529 | -22.353 | -22.7451 | -19.3277 | -23.4094 | -23.8896 | -25.0901 | -22.449 | -22.904 | -22.0397 | -23.0769 | -23.4226 | -22.5676 | -22.7027 | -22.8378
| 20 | -23.5294 | -22.745 | -22.7451 | -23.1372 | -18.1273 | -23.1692 | -21.4886 | -23.1692 | -22.449 | -21.694 | -23.4227 | -23.0769 | -22.7312 | -22.1622 | -22.2973 | -21.8919
| 21 | -24.7059 | -23.5294 | -23.1373 | -23.5294 | -18.3674 | -21.9687 | -22.689 | -21.4886 | -22.2088 | -22.3854 | -22.3854 | -23.0769 | -22.2126 | -22.1621 | -22.7027 | -21.7568
| 22 | -24.3137 | -25.4902 | -23.1373 | -24.3137 | -19.5678 | -21.7286 | -21.9687 | -22.9291 | -22.2088 | -22.3854 | -22.9041 | -22.9041 | -22.904 | -22.4324 | -21.8919 | -22.2973
| 23 | -25.098 | -24.7059 | -24.3137 | -24.7059 | -19.0876 | -23.6495 | -23.1692 | -23.1693 | -22.4489 | -22.3854 | -23.0769 | -22.7312 | -23.2498 | -22.1621 | -21.6216 | -21.8919
| 24 | -23.5295 | -24.3138 | -23.5294 | -25.098 | -18.8476 | -24.1297 | -23.1692 | -23.8896 | -22.6891 | -22.5583 | -22.904 | -22.904 | -22.904 | -22.1621 | -23.1081 | -22.7027
| 25 | -23.1373 | -23.9216 | -23.9215 | -22.3529 | -19.3277 | -23.1692 | -21.2485 | -22.9292 | -21.9688 | -22.5583 | -22.7312 | -24.1141 | -22.3855 | -21.2162 | -22.8379 | -22.2973
| 26 | -21.5686 | -24.7059 | -24.3137 | -23.1372 | -19.808 | -23.1693 | -22.2089 | -23.1693 | -22.2089 | -22.5583 | -22.904 | -22.3855 | -22.5583 | -22.5676 | -22.7027 | -22.4325
| 27 | -23.1372 | -23.1373 | -23.9215 | -25.4902 | -20.5282 | -23.6495 | -22.6891 | -22.9292 | -21.7287 | -22.0398 | -22.5583 | -22.2126 | -23.4226 | -22.027 | -22.5676 | -22.4325
| 28 | -24.3137 | -25.8824 | -25.4902 | -23.5294 | -20.7683 | -23.1693 | -23.6495 | -21.4886 | -21.9687 | -22.7312 | -23.2498 | -22.2126 | -23.7684 | -22.4324 | -22.8378 | -22.2973
| 29 | -24.7059 | -24.7059 | -23.1373 | -22.745 | -20.048 | -22.689 | -22.9292 | -21.9688 | -21.7287 | -23.0769 | -22.3854 | -22.3854 | -23.2498 | -21.8919 | -23.1081 | -23.1081
| 30 | -24.3137 | -21.9608 | -23.9215 | -24.7059 | -19.5679 | -22.689 | -23.6495 | -23.4093 | -22.2089 | -23.9412 | -22.9041 | -22.5583 | -22.9041 | -21.6216 | -22.1622 | -22.8379
| 31 | -23.9215 | -22.745 | -23.5294 | -22.353 | -19.5679 | -22.4489 | -22.689 | -23.4093 | -22.2088 | -22.7312 | -22.5583 | -22.3855 | -22.3855 | -21.7567 | -22.1621 | -22.4324
| 32 | -23.5294 | -23.9215 | -22.3529 | -24.3137 | -20.5282 | -23.4094 | -23.1692 | -22.689 | -21.2485 | -22.7312 | -23.0769 | -22.2126 | -23.0769 | -22.2973 | -22.5675 | -22.2973
| 33 | -24.3137 | -25.098 | -23.9215 | -24.3137 | -19.5678 | -22.449 | -23.1692 | -22.4489 | -21.4886 | -22.3854 | -23.7684 | -23.0769 | -23.0769 | -22.1622 | -22.5676 | -21.7567
| 34 | -24.3137 | -22.353 | -24.3137 | -23.1372 | -21.0084 | -23.1692 | -23.6495 | -23.8896 | -21.9688 | -22.2126 | -22.7312 | -22.2126 | -23.2498 | -22.2973 | -22.973 | -22.7027
| 35 | -24.3137 | -24.3138 | -22.7451 | -23.1372 | -19.5679 | -23.8896 | -22.9292 | -22.449 | -21.7287 | -21.8669 | -22.5583 | -22.3855 | -23.5955 | -21.6216 | -22.973 | -22.7027
| 36 | -25.098 | -23.1373 | -24.3137 | -23.1372 | -19.3278 | -22.6891 | -23.4094 | -22.9292 | -20.7683 | -23.0769 | -21.8669 | -22.7312 | -22.9041 | -22.1622 | -21.8919 | -22.8379
| 37 | -23.5294 | -24.3137 | -23.1372 | -23.1372 | -21.0084 | -22.9291 | -23.6495 | -22.9291 | -21.4886 | -22.9041 | -23.0769 | -22.904 | -22.904 | -21.6216 | -22.2973 | -22.4325
| 38 | -25.098 | -23.1372 | -24.3137 | -23.5295 | -20.7683 | -23.4094 | -24.1296 | -23.1692 | -21.7286 | -22.9041 | -23.4226 | -22.5583 | -23.0769 | -22.7027 | -22.8379 | -23.5135
| 39 | -24.3138 | -24.7059 | -23.1372 | -23.1372 | -20.7683 | -22.449 | -23.6495 | -23.1693 | -20.7683 | -22.7312 | -23.2498 | -22.904 | -22.904 | -22.2973 | -22.8379 | -22.973
| 40 | -23.9216 | -23.1372 | -21.9608 | -22.353 | -20.048 | -23.6495 | -22.689 | -23.4094 | -20.2881 | -22.7312 | -22.0397 | -23.2498 | -22.7312 | -22.2973 | -22.2973 | -22.8378
| 41 | -24.7059 | -23.5294 | -21.9608 | -22.3529 | -20.2881 | -24.1297 | -22.2089 | -22.2089 | -21.2485 | -22.5583 | -22.5583 | -22.5583 | -22.904 | -21.7567 | -22.2973 | -21.7568
| 42 | -25.4902 | -22.745 | -20.7843 | -23.9215 | -20.048 | -21.9688 | -21.9687 | -22.2089 | -20.048 | -22.7312 | -22.9041 | -22.7312 | -23.5955 | -22.2973 | -22.7027 | -22.4324
| 43 | -24.7059 | -25.098 | -22.745 | -24.7059 | -21.2485 | -22.6891 | -23.1693 | -21.7287 | -21.7287 | -23.4227 | -22.7312 | -23.0769 | -23.2498 | -22.4325 | -22.5676 | -22.4324
| 44 | -24.7059 | -23.1372 | -23.9215 | -25.098 | -20.7683 | -22.449 | -23.8896 | -23.1692 | -21.2485 | -22.3854 | -22.5583 | -22.7312 | -23.4226 | -22.7027 | -22.5675 | -22.4324
| 45 | -24.3137 | -21.9608 | -23.9215 | -25.8824 | -21.9687 | -21.4886 | -24.6099 | -21.9687 | -21.4886 | -22.9041 | -22.5583 | -23.0769 | -23.4227 | -22.1621 | -22.0271 | -22.7027
| 46 | -24.7059 | -23.5294 | -23.5294 | -23.1373 | -20.048 | -22.9291 | -22.4489 | -22.4489 | -21.4886 | -22.3854 | -22.7312 | -22.5583 | -23.2498 | -22.1622 | -22.1621 | -22.4325
| 47 | -24.3137 | -24.3137 | -22.745 | -24.3137 | -21.0084 | -22.449 | -21.9687 | -22.6891 | -21.2485 | -22.9041 | -22.3855 | -23.0769 | -23.0769 | -21.8919 | -22.5676 | -22.5675
| 48 | -24.7059 | -22.745 | -22.353 | -23.5294 | -21.4885 | -21.9688 | -23.1693 | -22.6891 | -20.7683 | -22.3855 | -22.904 | -22.2126 | -23.9412 | -22.8378 | -22.4324 | -22.1622
| 49 | -23.1373 | -24.3137 | -25.098 | -22.3529 | -21.7286 | -23.4093 | -22.9292 | -22.4489 | -21.0084 | -22.3854 | -23.4226 | -22.5583 | -23.0769 | -22.7027 | -22.7027 | -23.1081
| 50 | -23.5295 | -25.4902 | -23.1372 | -22.353 | -21.4886 | -23.4094 | -23.4093 | -23.1692 | -21.2485 | -23.0769 | -22.9041 | -22.9041 | -22.2126 | -22.7027 | -22.8378 | -22.1621
| 51 | -23.9215 | -22.7451 | -22.353 | -22.745 | -22.449 | -23.4093 | -22.9291 | -23.8895 | -20.2881 | -22.904 | -22.9041 | -22.5583 | -22.2126 | -21.8919 | -22.7027 | -22.1621
| 52 | -23.9215 | -23.5294 | -23.5294 | -24.7059 | -21.9688 | -23.6495 | -22.9291 | -23.6495 | -20.048 | -22.7312 | -22.7312 | -23.7684 | -23.0769 | -22.027 | -22.4324 | -22.027
| 53 | -24.3137 | -25.4902 | -25.8824 | -25.098 | -20.5282 | -23.1692 | -22.6891 | -23.1692 | -21.0084 | -23.0769 | -23.5955 | -23.4227 | -23.0769 | -22.2973 | -22.2973 | -22.1622
| 54 | -24.3137 | -24.3137 | -23.1372 | -25.098 | -21.7287 | -23.6495 | -22.2088 | -22.4489 | -21.0084 | -23.4227 | -22.5583 | -22.7312 | -23.0769 | -22.2973 | -22.4325 | -22.4324
| 55 | -23.5294 | -24.7059 | -23.1372 | -23.1373 | -22.2089 | -22.9291 | -22.6891 | -23.1693 | -21.0084 | -22.3855 | -23.2498 | -22.5583 | -22.5583 | -22.1622 | -22.4324 | -22.4324
| 56 | -23.9216 | -23.9215 | -23.9215 | -23.9216 | -22.2089 | -22.2089 | -22.449 | -22.2088 | -20.2881 | -22.3854 | -22.3854 | -22.2126 | -22.7312 | -21.8919 | -22.973 | -22.5676
| 57 | -23.1372 | -22.353 | -21.9608 | -23.5294 | -20.7683 | -22.9292 | -22.4489 | -23.6495 | -20.7683 | -22.3855 | -23.7684 | -22.3855 | -22.7312 | -22.2973 | -22.7027 | -22.5676
| 58 | -25.098 | -23.9216 | -22.7451 | -23.5294 | -22.2089 | -22.689 | -23.8896 | -22.9291 | -21.0084 | -22.5583 | -23.2498 | -21.8669 | -22.3855 | -22.1621 | -23.2432 | -21.8919
| 59 | -24.3137 | -24.3138 | -22.7451 | -23.5294 | -22.9292 | -22.689 | -23.6495 | -22.9291 | -21.4886 | -23.5955 | -22.904 | -22.7312 | -22.2126 | -22.5676 | -23.1081 | -22.1622
| 60 | -23.9216 | -24.7059 | -25.4901 | -21.9608 | -22.2089 | -22.6891 | -23.1692 | -22.449 | -20.2881 | -22.7312 | -22.7312 | -22.904 | -22.9041 | -22.7027 | -22.7027 | -22.4325
| 61 | -23.5294 | -25.098 | -22.745 | -24.3137 | -22.2088 | -22.6891 | -22.6891 | -21.7287 | -20.5282 | -22.0397 | -22.7312 | -22.904 | -22.2126 | -22.1621 | -22.4324 | -22.5676
| 62 | -23.5294 | -23.1372 | -23.9215 | -24.7059 | -21.2485 | -23.1692 | -22.9291 | -22.9292 | -21.4886 | -22.3854 | -22.904 | -22.5583 | -22.7312 | -22.2973 | -22.4324 | -22.7027
| 63 | -24.3137 | -23.5294 | -22.745 | -24.3137 | -21.9688 | -23.6495 | -23.4094 | -21.9688 | -20.7683 | -22.904 | -22.7312 | -23.2498 | -22.904 | -21.8919 | -22.5676 | -22.7027
| 64 | -25.8824 | -24.3137 | -23.1372 | -22.7451 | -24.1297 | -23.8896 | -22.689 | -23.1692 | -21.2485 | -22.904 | -23.0769 | -23.2498 | -22.7312 | -23.1081 | -22.4324 | -22.4324
| 65 | -23.5295 | -23.9215 | -22.745 | -23.1372 | -22.2089 | -22.9291 | -23.1692 | -22.449 | -21.4886 | -22.904 | -22.9041 | -22.7312 | -22.5583 | -22.5676 | -22.4325 | -22.2973
| 66 | -22.353 | -23.9215 | -23.9215 | -22.7451 | -21.4886 | -21.4886 | -23.1693 | -22.9291 | -20.7683 | -22.3854 | -22.7312 | -22.904 | -22.0397 | -22.2973 | -22.4324 | -22.2973
| 67 | -23.9215 | -24.7059 | -22.3529 | -22.7451 | -21.7287 | -22.2089 | -22.2088 | -22.4489 | -20.048 | -23.0769 | -22.3854 | -22.3855 | -22.5583 | -22.4324 | -22.2973 | -22.1622
| 68 | -24.3137 | -23.5295 | -23.1372 | -24.3137 | -21.4886 | -23.6495 | -23.6495 | -21.9688 | -21.0084 | -23.4227 | -23.0769 | -23.2498 | -23.0769 | -22.027 | -22.5675 | -22.5675
| 69 | -22.7451 | -23.1372 | -23.1372 | -23.1372 | -23.4094 | -22.4489 | -24.1297 | -22.9292 | -20.7683 | -22.5583 | -23.0769 | -22.5583 | -22.3854 | -22.2973 | -22.2973 | -22.2973
| 70 | -24.7059 | -23.1373 | -23.5294 | -23.9215 | -22.9292 | -22.689 | -23.1692 | -24.1297 | -22.2088 | -22.7312 | -23.2498 | -22.5583 | -22.3855 | -22.2973 | -22.0271 | -22.8379
| 71 | -24.3137 | -22.353 | -21.5686 | -23.5294 | -22.689 | -21.9688 | -22.9291 | -21.2485 | -19.8079 | -22.9041 | -22.904 | -22.5583 | -22.5583 | -22.4324 | -22.4324 | -22.1622
| 72 | -24.3137 | -23.9215 | -23.9215 | -23.1373 | -21.7287 | -22.6891 | -22.2088 | -23.1693 | -19.5678 | -23.0769 | -22.3855 | -21.8669 | -23.4227 | -22.5676 | -22.5675 | -22.1621
| 73 | -24.3137 | -23.9216 | -25.098 | -25.098 | -22.9291 | -23.1693 | -22.2088 | -22.4489 | -20.7683 | -22.5584 | -22.2126 | -22.0398 | -22.5583 | -22.4324 | -22.027 | -22.027
| 74 | -23.9216 | -23.9215 | -25.098 | -22.7451 | -23.4093 | -23.4094 | -23.1693 | -23.6495 | -20.7683 | -22.7312 | -23.2498 | -22.3854 | -22.7312 | -23.2432 | -22.8378 | -22.7027
| 75 | -22.353 | -24.3137 | -21.1765 | -23.1373 | -22.449 | -22.9291 | -23.6495 | -23.8896 | -20.2882 | -23.0769 | -22.904 | -22.5583 | -22.0398 | -22.027 | -21.7567 | -22.7027
| 76 | -22.353 | -23.9215 | -22.3529 | -24.3137 | -22.6891 | -22.689 | -22.6891 | -22.449 | -20.2881 | -22.0398 | -22.7312 | -22.5583 | -21.6941 | -22.2973 | -21.6216 | -22.2973
| 77 | -22.7451 | -23.9215 | -23.9215 | -22.353 | -23.4093 | -22.2088 | -23.4094 | -23.8896 | -20.7683 | -22.5583 | -22.7312 | -23.0769 | -22.7312 | -21.8919 | -22.4324 | -22.2973
| 78 | -24.3137 | -24.7059 | -24.3137 | -24.3137 | -22.2089 | -22.689 | -22.6891 | -23.6495 | -20.7683 | -22.5583 | -22.904 | -22.5583 | -22.3855 | -22.973 | -22.7027 | -22.7027
| 79 | -24.3137 | -23.1372 | -23.9215 | -23.9215 | -21.2485 | -23.4093 | -23.4094 | -21.9688 | -20.7683 | -23.5955 | -22.5583 | -22.7312 | -22.5583 | -22.2973 | -22.973 | -22.027
| 80 | -22.7451 | -24.7059 | -25.098 | -24.3137 | -22.449 | -22.4489 | -22.6891 | -22.9292 | -21.0084 | -22.7312 | -22.9041 | -23.2498 | -22.0398 | -22.1621 | -22.4325 | -22.4324
| 81 | -24.7059 | -22.353 | -24.3137 | -23.5294 | -22.449 | -23.1693 | -22.4489 | -22.2089 | -20.2881 | -22.904 | -22.3855 | -22.3854 | -22.7312 | -21.8919 | -21.7567 | -22.1622
| 82 | -25.4902 | -21.9608 | -25.098 | -22.7451 | -22.9292 | -23.8896 | -22.9291 | -22.449 | -20.7683 | -21.694 | -22.5583 | -22.3854 | -22.0397 | -22.7027 | -22.2973 | -22.4324
| 83 | -23.5294 | -23.9215 | -23.5294 | -23.9215 | -23.1693 | -22.9291 | -23.4093 | -22.449 | -20.7683 | -23.2498 | -22.9041 | -22.5583 | -22.7312 | -22.1622 | -22.7027 | -22.4324
| 84 | -24.7059 | -21.9608 | -21.9608 | -22.353 | -22.9292 | -23.1692 | -22.689 | -23.1693 | -20.5282 | -22.9041 | -23.0769 | -22.0397 | -22.3855 | -22.5676 | -22.5676 | -22.8378
| 85 | -23.5294 | -24.3137 | -22.745 | -21.9607 | -22.449 | -22.6891 | -21.9688 | -22.9291 | -19.5678 | -21.8669 | -22.3855 | -21.8669 | -22.3855 | -22.973 | -22.2973 | -22.1622
| 86 | -23.9215 | -23.5294 | -23.5294 | -23.1372 | -22.2088 | -23.1693 | -21.7287 | -21.0084 | -20.7683 | -22.2126 | -23.2498 | -22.7312 | -22.5583 | -22.973 | -22.5676 | -22.0271
| 87 | -25.8823 | -23.5294 | -23.1372 | -22.7451 | -22.689 | -22.449 | -21.9688 | -20.7683 | -20.7683 | -22.2126 | -22.5583 | -22.5583 | -23.0769 | -22.1622 | -22.1622 | -22.0271
| 88 | -23.1373 | -24.3137 | -23.9215 | -22.745 | -22.449 | -23.1692 | -21.9688 | -20.5282 | -20.2881 | -23.0769 | -23.4226 | -22.3855 | -23.5955 | -22.973 | -21.7568 | -22.4325
| 89 | -23.1372 | -23.5294 | -21.5686 | -23.9215 | -21.9688 | -23.1693 | -23.1692 | -21.7287 | -20.2881 | -23.0769 | -23.2498 | -22.904 | -22.0398 | -22.2973 | -22.973 | -22.2973
| 90 | -23.1373 | -22.7451 | -21.9608 | -23.1373 | -22.9292 | -23.4094 | -23.1693 | -20.7683 | -20.5282 | -22.904 | -22.3855 | -22.2126 | -22.0397 | -22.2973 | -22.973 | -22.2973
| 91 | -20.7843 | -23.5294 | -23.9215 | -23.1372 | -22.449 | -22.4489 | -23.1693 | -22.4489 | -19.3277 | -23.0769 | -22.7312 | -22.904 | -22.7312 | -21.7567 | -21.8919 | -21.7568
| 92 | -22.745 | -23.1372 | -22.7451 | -23.5294 | -22.689 | -23.4094 | -21.9688 | -21.2485 | -20.5282 | -22.5583 | -22.3855 | -22.7312 | -22.7312 | -22.1622 | -22.2973 | -22.2973
| 93 | -24.3137 | -24.3137 | -24.7059 | -24.7059 | -23.1693 | -22.6891 | -23.4093 | -22.9291 | -21.2485 | -23.0769 | -22.7312 | -22.7312 | -23.0769 | -22.973 | -22.4325 | -22.1621
| 94 | -24.7059 | -23.5295 | -23.5294 | -22.7451 | -24.1297 | -23.1692 | -23.1692 | -23.1692 | -21.4886 | -22.5583 | -23.2498 | -22.5583 | -23.0769 | -22.4325 | -22.1621 | -22.4324
| 95 | -23.5295 | -22.745 | -24.3137 | -22.7451 | -22.449 | -22.2089 | -22.9291 | -24.1297 | -20.7683 | -22.0398 | -22.7312 | -21.8669 | -21.8669 | -22.5676 | -23.3784 | -22.027
| 96 | -21.9608 | -25.098 | -23.9215 | -23.1373 | -21.7287 | -24.1297 | -22.449 | -24.1296 | -19.808 | -22.3855 | -22.7312 | -23.0769 | -21.5212 | -22.5675 | -22.5676 | -22.4324
| 97 | -23.5294 | -23.1372 | -24.7059 | -23.9215 | -22.689 | -22.449 | -22.4489 | -24.85 | -21.2485 | -23.0769 | -21.8669 | -22.904 | -22.9041 | -22.2973 | -22.4324 | -22.4324
| 98 | -22.7451 | -24.3137 | -24.7059 | -24.7059 | -22.2088 | -22.689 | -23.4093 | -25.0901 | -21.2485 | -22.2126 | -23.2498 | -22.3855 | -22.3855 | -22.973 | -22.7027 | -22.2973
| 99 | -23.1372 | -23.9215 | -24.3137 | -24.7059 | -22.9291 | -23.6495 | -23.6495 | -25.8103 | -20.7683 | -23.0769 | -22.7312 | -22.904 | -23.5955 | -22.4324 | -22.973 | -23.1081
| 100 | -23.1373 | -23.1373 | -26.2745 | -23.9215 | -23.8895 | -22.9292 | -23.1693 | -26.0504 | -20.7683 | -22.5583 | -22.0398 | -22.7312 | -21.6941 | -21.8919 | -22.7027 | -22.5676
| 101 | -22.7451 | -23.5294 | -23.1372 | -23.5295 | -21.2485 | -21.9688 | -22.9292 | -25.5703 | -20.7683 | -22.2126 | -22.0398 | -23.0769 | -22.3854 | -21.6216 | -22.0271 | -21.8919
| 102 | -22.3529 | -23.9215 | -23.5295 | -21.5686 | -21.9688 | -22.6891 | -23.4094 | -26.5306 | -21.4886 | -22.3855 | -22.904 | -22.3854 | -23.0769 | -22.2973 | -21.8919 | -22.0271
| 103 | -25.098 | -23.5294 | -23.1372 | -23.5294 | -22.4489 | -22.4489 | -22.9292 | -26.7707 | -20.2881 | -22.7312 | -23.0769 | -22.3855 | -22.3854 | -22.7027 | -22.5676 | -22.7027
| 104 | -23.5294 | -23.5294 | -21.9608 | -23.1372 | -23.4093 | -22.689 | -23.8896 | -27.491 | -21.2484 | -23.2498 | -23.0769 | -22.9041 | -22.0397 | -22.1622 | -22.5676 | -22.4325
| 105 | -20.7844 | -23.9215 | -23.9215 | -23.5294 | -22.9291 | -22.2089 | -22.2089 | -27.0108 | -20.7683 | -22.9041 | -22.5583 | -22.3854 | -21.8669 | -21.7568 | -22.5676 | -22.2973
| 106 | -24.3137 | -24.3137 | -23.9215 | -23.5294 | -22.9291 | -23.1693 | -22.6891 | -27.2509 | -20.5282 | -22.7312 | -23.5955 | -22.5583 | -22.7312 | -22.5676 | -22.5675 | -22.2973
| 107 | -23.5294 | -23.5294 | -23.9216 | -24.3137 | -21.9688 | -22.2088 | -22.2089 | -27.7311 | -21.0084 | -23.2498 | -22.5583 | -23.2497 | -22.2126 | -21.8919 | -22.2973 | -22.027
| 108 | -23.9215 | -22.353 | -23.9215 | -23.1372 | -23.1692 | -24.1297 | -22.449 | -27.731 | -20.7683 | -22.7312 | -22.5583 | -22.5583 | -22.2126 | -22.7027 | -22.027 | -22.1622
| 109 | -23.1373 | -23.1373 | -24.3137 | -23.5294 | -24.3698 | -23.6495 | -24.3698 | -27.491 | -20.7683 | -22.9041 | -23.2498 | -22.7312 | -21.8669 | -22.8378 | -22.7027 | -22.2973
| 110 | -22.745 | -22.3529 | -22.745 | -21.5686 | -23.1692 | -22.9291 | -23.4093 | -26.2905 | -20.5282 | -22.7312 | -22.904 | -22.7312 | -22.5583 | -22.973 | -22.1622 | -22.973
| 111 | -24.3137 | -23.5294 | -23.1373 | -22.745 | -22.4489 | -22.9292 | -23.4094 | -26.2905 | -21.0084 | -22.0398 | -22.5583 | -22.904 | -21.8669 | -22.027 | -22.7027 | -21.8919
| 112 | -24.3137 | -22.745 | -22.745 | -23.1372 | -21.7287 | -23.1693 | -22.4489 | -26.0504 | -20.0481 | -22.3854 | -22.5583 | -22.2126 | -22.3855 | -22.2973 | -22.973 | -22.027
| 113 | -24.3137 | -21.9608 | -23.1372 | -23.5294 | -22.9292 | -23.4093 | -22.9292 | -25.5702 | -21.7287 | -22.9041 | -22.5584 | -22.9041 | -22.7312 | -22.1621 | -22.5676 | -22.1621
| 114 | -23.9215 | -23.5294 | -25.098 | -24.3137 | -22.9292 | -23.6495 | -22.449 | -26.5306 | -21.0084 | -22.2126 | -22.904 | -23.0769 | -22.5583 | -22.4324 | -22.7027 | -22.4324
| 115 | -23.5294 | -24.3137 | -22.353 | -23.5294 | -22.6891 | -23.1693 | -22.9291 | -25.5702 | -21.2485 | -22.7312 | -22.7312 | -22.9041 | -22.5583 | -22.4324 | -22.4324 | -22.1622
| 116 | -24.3138 | -24.3137 | -23.5294 | -23.1372 | -22.9291 | -22.6891 | -21.9688 | -26.2905 | -21.0084 | -22.5583 | -23.2498 | -22.2127 | -22.2126 | -21.8919 | -21.8919 | -22.1622
| 117 | -25.098 | -24.7059 | -22.7451 | -23.9215 | -23.6494 | -23.4094 | -21.7286 | -24.85 | -21.4886 | -22.2126 | -23.2498 | -21.6941 | -22.0398 | -22.027 | -22.4324 | -22.4325
| 118 | -23.9215 | -23.5294 | -22.353 | -53.1813 | -22.2089 | -23.1692 | -22.689 | -25.3302 | -21.2485 | -23.2498 | -23.0769 | -22.5583 | -22.0397 | -22.1621 | -22.7027 | -22.5676
| 119 | -22.353 | -23.1372 | -22.745 | -52.7011 | -22.9292 | -23.4094 | -22.449 | -25.0901 | -20.048 | -23.2498 | -22.7312 | -22.3854 | -23.5955 | -22.2973 | -22.973 | -22.7027
| 120 | -23.1372 | -22.745 | -23.5294 | -54.3818 | -22.4489 | -23.6495 | -23.6495 | -25.5703 | -21.0084 | -23.2498 | -22.0398 | -22.3855 | -22.2126 | -22.7027 | -22.4324 | -22.027
| 121 | -21.9607 | -22.745 | -22.745 | -55.8223 | -22.689 | -21.7287 | -22.2088 | -24.85 | -19.8079 | -22.5583 | -22.3855 | -22.2126 | -22.7312 | -22.027 | -22.1622 | -22.1622
| 122 | -23.5294 | -23.5294 | -23.5294 | -57.9832 | -22.2089 | -21.7287 | -22.9291 | -24.6099 | -20.5282 | -22.7312 | -22.3855 | -22.2127 | -22.904 | -22.7027 | -22.1622 | -22.1622
| 123 | -24.3137 | -23.9216 | -23.9215 | -58.9436 | -23.4094 | -21.7287 | -23.8895 | -24.3698 | -20.2881 | -22.9041 | -22.3854 | -22.9041 | -22.2127 | -22.5676 | -22.4325 | -22.4324
| 124 | -22.745 | -21.9608 | -23.5294 | -59.6639 | -23.8896 | -23.1693 | -22.9292 | -24.1297 | -21.2485 | -22.3855 | -23.5955 | -22.5584 | -23.4226 | -22.973 | -23.3784 | -22.027
| 125 | -22.353 | -22.745 | -23.5294 | -60.3842 | -22.689 | -23.4094 | -22.449 | -24.1297 | -20.5282 | -22.2126 | -22.5583 | -22.904 | -22.0397 | -22.7027 | -22.2973 | -22.1622
| 126 | -23.9215 | -22.745 | -22.3529 | -60.8643 | -22.9291 | -22.689 | -22.9292 | -23.8896 | -20.7683 | -22.5583 | -23.0769 | -22.7312 | -22.2126 | -23.2432 | -22.027 | -22.7027
| 127 | -25.098 | -23.1373 | -23.1372 | -58.9436 | -23.1693 | -23.1693 | -22.9291 | -24.1296 | -20.7683 | -23.0769 | -22.5583 | -22.3855 | -22.3855 | -22.1622 | -22.8379 | -22.7027
| 128 | -23.1372 | -23.5294 | -22.745 | -53.9015 | -21.9688 | -21.7287 | -22.9291 | -23.1693 | -20.048 | -22.0398 | -23.0769 | -22.7312 | -22.3855 | -22.5676 | -22.8379 | -22.2973
| 129 | -23.1372 | -21.9608 | -23.9215 | -47.4189 | -23.6495 | -23.4093 | -23.4094 | -23.8896 | -21.0084 | -23.0769 | -22.904 | -22.5583 | -22.5583 | -22.1622 | -22.2973 | -22.1622
| 130 | -23.5294 | -23.1372 | -23.9215 | -40.9364 | -21.9687 | -22.9292 | -21.9687 | -24.1297 | -21.0084 | -22.9041 | -22.7312 | -22.3854 | -22.5583 | -22.2973 | -22.4325 | -22.5676
| 131 | -25.098 | -21.5686 | -23.1372 | -38.5355 | -23.1692 | -21.7287 | -23.4093 | -23.8896 | -19.8079 | -22.9041 | -22.2126 | -22.2126 | -22.5583 | -22.4324 | -22.1621 | -22.1622
| 132 | -21.5686 | -23.9215 | -21.5686 | -34.6939 | -22.2088 | -22.9292 | -23.1693 | -22.449 | -20.2881 | -22.5583 | -23.0769 | -23.2498 | -22.0398 | -22.5676 | -22.027 | -22.4324
| 133 | -23.9215 | -23.9215 | -24.3137 | -28.9316 | -22.449 | -23.6495 | -23.6495 | -22.6891 | -21.7287 | -23.0769 | -23.4227 | -22.7312 | -23.0769 | -22.2973 | -22.2973 | -22.7027
| 134 | -23.1372 | -24.3137 | -21.1765 | -24.6099 | -23.6495 | -23.1692 | -23.1693 | -22.9291 | -21.9687 | -22.5583 | -23.0769 | -22.3854 | -22.7312 | -22.8379 | -22.8378 | -22.1622
| 135 | -23.9215 | -24.3137 | -20.7843 | -21.4886 | -23.4094 | -23.1692 | -21.7286 | -22.9291 | -20.5282 | -22.7312 | -23.0769 | -22.2126 | -22.7312 | -22.5676 | -22.4324 | -22.4325
| 136 | -22.745 | -22.745 | -21.9608 | -16.6866 | -22.6891 | -22.9292 | -22.9291 | -23.4093 | -19.808 | -22.3854 | -22.5583 | -22.7312 | -22.0398 | -22.2973 | -22.1621 | -22.2973
| 137 | -24.3137 | -22.745 | -23.1372 | -13.8055 | -22.6891 | -22.9291 | -22.449 | -22.449 | -21.0084 | -22.2126 | -22.5583 | -22.2126 | -22.904 | -22.0271 | -22.1622 | -22.973
| 138 | -24.7059 | -22.745 | -23.5294 | -12.1249 | -22.449 | -22.9291 | -22.689 | -22.449 | -20.7683 | -22.5583 | -23.4226 | -22.904 | -22.3854 | -22.4325 | -22.7027 | -22.2973
| 139 | -25.098 | -24.3137 | -23.1372 | -13.8056 | -22.689 | -23.4093 | -23.6495 | -22.9291 | -21.0084 | -22.7312 | -22.9041 | -22.3854 | -22.3855 | -22.7027 | -22.973 | -22.5675
| 140 | -23.9215 | -23.9215 | -23.5294 | -13.5654 | -22.6891 | -21.7287 | -23.1693 | -23.1692 | -20.048 | -22.2126 | -22.0398 | -22.3854 | -22.2126 | -22.7027 | -22.4325 | -21.7568
| 141 | -22.745 | -21.9607 | -23.1372 | -15.7263 | -23.4094 | -22.689 | -22.449 | -23.1692 | -20.048 | -22.2126 | -22.0397 | -22.3855 | -22.3855 | -21.4865 | -22.2973 | -22.2973
| 142 | -23.1373 | -24.7059 | -24.7059 | -18.3674 | -23.4094 | -23.6495 | -22.9291 | -23.4093 | -20.5282 | -22.5583 | -22.0397 | -23.0769 | -23.2498 | -22.2973 | -21.8919 | -22.4324
| 143 | -23.5295 | -24.7059 | -22.353 | -19.8079 | -23.4094 | -22.4489 | -23.4094 | -21.7287 | -21.9688 | -23.4227 | -22.9041 | -23.0769 | -22.3854 | -22.2973 | -22.4324 | -21.8919
| 144 | -24.7059 | -23.1372 | -23.9215 | -16.4466 | -22.6891 | -22.9291 | -22.449 | -22.449 | -21.2485 | -23.7684 | -23.4226 | -23.2498 | -23.0769 | -22.5676 | -22.7027 | -22.8378
| 145 | -23.5294 | -22.3529 | -22.745 | -15.9664 | -22.9292 | -22.689 | -23.6495 | -22.2088 | -20.7683 | -22.9041 | -23.0769 | -23.4226 | -22.904 | -22.4325 | -22.1621 | -22.7027
| 146 | -22.353 | -23.1372 | -23.1372 | -16.9268 | -21.9688 | -22.2089 | -23.8896 | -22.449 | -20.7683 | -21.8669 | -22.3855 | -22.2126 | -21.8669 | -22.2973 | -22.0271 | -22.2973
| 147 | -23.1372 | -23.5294 | -23.9215 | -18.6075 | -22.2089 | -22.9292 | -24.3698 | -22.2089 | -20.5282 | -22.7312 | -22.5583 | -22.5583 | -22.7312 | -22.2973 | -22.2973 | -22.8378
| 148 | -25.8824 | -23.5294 | -24.3137 | -19.0876 | -23.6495 | -22.9291 | -22.4489 | -21.7287 | -20.2881 | -22.7312 | -23.0769 | -23.0769 | -22.5583 | -22.5676 | -22.2973 | -22.973
| 149 | -24.7059 | -23.1372 | -22.7451 | -20.2881 | -23.1692 | -23.6495 | -23.1693 | -22.6891 | -21.0084 | -23.4226 | -22.3855 | -22.904 | -22.904 | -22.973 | -23.1081 | -22.7027
| 150 | -25.098 | -22.7451 | -24.7059 | -16.4466 | -22.2089 | -22.9291 | -22.6891 | -22.449 | -20.5282 | -22.904 | -22.7312 | -22.0398 | -22.904 | -22.4325 | -22.2973 | -22.0271
| 151 | -24.3137 | -24.3137 | -22.745 | -16.9268 | -22.449 | -21.9687 | -21.7287 | -21.9687 | -21.0084 | -22.3854 | -23.2498 | -23.2498 | -22.904 | -22.027 | -22.2973 | -22.1622
| 152 | -23.1372 | -21.9608 | -23.9215 | -18.8475 | -22.689 | -22.9292 | -22.6891 | -21.7287 | -21.0084 | -22.9041 | -22.3855 | -22.9041 | -22.3854 | -21.7567 | -22.4325 | -22.1621
| 153 | -23.5294 | -26.2745 | -23.1372 | -19.3277 | -23.6495 | -23.6495 | -22.689 | -22.449 | -21.0084 | -22.5583 | -22.904 | -22.9041 | -23.0769 | -21.8919 | -22.5676 | -22.1622
| 154 | -22.3529 | -22.353 | -22.7451 | -18.1273 | -23.4094 | -23.8896 | -25.0901 | -22.449 | -21.2485 | -22.0397 | -23.0769 | -23.4226 | -22.5583 | -22.7027 | -22.8378 | -22.8379
| 155 | -22.745 | -22.7451 | -23.1372 | -18.3674 | -23.1692 | -21.4886 | -23.1692 | -22.449 | -21.4885 | -23.4227 | -23.0769 | -22.7312 | -22.7312 | -22.2973 | -21.8919 | -22.5676
| 156 | -23.5294 | -23.1373 | -23.5294 | -19.5678 | -21.9687 | -22.689 | -21.4886 | -22.2088 | -21.0084 | -22.3854 | -23.0769 | -22.2126 | -22.2126 | -22.7027 | -21.7568 | -22.4324
| 157 | -25.4902 | -23.1373 | -24.3137 | -19.0876 | -21.7286 | -21.9687 | -22.9291 | -22.2088 | -20.7683 | -22.9041 | -22.9041 | -22.904 | -22.3855 | -21.8919 | -22.2973 | -22.1621
| 158 | -24.7059 | -24.3137 | -24.7059 | -18.8476 | -23.6495 | -23.1692 | -23.1693 | -22.4489 | -21.0084 | -23.0769 | -22.7312 | -23.2498 | -22.7312 | -21.6216 | -21.8919 | -22.5676
| 159 | -24.3138 | -23.5294 | -25.098 | -19.3277 | -24.1297 | -23.1692 | -23.8896 | -22.6891 | -21.0084 | -22.904 | -22.904 | -22.904 | -23.0769 | -23.1081 | -22.7027 | -22.1621
| 160 | -23.9216 | -23.9215 | -22.3529 | -19.808 | -23.1692 | -21.2485 | -22.9292 | -21.9688 | -20.7683 | -22.7312 | -24.1141 | -22.3855 | -22.7312 | -22.8379 | -22.2973 | -22.5676
| 161 | -24.7059 | -24.3137 | -23.1372 | -20.5282 | -23.1693 | -22.2089 | -23.1693 | -22.2089 | -21.2485 | -22.904 | -22.3855 | -22.5583 | -23.0769 | -22.7027 | -22.4325 | -22.027
| 162 | -23.1373 | -23.9215 | -25.4902 | -20.7683 | -23.6495 | -22.6891 | -22.9292 | -21.7287 | -19.5678 | -22.5583 | -22.2126 | -23.4226 | -23.2498 | -22.5676 | -22.4325 | -22.973
| 163 | -25.8824 | -25.4902 | -23.5294 | -20.048 | -23.1693 | -23.6495 | -21.4886 | -21.9687 | -20.048 | -23.2498 | -22.2126 | -23.7684 | -22.5583 | -22.8378 | -22.2973 | -22.4324
| 164 | -24.7059 | -23.1373 | -22.745 | -19.5679 | -22.689 | -22.9292 | -21.9688 | -21.7287 | -43.8202 | -22.3854 | -22.3854 | -23.2498 | -23.2498 | -23.1081 | -23.1081 | -22.1622
| 165 | -21.9608 | -23.9215 | -24.7059 | -19.5679 | -22.689 | -23.6495 | -23.4093 | -22.2089 | -42.4373 | -22.9041 | -22.5583 | -22.9041 | -22.9041 | -22.1622 | -22.8379 | -22.5676
| 166 | -22.745 | -23.5294 | -22.353 | -20.5282 | -22.4489 | -22.689 | -23.4093 | -22.2088 | -43.9931 | -22.5583 | -22.3855 | -22.3855 | -22.9041 | -22.1621 | -22.4324 | -22.4324
| 167 | -23.9215 | -22.3529 | -24.3137 | -19.5678 | -23.4094 | -23.1692 | -22.689 | -21.2485 | -45.2031 | -23.0769 | -22.2126 | -23.0769 | -23.0769 | -22.5675 | -22.2973 | -22.8379
| 168 | -25.098 | -23.9215 | -24.3137 | -21.0084 | -22.449 | -23.1692 | -22.4489 | -21.4886 | -47.4503 | -23.7684 | -23.0769 | -23.0769 | -22.3854 | -22.5676 | -21.7567 | -22.4325
| 169 | -22.353 | -24.3137 | -23.1372 | -19.5679 | -23.1692 | -23.6495 | -23.8896 | -21.9688 | -47.9688 | -22.7312 | -22.2126 | -23.2498 | -23.2498 | -22.973 | -22.7027 | -22.4324
| 170 | -24.3138 | -22.7451 | -23.1372 | -19.3278 | -23.8896 | -22.9292 | -22.449 | -21.7287 | -48.4875 | -22.5583 | -22.3855 | -23.5955 | -22.904 | -22.973 | -22.7027 | -22.5675
| 171 | -23.1373 | -24.3137 | -23.1372 | -21.0084 | -22.6891 | -23.4094 | -22.9292 | -20.7683 | -49.6975 | -21.8669 | -22.7312 | -22.9041 | -22.2126 | -21.8919 | -22.8379 | -22.027
| 172 | -24.3137 | -23.1372 | -23.1372 | -20.7683 | -22.9291 | -23.6495 | -22.9291 | -21.4886 | -49.6975 | -23.0769 | -22.904 | -22.904 | -23.9413 | -22.2973 | -22.4325 | -22.7027
| 173 | -23.1372 | -24.3137 | -23.5295 | -20.7683 | -23.4094 | -24.1296 | -23.1692 | -21.7286 | -48.3146 | -23.4226 | -22.5583 | -23.0769 | -22.7312 | -22.8379 | -23.5135 | -22.5676
| 174 | -24.7059 | -23.1372 | -23.1372 | -20.048 | -22.449 | -23.6495 | -23.1693 | -20.7683 | -44.1659 | -23.2498 | -22.904 | -22.904 | -23.5955 | -22.8379 | -22.973 | -23.1081
| 175 | -23.1372 | -21.9608 | -22.353 | -20.2881 | -23.6495 | -22.689 | -23.4094 | -20.2881 | -39.6716 | -22.0397 | -23.2498 | -22.7312 | -23.4227 | -22.2973 | -22.8378 | -22.2973
| 176 | -23.5294 | -21.9608 | -22.3529 | -20.048 | -24.1297 | -22.2089 | -22.2089 | -21.2485 | -35.1772 | -22.5583 | -22.5583 | -22.904 | -22.9041 | -22.2973 | -21.7568 | -22.1622
| 177 | -22.745 | -20.7843 | -23.9215 | -21.2485 | -21.9688 | -21.9687 | -22.2089 | -20.048 | -32.4114 | -22.9041 | -22.7312 | -23.5955 | -22.2126 | -22.7027 | -22.4324 | -22.4324
| 178 | -25.098 | -22.745 | -24.7059 | -20.7683 | -22.6891 | -23.1693 | -21.7287 | -21.7287 | -30.1642 | -22.7312 | -23.0769 | -23.2498 | -22.9041 | -22.5676 | -22.4324 | -22.7027
| 179 | -23.1372 | -23.9215 | -25.098 | -21.9687 | -22.449 | -23.8896 | -23.1692 | -21.2485 | -27.2256 | -22.5583 | -22.7312 | -23.4226 | -22.904 | -22.5675 | -22.4324 | -22.4324
| 180 | -21.9608 | -23.9215 | -25.8824 | -20.048 | -21.4886 | -24.6099 | -21.9687 | -21.4886 | -23.7684 | -22.5583 | -23.0769 | -23.4227 | -23.2498 | -22.0271 | -22.7027 | -22.973
| 181 | -23.5294 | -23.5294 | -23.1373 | -21.0084 | -22.9291 | -22.4489 | -22.4489 | -21.4886 | -21.1755 | -22.7312 | -22.5583 | -23.2498 | -22.7312 | -22.1621 | -22.4325 | -21.6216
| 182 | -24.3137 | -22.745 | -24.3137 | -21.4885 | -22.449 | -21.9687 | -22.6891 | -21.2485 | -18.2368 | -22.3855 | -23.0769 | -23.0769 | -22.904 | -22.5676 | -22.5675 | -22.027
| 183 | -22.745 | -22.353 | -23.5294 | -21.7286 | -21.9688 | -23.1693 | -22.6891 | -20.7683 | -15.9896 | -22.904 | -22.2126 | -23.9412 | -22.9041 | -22.4324 | -22.1622 | -22.8378
| 184 | -24.3137 | -25.098 | -22.3529 | -21.4886 | -23.4093 | -22.9292 | -22.4489 | -21.0084 | -13.9153 | -23.4226 | -22.5583 | -23.0769 | -23.0769 | -22.7027 | -23.1081 | -22.7027
| 185 | -25.4902 | -23.1372 | -22.353 | -22.449 | -23.4094 | -23.4093 | -23.1692 | -21.2485 | -14.9524 | -22.9041 | -22.9041 | -22.2126 | -22.2126 | -22.8378 | -22.1621 | -22.7027
| 186 | -22.7451 | -22.353 | -22.745 | -21.9688 | -23.4093 | -22.9291 | -23.8895 | -20.2881 | -15.9896 | -22.9041 | -22.5583 | -22.2126 | -22.5583 | -22.7027 | -22.1621 | -22.4324
| 187 | -23.5294 | -23.5294 | -24.7059 | -20.5282 | -23.6495 | -22.9291 | -23.6495 | -20.048 | -17.8911 | -22.7312 | -23.7684 | -23.0769 | -22.5583 | -22.4324 | -22.027 | -22.4324
| 188 | -25.4902 | -25.8824 | -25.098 | -21.7287 | -23.1692 | -22.6891 | -23.1692 | -21.0084 | -19.274 | -23.5955 | -23.4227 | -23.0769 | -23.0769 | -22.2973 | -22.1622 | -22.2973
| 189 | -24.3137 | -23.1372 | -25.098 | -22.2089 | -23.6495 | -22.2088 | -22.4489 | -21.0084 | -19.274 | -22.5583 | -22.7312 | -23.0769 | -23.0769 | -22.4325 | -22.4324 | -22.4325
| 190 | -24.7059 | -23.1372 | -23.1373 | -22.2089 | -22.9291 | -22.6891 | -23.1693 | -21.0084 | -18.5825 | -23.2498 | -22.5583 | -22.5583 | -22.3855 | -22.4324 | -22.4324 | -22.1622
| 191 | -23.9215 | -23.9215 | -23.9216 | -20.7683 | -22.2089 | -22.449 | -22.2088 | -20.2881 | -17.1996 | -22.3854 | -22.2126 | -22.7312 | -22.3855 | -22.973 | -22.5676 | -22.027
| 192 | -22.353 | -21.9608 | -23.5294 | -22.2089 | -22.9292 | -22.4489 | -23.6495 | -20.7683 | -17.3725 | -23.7684 | -22.3855 | -22.7312 | -22.9041 | -22.7027 | -22.5676 | -21.8919
| 193 | -23.9216 | -22.7451 | -23.5294 | -22.9292 | -22.689 | -23.8896 | -22.9291 | -21.0084 | -19.2739 | -23.2498 | -21.8669 | -22.3855 | -22.904 | -23.2432 | -21.8919 | -21.6216
| 194 | -24.3138 | -22.7451 | -23.5294 | -22.2089 | -22.689 | -23.6495 | -22.9291 | -21.4886 | -20.1383 | -22.904 | -22.7312 | -22.2126 | -23.0769 | -23.1081 | -22.1622 | -22.5676
| 195 | -24.7059 | -25.4901 | -21.9608 | -22.2088 | -22.6891 | -23.1692 | -22.449 | -20.2881 | -19.274 | -22.7312 | -22.904 | -22.9041 | -23.2498 | -22.7027 | -22.4325 | -21.8919

## SavedData
[5 empty] | SS | Total CCN | Total CN |  | Max. T | Dp50 | Dp50 | Dp50+20 | B | D | c
DataSource | Meas Date | Start Time | End Time | Shift | (%) | Conc | Conc | CCN/CN | Error (%) | (dry) | (wet) | (wet) | (sigmd) | (sigmd) | (sigmd)
Data!R | 2010-02-03 | 10:25 | N/A | 8 | 0.8 | 142528.8793 | 218140.0657 | 0.65338 | 23.514 | 27.4516 | 4.6306 | 4.7403 | 0.97034 | 27.4516 | -16.8245
Data!Q | 2010-02-03 | 10:22 | 10:25 | 8 | 0.8 | 153798.7659 | 226619.8644 | 0.67866 | 23.514 | 27.2182 | 4.4406 | 4.6702 | 0.98073 | 27.2182 | -18.5414
Data!P | 2010-02-03 | 10:20 | 10:22 | 8 | 0.8 | 155238.1462 | 225650.8546 | 0.68796 | 23.243 | 26.6417 | 4.3037 | 4.5597 | 0.9793 | 26.6417 | -17.3094
Data!O | 2010-02-03 | 10:16 | 10:18 | 8 | 0.6 | 129594.7119 | 227293.1887 | 0.57017 | 23.941 | 33.6428 | 3.2738 | 3.5431 | 1.0027 | 33.6428 | -18.7396
Data!N | 2010-02-03 | 10:13 | 10:16 | 8 | 0.6 | 129638.6402 | 227891.9952 | 0.56886 | 24.114 | 33.6904 | 3.2291 | 3.5099 | 1.011 | 33.6904 | -17.0254
Data!M | 2010-02-03 | 10:11 | 10:13 | 8 | 0.6 | 122534.867 | 220001.4839 | 0.55697 | 24.114 | 33.462 | 3.2094 | 3.5736 | 1.0015 | 33.462 | -17.5064
Data!L | 2010-02-03 | 10:09 | 10:11 | 8 | 0.6 | 132964.45 | 229458.1053 | 0.57947 | 23.941 | 33.2758 | 3.1643 | 3.4453 | 1.0243 | 33.2758 | -15.7142
Data!I | 2010-02-03 | 10:00 | 10:02 | 9 | 0.4 | 97838.2955 | 228518.3784 | 0.42814 | 25.09 | 42.1052 | 1.9472 | 2.2745 | 0.98932 | 42.1052 | -23.0523
Data!H | 2010-02-03 | 9:58 | 10:00 | 9 | 0.4 | 91965.6473 | 217921.2803 | 0.42201 | 24.61 | 41.5917 | 1.9499 | 2.1449 | 0.99831 | 41.5917 | -21.6199
Data!E | 2010-02-03 | 9:51 | 9:53 | 9 | 0.2 | 40448.8435 | 228135.7432 | 0.1773 | 26.275 | 65.8919 | 1.1343 | 1.322 | 0.97402 | 65.8919 | -14.9328
Data!D | 2010-02-03 | 9:49 | 9:51 | 9 | 0.2 | 37384.7791 | 218361.5177 | 0.17121 | 26.274 | 64.1593 | 1.1344 | 1.3326 | 0.90008 | 64.1593 | -15.5948
Data!C | 2010-02-03 | 9:46 | 9:49 | 9 | 0.2 | 39542.1851 | 226569.8973 | 0.17453 | 26.275 | 66.1951 | 1.1238 | 1.357 | 0.96463 | 66.1951 | -15.3936

## Summary
Aerosol Thermodynamic Properties and Köhler
Theory Parameters (Copy and Paste From Tables
on Aerosol_Props Worksheet for Different Solutes):
Solute - Ammonium Sulfate
A |  | 2.0311e-09 | m
Density |  | 1770 | kg/m3
Molecular Weight |  | 0.13214 | kg/mol
Van't Hoff factor: |  | 2.5 | (nominal)
| a0 | 1.8853 | VHF(Dp) = a0 + a1*Dp
| a1 | 0.024642 | + a2*Dp^2 +
| a2 | -0.00036942 | a3*Dp^3 +
| a3 | 2.8686e-06 | a4*Dp^4
| a4 | -8.5855e-09
Dynamic Shape Factor |  | 1
Solvent - Water
Density |  | 994.323 | kg/m3
Molecular Weight |  | 0.01802 | kg/mol
Surface Tension |  | 0.070581 | N/m
Ideal Gas Constant |  | 8.314 | J/mol*K | Data Summary:
Temperature |  | 303 | K
DMA |  |  |  | Data Input |  |  |  |  | Calculated Parameters | [6 empty] | Theoretical Supersaturation Uncertainty Calculations
Temperature |  | 298.15 | K | %SS nominal | dTset (°C) | dTread (°C) | Dp avg (nm) | Dp stdev (nm) | Dp corrected (nm) | Pitzer VHF | ns | B | %SS theory | %SS Uncertainty + | %SS Uncertainty - | Dp + corrected (nm) | Pitzer VHF + | ns + | B + | %SS theory + | Dp - corrected (nm) | Pitzer VHF - | ns - | B - | %SS theory -
Pressure |  | 1000 | mb | [14 empty] | ns Plus | B Plus | %SS_theory Plus |  |  | ns Plus | B Plus | %SS_theory Plus
Qsheath |  | 10 | LPM | [14 empty] | ns Plus | B Plus | %SS_theory Plus |  |  | ns Plus | B Plus | %SS_theory Plus
Qaerosol |  | 1 | LPM | 0.2 | 5.1 | 4.44 | 65.4154 | 1.0984 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
R1 |  | 0.937 | cm | 0.4 | 8.33 | 7.3 | 41.8484 | 0.36307 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
R2 |  | 1.961 | cm | 0.6 | 11.57 | 10.09 | 33.5177 | 0.18894 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
Length |  | 43.6 | cm | 0.8 | 14.8 | 12.95 | 27.1038 | 0.4169 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
Gas Const, R |  | 8.314 | J/mol-K
Molar Mass of Air |  | 28.84 | g/mol
e |  | 1.6022e-19 | coulomb
Air Viscosity |  | 1.8372e-05 | Pa-s
Air Mean Free Path |  | 6.7507e-08 | m

## Notes
Notes:
This demonstration SMCA processor contains a subset of the calibration data obtained on the Georgia Tech CCNC on 3 February 2010.
Ammonium sulfate aerosol was generated via atomization and subsequent drying with two diffusion dryers.
The aerosol was charge-neutralized with Kr-85 ionizing radiation sources before being introduced into a TSI SMPS
consisting of a 3085 nano-Differential Mobility Analyzer and a 3022A Condensation Particle Counter
16 SMCA scans are included in this workbook, of which, 12 scans were analyzed and 4 scans were discarded during analysis.
Köhler theory was used to produce a 4-point calibration curve of %SS versus DT, which was used to calibrate the instrument %SS set points.
Data reanalysis performed on 30 July 2010, RHM

## ShiftCalculatorSmooth
Formulas for Data Alignment and Sigmoidal Fit: Modification Could Result In Errors
[19 empty] | Calculaing CCN |  |  |  |  | CCN/CN sigmoid
Diameter | CCN | CN | CCN/CN | 3 pt Run Avg | DropDiameter | 3 pt Run Avg | Scantime | SMPS | CCN Scantime (Adj) | Raw CCN | 3ptRunAvg | 5ptRunAvg | Raw Droplet Size | 3ptRunAvg | 5ptRunAvg | T3-T2 %Error | Diameter | dNdlogDp | Normalized | dNdlogDp | delta logDp | CCN | CN | Diameter | CCN/CN | CCN/CN simulated | error^2 | Asymptote
#REF! | 0 | #REF! | #REF! |  | #REF! |  | 1 | #REF! | -7 | #REF! |  |  | #REF! |  |  | #REF! | #REF! | #REF! | #REF! |  |  |  |  | #REF! | 456884998.35 |  | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 2 | #REF! | -6 | #REF! | #REF! |  | #REF! | #REF! |  | #REF! | #REF! | #REF! | #REF! |  |  |  |  | #REF! | 162194998.9 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 3 | #REF! | -5 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  |  | #REF! | 0.096563 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 4 | #REF! | -4 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  |  | #REF! | 0.069355 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 5 | #REF! | -3 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  |  | #REF! | 0.10604 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 6 | #REF! | -2 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  |  | #REF! | 0.11519 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 7 | #REF! | -1 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0.18897 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 8 | #REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 3.5948e-08 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 9 | #REF! | 1 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 3.2353e-08 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 10 | #REF! | 2 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 5.6122e-08 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 11 | #REF! | 3 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 5.0926e-08 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 12 | #REF! | 4 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 5.9783e-08 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 13 | #REF! | 5 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 5.2885e-08 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 14 | #REF! | 6 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 8.871e-08 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 15 | #REF! | 7 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 7.3333e-08 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 16 | #REF! | 8 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1.0185e-07 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 17 | #REF! | 9 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 6.962e-08 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 18 | #REF! | 10 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 8.5937e-08 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 19 | #REF! | 11 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1.0577e-07 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 20 | #REF! | 12 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1.25e-07 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 21 | #REF! | 13 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 7.7465e-08 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 22 | #REF! | 14 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1.3415e-07 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 23 | #REF! | 15 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1.3095e-07 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 24 | #REF! | 16 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 9.0164e-08 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 25 | #REF! | 17 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1.0185e-07 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 26 | #REF! | 18 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1.2791e-07 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 27 | #REF! | 19 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1.1224e-07 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 28 | #REF! | 20 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1.4474e-07 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 29 | #REF! | 21 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1.1e-07 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 30 | #REF! | 22 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1.0377e-07 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 31 | #REF! | 23 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1.4865e-07 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 32 | #REF! | 24 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1.4865e-07 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 33 | #REF! | 25 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1.1702e-07 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 34 | #REF! | 26 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1e-07 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 35 | #REF! | 27 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1.0185e-07 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 36 | #REF! | 28 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 9.0164e-08 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 37 | #REF! | 29 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 9.1667e-08 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 38 | #REF! | 30 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 7.1429e-08 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 39 | #REF! | 31 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 6.875e-08 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 40 | #REF! | 32 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 5.914e-08 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 41 | #REF! | 33 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 5.2381e-08 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 42 | #REF! | 34 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 3.9568e-08 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 43 | #REF! | 35 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 3.4375e-08 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 44 | #REF! | 36 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 2.6699e-08 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 45 | #REF! | 37 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 2.9255e-08 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 46 | #REF! | 38 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 2.2358e-08 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 47 | #REF! | 39 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1.8272e-08 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 48 | #REF! | 40 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1.5406e-08 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 49 | #REF! | 41 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1.3382e-08 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 50 | #REF! | 42 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1.1044e-08 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 51 | #REF! | 43 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1.1224e-08 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 52 | #REF! | 44 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 9.3857e-09 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 53 | #REF! | 45 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 7.3333e-09 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 54 | #REF! | 46 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 6.4935e-09 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 55 | #REF! | 47 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 6.0373e-09 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 56 | #REF! | 48 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 5.8201e-09 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 57 | #REF! | 49 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 4.717e-09 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 58 | #REF! | 50 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 4.1857e-09 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 59 | #REF! | 51 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 3.7187e-09 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 60 | #REF! | 52 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 3.2506e-09 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 61 | #REF! | 53 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 3.0726e-09 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 62 | #REF! | 54 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 2.7214e-09 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 63 | #REF! | 55 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 2.3534e-09 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 64 | #REF! | 56 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 2.0984e-09 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 65 | #REF! | 57 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1.9312e-09 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 66 | #REF! | 58 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1.7921e-09 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 67 | #REF! | 59 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1.6134e-09 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 68 | #REF! | 60 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1.4946e-09 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 69 | #REF! | 61 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1.3504e-09 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 70 | #REF! | 62 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 1.241e-09 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 71 | #REF! | 63 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0.011585 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 72 | #REF! | 64 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0.0078681 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 73 | #REF! | 65 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0.015492 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 74 | #REF! | 66 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0.021941 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 75 | #REF! | 67 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0.01177 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 76 | #REF! | 68 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0.014404 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 77 | #REF! | 69 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0.019084 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 78 | #REF! | 70 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0.013702 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 79 | #REF! | 71 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0.020365 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 80 | #REF! | 72 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0.027408 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 81 | #REF! | 73 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0.065466 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 82 | #REF! | 74 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0.20804 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 83 | #REF! | 75 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0.21931 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 84 | #REF! | 76 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0.41233 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 85 | #REF! | 77 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0.57254 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 86 | #REF! | 78 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0.72888 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 87 | #REF! | 79 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0.80862 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 88 | #REF! | 80 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0.84473 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 89 | #REF! | 81 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0.85578 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 90 | #REF! | 82 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0.87569 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 91 | #REF! | 83 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0.93125 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 92 | #REF! | 84 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0.93248 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 93 | #REF! | 85 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0.90568 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 94 | #REF! | 86 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0.98417 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 95 | #REF! | 87 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0.97106 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 96 | #REF! | 88 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0.90127 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 97 | #REF! | 89 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0.95739 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 98 | #REF! | 90 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0.92742 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 99 | #REF! | 91 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0.96502 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 100 | #REF! | 92 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0.96266 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 101 | #REF! | 93 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 0.97824 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 102 | #REF! | 94 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 0.95661 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 103 | #REF! | 95 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 0.96159 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 104 | #REF! | 96 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 1.007 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 105 | #REF! | 97 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 0.95517 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 106 | #REF! | 98 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 1.0251 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 107 | #REF! | 99 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 0.96415 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 108 | #REF! | 100 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 1.0297 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 109 | #REF! | 101 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 0.99503 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 110 | #REF! | 102 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 1.0224 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 111 | #REF! | 103 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 1.0426 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 112 | #REF! | 104 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 1.015 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 113 | #REF! | 105 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 1.0056 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 114 | #REF! | 106 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 1.0378 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 115 | #REF! | 107 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 1.0794 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 116 | #REF! | 108 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 1.0156 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 117 | #REF! | 109 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 1.023 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 118 | #REF! | 110 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 0.97342 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 119 | #REF! | 111 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 1.0227 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 120 | #REF! | 112 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 0.98103 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 121 | #REF! | 113 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 0.95699 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 122 | #REF! | 114 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 1.0309 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 123 | #REF! | 115 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 1.0694 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 124 | #REF! | 116 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 1.0368 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 125 | #REF! | 117 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 1.0061 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 126 | #REF! | 118 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 1.0387 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 127 | #REF! | 119 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 0.95412 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 128 | #REF! | 120 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 0.78143 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 129 | #REF! | 121 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 0.79016 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 130 | #REF! | 122 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 0.82611 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 131 | #REF! | 123 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 0.94181 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 132 | #REF! | 124 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 1.065 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 133 | #REF! | 125 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 0.75008 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! | #REF! | #REF! | #REF! | 134 | #REF! | 126 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 0.418 | #REF! | #REF! | #REF!
#REF! | 0 | #REF! | #REF! |  | #REF! |  | 135 | #REF! | 127 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | 0.28221 | #REF! | #REF! | #REF!
[7 empty] | 136 | 1e-06 | 128 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 137 | 1e-06 | 129 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 138 | 1e-06 | 130 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 139 | 1e-06 | 131 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 140 | 1e-06 | 132 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 141 | 1e-06 | 133 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 142 | 1e-06 | 134 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 143 | 1e-06 | 135 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 144 | 1e-06 | 136 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 145 | 1e-06 | 137 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 146 | 1e-06 | 138 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 147 | 1e-06 | 139 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 148 | 1e-06 | 140 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 149 | 1e-06 | 141 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 150 | 1e-06 | 142 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 151 | 1e-06 | 143 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 152 | 1e-06 | 144 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 153 | 1e-06 | 145 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 154 | 1e-06 | 146 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 155 | 1e-06 | 147 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 156 | 1e-06 | 148 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 157 | 1e-06 | 149 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 158 | 1e-06 | 150 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 159 | 1e-06 | 151 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 160 | 1e-06 | 152 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 161 | 1e-06 | 153 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 162 | 1e-06 | 154 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 163 | 1e-06 | 155 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 164 | 1e-06 | 156 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 165 | 1e-06 | 157 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 166 | 1e-06 | 158 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 167 | 1e-06 | 159 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 168 | 1e-06 | 160 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 169 | 1e-06 | 161 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 170 | 1e-06 | 162 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 171 | 1e-06 | 163 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 172 | 1e-06 | 164 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 173 | 1e-06 | 165 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 174 | 1e-06 | 166 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 175 | 1e-06 | 167 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 176 | 1e-06 | 168 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 177 | 1e-06 | 169 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 178 | 1e-06 | 170 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 179 | 1e-06 | 171 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 180 | 1e-06 | 172 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 181 | 1e-06 | 173 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 182 | 1e-06 | 174 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 183 | 1e-06 | 175 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 184 | 1e-06 | 176 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 185 | 1e-06 | 177 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 186 | 1e-06 | 178 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 187 | 1e-06 | 179 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 188 | 1e-06 | 180 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 189 | 1e-06 | 181 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 190 | 1e-06 | 182 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 191 | 1e-06 | 183 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 192 | 1e-06 | 184 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 193 | 1e-06 | 185 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 194 | 1e-06 | 186 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 195 | 1e-06 | 187 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 196 | 1e-06 | 188 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 197 | 1e-06 | 189 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 198 | 1e-06 | 190 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 199 | 1e-06 | 191 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 200 | 1e-06 | 192 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 201 | 1e-06 | 193 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 202 | 1e-06 | 194 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 203 | 1e-06 | 195 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 204 | 1e-06 | 196 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 205 | 1e-06 | 197 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 206 | 1e-06 | 198 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
[7 empty] | 207 | 1e-06 | 199 | #REF! | #REF! | #REF! | #REF! | #REF! |  | #REF!
[7 empty] | 208 | 1e-06 | 200 | #REF! | #REF! |  | #REF! | #REF! |  | #REF!
[7 empty] | 209 | 1e-06 | 201 | #REF! |  |  | #REF! |  |  | #REF!

## ChargeCorrection
Formulas to Correct for Multiple Charging: Modification Could Result In Errors
Wiedensohler, 1988 Coefficients: |  |  |  |  | Gunn, 1956 Coefficients: |  |  | Variables
a0 = | -0.0003 | -2.3484 | -44.4756 |  | e = | 1.6022e-19 | coulomb | T = | 298.15 |  |  | K
a1 = | -0.1014 | 0.6044 | 79.3772 |  | e0 = | 8.8542e-12 | farad/m | P = | 1013 |  |  | mb
a2 = | 0.3073 | 0.48 | -62.89 |  | k = | 1.3807e-23 | J/K
a3 = | -0.3372 | 0.0013 | 26.4492 |  | T = | 298.15 | K | ηair = | 1.8372e-05 |  |  | Pa-s | (Hinds, 1999)
a4 = | 0.1023 | -0.1553 | -5.748 |  | z_+/z_- = | 0.875 |  | λair = | 6.6641e-08 |  |  | m | (Seinfeld and Pandis, 2006)
a5 = | -0.0105 | 0.032 | 0.5049
n = | 0 | 1 | 2 | 3
Diameter (nm) | +0 Fraction | +1 Fraction | +2 Fraction | +3 Fraction | +1 Diameter (nm) | +2 Diameter (nm) | +3 Diameter (nm) | Raw CCN | Raw CN | Previous Iteration (gCN) | Previous Iteration (gCCN) | CN (corrected) | CCN (corrected) | Activated Ratio (Corrected)
#REF! | #REF! | #REF! | #REF! | #REF! | 3.1331 | 4.4408 | 5.446 | 0 | #REF! | 1.8182e-07 | 83.07 | 1.8182e-07 | 83.07 | 456884998.35
#REF! | #REF! | #REF! | #REF! | #REF! | 3.2247 | 4.5704 | 5.6051 | 0 | #REF! | 1.8182e-07 | 29.49 | 1.8182e-07 | 29.49 | 162194998.9
#REF! | #REF! | #REF! | #REF! | #REF! | 3.3189 | 4.7036 | 5.7687 | 0 | #REF! | 175.6364 | 16.96 | 175.6364 | 16.96 | 0.096563
#REF! | #REF! | #REF! | #REF! | #REF! | 3.4158 | 4.8409 | 5.9373 | 0 | #REF! | 115.6364 | 8.02 | 115.6364 | 8.02 | 0.069355
#REF! | #REF! | #REF! | #REF! | #REF! | 3.5155 | 4.9822 | 6.1108 | 0 | #REF! | 75.8182 | 8.04 | 75.8182 | 8.04 | 0.10604
#REF! | #REF! | #REF! | #REF! | #REF! | 3.6182 | 5.1276 | 6.2895 | 0 | #REF! | 54 | 6.22 | 54 | 6.22 | 0.11519
#REF! | #REF! | #REF! | #REF! | #REF! | 3.7237 | 5.2772 | 6.4733 | 0 | #REF! | 42.5455 | 8.04 | 42.5455 | 8.04 | 0.18897
#REF! | #REF! | #REF! | #REF! | #REF! | 3.8323 | 5.4313 | 6.6626 | 0 | #REF! | 27.8182 | 1e-06 | 27.8182 | 1e-06 | 3.5948e-08
#REF! | #REF! | #REF! | #REF! | #REF! | 3.9441 | 5.59 | 6.8576 | 0 | #REF! | 30.9091 | 1e-06 | 30.9091 | 1e-06 | 3.2353e-08
#REF! | #REF! | #REF! | #REF! | #REF! | 4.059 | 5.7532 | 7.0581 | 0 | #REF! | 17.8182 | 1e-06 | 17.8182 | 1e-06 | 5.6122e-08
#REF! | #REF! | #REF! | #REF! | #REF! | 4.1774 | 5.9213 | 7.2647 | 0 | #REF! | 19.6364 | 1e-06 | 19.6364 | 1e-06 | 5.0926e-08
#REF! | #REF! | #REF! | #REF! | #REF! | 4.2992 | 6.0942 | 7.4773 | 0 | #REF! | 16.7273 | 1e-06 | 16.7273 | 1e-06 | 5.9783e-08
#REF! | #REF! | #REF! | #REF! | #REF! | 4.4246 | 6.2724 | 7.6963 | 0 | #REF! | 18.9091 | 1e-06 | 18.9091 | 1e-06 | 5.2885e-08
#REF! | #REF! | #REF! | #REF! | #REF! | 4.5536 | 6.4557 | 7.9217 | 0 | #REF! | 11.2727 | 1e-06 | 11.2727 | 1e-06 | 8.871e-08
#REF! | #REF! | #REF! | #REF! | #REF! | 4.6864 | 6.6444 | 8.1537 | 0 | #REF! | 13.6364 | 1e-06 | 13.6364 | 1e-06 | 7.3333e-08
#REF! | #REF! | #REF! | #REF! | #REF! | 4.8231 | 6.8387 | 8.3927 | 0 | #REF! | 9.8182 | 1e-06 | 9.8182 | 1e-06 | 1.0185e-07
#REF! | #REF! | #REF! | #REF! | #REF! | 4.9637 | 7.0387 | 8.6386 | 0 | #REF! | 14.3636 | 1e-06 | 14.3636 | 1e-06 | 6.962e-08
#REF! | #REF! | #REF! | #REF! | #REF! | 5.1086 | 7.2447 | 8.892 | 0 | #REF! | 11.6364 | 1e-06 | 11.6364 | 1e-06 | 8.5937e-08
#REF! | #REF! | #REF! | #REF! | #REF! | 5.2576 | 7.4567 | 9.1528 | 0 | #REF! | 9.4545 | 1e-06 | 9.4545 | 1e-06 | 1.0577e-07
#REF! | #REF! | #REF! | #REF! | #REF! | 5.4111 | 7.675 | 9.4213 | 0 | #REF! | 8 | 1e-06 | 8 | 1e-06 | 1.25e-07
#REF! | #REF! | #REF! | #REF! | #REF! | 5.569 | 7.8997 | 9.6979 | 0 | #REF! | 12.9091 | 1e-06 | 12.9091 | 1e-06 | 7.7465e-08
#REF! | #REF! | #REF! | #REF! | #REF! | 5.7315 | 8.1309 | 9.9825 | 0 | #REF! | 7.4545 | 1e-06 | 7.4545 | 1e-06 | 1.3415e-07
#REF! | #REF! | #REF! | #REF! | #REF! | 5.8989 | 8.3692 | 10.2758 | 0 | #REF! | 7.6364 | 1e-06 | 7.6364 | 1e-06 | 1.3095e-07
#REF! | #REF! | #REF! | #REF! | #REF! | 6.0711 | 8.6143 | 10.5776 | 0 | #REF! | 11.0909 | 1e-06 | 11.0909 | 1e-06 | 9.0164e-08
#REF! | #REF! | #REF! | #REF! | #REF! | 6.2485 | 8.8668 | 10.8885 | 0 | #REF! | 9.8182 | 1e-06 | 9.8182 | 1e-06 | 1.0185e-07
#REF! | #REF! | #REF! | #REF! | #REF! | 6.431 | 9.1268 | 11.2086 | 0 | #REF! | 7.8182 | 1e-06 | 7.8182 | 1e-06 | 1.2791e-07
#REF! | #REF! | #REF! | #REF! | #REF! | 6.619 | 9.3946 | 11.5384 | 0 | #REF! | 8.9091 | 1e-06 | 8.9091 | 1e-06 | 1.1224e-07
#REF! | #REF! | #REF! | #REF! | #REF! | 6.8124 | 9.6702 | 11.878 | 0 | #REF! | 6.9091 | 1e-06 | 6.9091 | 1e-06 | 1.4474e-07
#REF! | #REF! | #REF! | #REF! | #REF! | 7.0115 | 9.9539 | 12.2276 | 0 | #REF! | 9.0909 | 1e-06 | 9.0909 | 1e-06 | 1.1e-07
#REF! | #REF! | #REF! | #REF! | #REF! | 7.2166 | 10.2462 | 12.5878 | 0 | #REF! | 9.6364 | 1e-06 | 9.6364 | 1e-06 | 1.0377e-07
#REF! | #REF! | #REF! | #REF! | #REF! | 7.4275 | 10.5471 | 12.9586 | 0 | #REF! | 6.7273 | 1e-06 | 6.7273 | 1e-06 | 1.4865e-07
#REF! | #REF! | #REF! | #REF! | #REF! | 7.6449 | 10.857 | 13.3408 | 0 | #REF! | 6.7273 | 1e-06 | 6.7273 | 1e-06 | 1.4865e-07
#REF! | #REF! | #REF! | #REF! | #REF! | 7.8686 | 11.1761 | 13.7343 | 0 | #REF! | 8.5455 | 1e-06 | 8.5455 | 1e-06 | 1.1702e-07
#REF! | #REF! | #REF! | #REF! | #REF! | 8.0988 | 11.5048 | 14.1396 | 0 | #REF! | 10 | 1e-06 | 10 | 1e-06 | 1e-07
#REF! | #REF! | #REF! | #REF! | #REF! | 8.3358 | 11.8431 | 14.557 | 0 | #REF! | 9.8182 | 1e-06 | 9.8182 | 1e-06 | 1.0185e-07
#REF! | #REF! | #REF! | #REF! | #REF! | 8.5799 | 12.1916 | 14.987 | 0 | #REF! | 11.0909 | 1e-06 | 11.0909 | 1e-06 | 9.0164e-08
#REF! | #REF! | #REF! | #REF! | #REF! | 8.8313 | 12.5506 | 15.4301 | 0 | #REF! | 10.9091 | 1e-06 | 10.9091 | 1e-06 | 9.1667e-08
#REF! | #REF! | #REF! | #REF! | #REF! | 9.09 | 12.9201 | 15.8864 | 0 | #REF! | 14 | 1e-06 | 14 | 1e-06 | 7.1429e-08
#REF! | #REF! | #REF! | #REF! | #REF! | 9.3565 | 13.301 | 16.3567 | 0 | #REF! | 14.5455 | 1e-06 | 14.5455 | 1e-06 | 6.875e-08
#REF! | #REF! | #REF! | #REF! | #REF! | 9.6307 | 13.693 | 16.8408 | 0 | #REF! | 16.9091 | 1e-06 | 16.9091 | 1e-06 | 5.914e-08
#REF! | #REF! | #REF! | #REF! | #REF! | 9.9131 | 14.0969 | 17.3399 | 0 | #REF! | 19.0909 | 1e-06 | 19.0909 | 1e-06 | 5.2381e-08
#REF! | #REF! | #REF! | #REF! | #REF! | 10.2039 | 14.5129 | 17.854 | 0 | #REF! | 25.2727 | 1e-06 | 25.2727 | 1e-06 | 3.9568e-08
#REF! | #REF! | #REF! | #REF! | #REF! | 10.5033 | 14.9414 | 18.3838 | 0 | #REF! | 29.0909 | 1e-06 | 29.0909 | 1e-06 | 3.4375e-08
#REF! | #REF! | #REF! | #REF! | #REF! | 10.8116 | 15.3828 | 18.9296 | 0 | #REF! | 37.4545 | 1e-06 | 37.4545 | 1e-06 | 2.6699e-08
#REF! | #REF! | #REF! | #REF! | #REF! | 11.1292 | 15.8376 | 19.4923 | 0 | #REF! | 34.1818 | 1e-06 | 34.1818 | 1e-06 | 2.9255e-08
#REF! | #REF! | #REF! | #REF! | #REF! | 11.4561 | 16.3059 | 20.0717 | 0 | #REF! | 44.7273 | 1e-06 | 44.7273 | 1e-06 | 2.2358e-08
#REF! | #REF! | #REF! | #REF! | #REF! | 11.7927 | 16.7885 | 20.669 | 0 | #REF! | 54.7273 | 1e-06 | 54.7273 | 1e-06 | 1.8272e-08
#REF! | #REF! | #REF! | #REF! | #REF! | 12.1395 | 17.2858 | 21.2848 | 0 | #REF! | 64.9091 | 1e-06 | 64.9091 | 1e-06 | 1.5406e-08
#REF! | #REF! | #REF! | #REF! | #REF! | 12.4964 | 17.7978 | 21.9191 | 0 | #REF! | 74.7273 | 1e-06 | 74.7273 | 1e-06 | 1.3382e-08
#REF! | #REF! | #REF! | #REF! | #REF! | 12.8643 | 18.3257 | 22.5732 | 0 | #REF! | 90.5455 | 1e-06 | 90.5455 | 1e-06 | 1.1044e-08
#REF! | #REF! | #REF! | #REF! | #REF! | 13.243 | 18.8695 | 23.2473 | 0 | #REF! | 89.0909 | 1e-06 | 89.0909 | 1e-06 | 1.1224e-08
#REF! | #REF! | #REF! | #REF! | #REF! | 13.633 | 19.4298 | 23.9421 | 0 | #REF! | 106.5455 | 1e-06 | 106.5455 | 1e-06 | 9.3857e-09
#REF! | #REF! | #REF! | #REF! | #REF! | 14.0346 | 20.0072 | 24.6584 | 0 | #REF! | 136.3636 | 1e-06 | 136.3636 | 1e-06 | 7.3333e-09
#REF! | #REF! | #REF! | #REF! | #REF! | 14.4482 | 20.6019 | 25.3966 | 0 | #REF! | 154 | 1e-06 | 154 | 1e-06 | 6.4935e-09
#REF! | #REF! | #REF! | #REF! | #REF! | 14.8745 | 21.2153 | 26.1582 | 0 | #REF! | 165.6364 | 1e-06 | 165.6364 | 1e-06 | 6.0373e-09
#REF! | #REF! | #REF! | #REF! | #REF! | 15.3134 | 21.8472 | 26.9432 | 0 | #REF! | 171.8182 | 1e-06 | 171.8182 | 1e-06 | 5.8201e-09
#REF! | #REF! | #REF! | #REF! | #REF! | 15.7655 | 22.4985 | 27.7526 | 0 | #REF! | 212 | 1e-06 | 212 | 1e-06 | 4.717e-09
#REF! | #REF! | #REF! | #REF! | #REF! | 16.2312 | 23.1698 | 28.5873 | 0 | #REF! | 238.9091 | 1e-06 | 238.9091 | 1e-06 | 4.1857e-09
#REF! | #REF! | #REF! | #REF! | #REF! | 16.7108 | 23.8615 | 29.4479 | 0 | #REF! | 268.9091 | 1e-06 | 268.9091 | 1e-06 | 3.7187e-09
#REF! | #REF! | #REF! | #REF! | #REF! | 17.2052 | 24.575 | 30.336 | 0 | #REF! | 307.6364 | 1e-06 | 307.6364 | 1e-06 | 3.2506e-09
#REF! | #REF! | #REF! | #REF! | #REF! | 17.7142 | 25.31 | 31.2515 | 0 | #REF! | 325.4545 | 1e-06 | 325.4545 | 1e-06 | 3.0726e-09
#REF! | #REF! | #REF! | #REF! | #REF! | 18.2387 | 26.0681 | 32.1961 | 0 | #REF! | 367.4545 | 1e-06 | 367.4545 | 1e-06 | 2.7214e-09
#REF! | #REF! | #REF! | #REF! | #REF! | 18.7792 | 26.8497 | 33.1705 | 0 | #REF! | 424.9091 | 1e-06 | 424.9091 | 1e-06 | 2.3534e-09
#REF! | #REF! | #REF! | #REF! | #REF! | 19.336 | 27.6556 | 34.1759 | 0 | #REF! | 476.5455 | 1e-06 | 476.5455 | 1e-06 | 2.0984e-09
#REF! | #REF! | #REF! | #REF! | #REF! | 19.9099 | 28.4869 | 35.2136 | 0 | #REF! | 517.8182 | 1e-06 | 517.8182 | 1e-06 | 1.9312e-09
#REF! | #REF! | #REF! | #REF! | #REF! | 20.5009 | 29.3436 | 36.2838 | 0 | #REF! | 558 | 1e-06 | 558 | 1e-06 | 1.7921e-09
#REF! | #REF! | #REF! | #REF! | #REF! | 21.1103 | 30.2277 | 37.3889 | 0 | #REF! | 619.8182 | 1e-06 | 619.8182 | 1e-06 | 1.6134e-09
#REF! | #REF! | #REF! | #REF! | #REF! | 21.738 | 31.1389 | 38.5286 | 0 | #REF! | 669.0909 | 1e-06 | 669.0909 | 1e-06 | 1.4946e-09
#REF! | #REF! | #REF! | #REF! | #REF! | 22.385 | 32.0791 | 39.7056 | 0 | #REF! | 740.5455 | 1e-06 | 740.5455 | 1e-06 | 1.3504e-09
#REF! | #REF! | #REF! | #REF! | #REF! | 23.0518 | 33.0491 | 40.9205 | 0 | #REF! | 805.8182 | 1e-06 | 805.8182 | 1e-06 | 1.241e-09
#REF! | #REF! | #REF! | #REF! | #REF! | 23.7391 | 34.0496 | 42.1749 | 0 | #REF! | 847.6364 | 9.82 | 847.6364 | 9.82 | 0.011585
#REF! | #REF! | #REF! | #REF! | #REF! | 24.4475 | 35.082 | 43.4701 | 0 | #REF! | 911.2727 | 7.17 | 911.2727 | 7.17 | 0.0078681
#REF! | #REF! | #REF! | #REF! | #REF! | 25.178 | 36.1476 | 44.8082 | 0 | #REF! | 981.8182 | 15.21 | 981.8182 | 15.21 | 0.015492
#REF! | #REF! | #REF! | #REF! | #REF! | 25.9305 | 37.2466 | 46.1892 | 0 | #REF! | 1057.8182 | 23.21 | 1057.8182 | 23.21 | 0.021941
#REF! | #REF! | #REF! | #REF! | #REF! | 26.7068 | 38.3815 | 47.6167 | 0 | #REF! | 1124.9091 | 13.24 | 1124.9091 | 13.24 | 0.01177
#REF! | #REF! | #REF! | #REF! | #REF! | 27.5067 | 39.552 | 49.0905 | 0 | #REF! | 1181.6364 | 17.02 | 1181.6364 | 17.02 | 0.014404
#REF! | #REF! | #REF! | #REF! | #REF! | 28.3317 | 40.7608 | 50.6137 | 0 | #REF! | 1260.7273 | 24.06 | 1260.7273 | 24.06 | 0.019084
#REF! | #REF! | #REF! | #REF! | #REF! | 29.1827 | 42.0091 | 52.1883 | 0 | #REF! | 1367.6364 | 18.74 | 1367.6364 | 18.74 | 0.013702
#REF! | #REF! | #REF! | #REF! | #REF! | 30.0599 | 43.2976 | 53.8153 | 0 | #REF! | 1409.2727 | 28.7 | 1409.2727 | 28.7 | 0.020365
#REF! | #REF! | #REF! | #REF! | #REF! | 30.9646 | 44.6281 | 55.497 | 0 | #REF! | 1474 | 40.4 | 1474 | 40.4 | 0.027408
#REF! | #REF! | #REF! | #REF! | #REF! | 31.8977 | 46.0022 | 57.2358 | 0 | #REF! | 1566.9091 | 102.58 | 1566.9091 | 102.58 | 0.065466
#REF! | #REF! | #REF! | #REF! | #REF! | 32.8598 | 47.421 | 59.0331 | 0 | #REF! | 1617.8182 | 336.57 | 1617.8182 | 336.57 | 0.20804
#REF! | #REF! | #REF! | #REF! | #REF! | 33.8526 | 48.8872 | 60.8927 | 0 | #REF! | 1709.8182 | 374.98 | 1709.8182 | 374.98 | 0.21931
#REF! | #REF! | #REF! | #REF! | #REF! | 34.8772 | 50.4024 | 62.8168 | 0 | #REF! | 1776.5455 | 732.53 | 1776.5455 | 732.53 | 0.41233
#REF! | #REF! | #REF! | #REF! | #REF! | 35.9336 | 51.9672 | 64.8062 | 0 | #REF! | 1817.4545 | 1040.57 | 1817.4545 | 1040.57 | 0.57254
#REF! | #REF! | #REF! | #REF! | #REF! | 37.0243 | 53.5853 | 66.8662 | 0 | #REF! | 1849.6364 | 1348.17 | 1849.6364 | 1348.17 | 0.72888
#REF! | #REF! | #REF! | #REF! | #REF! | 38.1491 | 55.2568 | 68.997 | 0 | #REF! | 1926.5455 | 1557.84 | 1926.5455 | 1557.84 | 0.80862
#REF! | #REF! | #REF! | #REF! | #REF! | 39.3104 | 56.9855 | 71.2036 | 0 | #REF! | 1942.1818 | 1640.62 | 1942.1818 | 1640.62 | 0.84473
#REF! | #REF! | #REF! | #REF! | #REF! | 40.5094 | 58.7733 | 73.4891 | 0 | #REF! | 2022.5455 | 1730.86 | 2022.5455 | 1730.86 | 0.85578
#REF! | #REF! | #REF! | #REF! | #REF! | 41.7467 | 60.6217 | 75.8554 | 0 | #REF! | 2065.8182 | 1809.01 | 2065.8182 | 1809.01 | 0.87569
#REF! | #REF! | #REF! | #REF! | #REF! | 43.0241 | 62.5336 | 78.3068 | 0 | #REF! | 2120.5455 | 1974.75 | 2120.5455 | 1974.75 | 0.93125
#REF! | #REF! | #REF! | #REF! | #REF! | 44.3426 | 64.5108 | 80.8461 | 0 | #REF! | 2114 | 1971.26 | 2114 | 1971.26 | 0.93248
#REF! | #REF! | #REF! | #REF! | #REF! | 45.705 | 66.558 | 83.4794 | 0 | #REF! | 2135.4545 | 1934.03 | 2135.4545 | 1934.03 | 0.90568
#REF! | #REF! | #REF! | #REF! | #REF! | 47.1116 | 68.6762 | 86.2086 | 0 | #REF! | 2135.0909 | 2101.29 | 2135.0909 | 2101.29 | 0.98417
#REF! | #REF! | #REF! | #REF! | #REF! | 48.5645 | 70.8688 | 89.0387 | 0 | #REF! | 2220.3636 | 2156.1101 | 2220.3636 | 2156.1101 | 0.97106
#REF! | #REF! | #REF! | #REF! | #REF! | 50.0655 | 73.139 | 91.9742 | 0 | #REF! | 2225.2727 | 2005.57 | 2225.2727 | 2005.57 | 0.90127
#REF! | #REF! | #REF! | #REF! | #REF! | 51.6159 | 75.4892 | 95.019 | 0 | #REF! | 2201.4545 | 2107.6499 | 2201.4545 | 2107.6499 | 0.95739
#REF! | #REF! | #REF! | #REF! | #REF! | 53.2189 | 77.9252 | 98.1808 | 0 | #REF! | 2202.7273 | 2042.85 | 2202.7273 | 2042.85 | 0.92742
#REF! | #REF! | #REF! | #REF! | #REF! | 54.8753 | 80.4485 | 101.4623 | 0 | #REF! | 2210.1818 | 2132.8601 | 2210.1818 | 2132.8601 | 0.96502
#REF! | #REF! | #REF! | #REF! | #REF! | 56.587 | 83.0627 | 104.8689 | 0 | #REF! | 2159.0909 | 2078.46 | 2159.0909 | 2078.46 | 0.96266
#REF! | #REF! | #REF! | #REF! | #REF! | 58.3577 | 85.7742 | 108.4097 | 0 | #REF! | 2159.4545 | 2112.46 | 2159.4545 | 2112.46 | 0.97824
#REF! | #REF! | #REF! | #REF! | #REF! | 60.1877 | 88.5842 | 112.0869 | 0 | #REF! | 2138.3636 | 2045.59 | 2138.3636 | 2045.59 | 0.95661
#REF! | #REF! | #REF! | #REF! | #REF! | 62.0814 | 91.5002 | 115.9116 | 0 | #REF! | 2110 | 2028.9499 | 2110 | 2028.9499 | 0.96159
#REF! | #REF! | #REF! | #REF! | #REF! | 64.0399 | 94.5247 | 119.8872 | 0 | #REF! | 2071.0909 | 2085.52 | 2071.0909 | 2085.52 | 1.007
#REF! | #REF! | #REF! | #REF! | #REF! | 66.0656 | 97.6627 | 124.0218 | 0 | #REF! | 2042.9091 | 1951.32 | 2042.9091 | 1951.32 | 0.95517
#REF! | #REF! | #REF! | #REF! | #REF! | 68.1632 | 100.9218 | 128.3264 | 0 | #REF! | 1952 | 2000.97 | 1952 | 2000.97 | 1.0251
#REF! | #REF! | #REF! | #REF! | #REF! | 70.3337 | 104.3051 | 132.8059 | 0 | #REF! | 1917.4545 | 1848.72 | 1917.4545 | 1848.72 | 0.96415
#REF! | #REF! | #REF! | #REF! | #REF! | 72.5804 | 107.8187 | 137.4695 | 0 | #REF! | 1821.4545 | 1875.63 | 1821.4545 | 1875.63 | 1.0297
#REF! | #REF! | #REF! | #REF! | #REF! | 74.9075 | 111.4706 | 142.3292 | 0 | #REF! | 1763.0909 | 1754.32 | 1763.0909 | 1754.32 | 0.99503
#REF! | #REF! | #REF! | #REF! | #REF! | 77.318 | 115.2664 | 147.3936 | 0 | #REF! | 1689.0909 | 1727.01 | 1689.0909 | 1727.01 | 1.0224
#REF! | #REF! | #REF! | #REF! | #REF! | 79.816 | 119.2142 | 152.6751 | 0 | #REF! | 1626 | 1695.21 | 1626 | 1695.21 | 1.0426
#REF! | #REF! | #REF! | #REF! | #REF! | 82.404 | 123.3196 | 158.1822 | 0 | #REF! | 1514.9091 | 1537.62 | 1514.9091 | 1537.62 | 1.015
#REF! | #REF! | #REF! | #REF! | #REF! | 85.0868 | 127.5915 | 163.9288 | 0 | #REF! | 1448 | 1456.12 | 1448 | 1456.12 | 1.0056
#REF! | #REF! | #REF! | #REF! | #REF! | 87.8686 | 132.0382 | 169.9274 | 0 | #REF! | 1360.1818 | 1411.58 | 1360.1818 | 1411.58 | 1.0378
#REF! | #REF! | #REF! | #REF! | #REF! | 90.7537 | 136.6688 | 176.1921 | 0 | #REF! | 1278.1818 | 1379.64 | 1278.1818 | 1379.64 | 1.0794
#REF! | #REF! | #REF! | #REF! | #REF! | 93.7471 | 141.4927 | 182.7371 | 0 | #REF! | 1215.0909 | 1234.02 | 1215.0909 | 1234.02 | 1.0156
#REF! | #REF! | #REF! | #REF! | #REF! | 96.8535 | 146.52 | 189.5785 | 0 | #REF! | 1131.6364 | 1157.68 | 1131.6364 | 1157.68 | 1.023
#REF! | #REF! | #REF! | #REF! | #REF! | 100.0783 | 151.7616 | 196.7326 | 0 | #REF! | 1059.0909 | 1030.9399 | 1059.0909 | 1030.9399 | 0.97342
#REF! | #REF! | #REF! | #REF! | #REF! | 103.4272 | 157.2286 | 204.2168 | 0 | #REF! | 960 | 981.79 | 960 | 981.79 | 1.0227
#REF! | #REF! | #REF! | #REF! | #REF! | 106.906 | 162.9336 | 212.0502 | 0 | #REF! | 942.9091 | 925.02 | 942.9091 | 925.02 | 0.98103
#REF! | #REF! | #REF! | #REF! | #REF! | 110.5211 | 168.889 | 220.2523 | 0 | #REF! | 897.0909 | 858.51 | 897.0909 | 858.51 | 0.95699
#REF! | #REF! | #REF! | #REF! | #REF! | 114.2782 | 175.1077 | 228.8429 | 0 | #REF! | 821.6364 | 847 | 821.6364 | 847 | 1.0309
#REF! | #REF! | #REF! | #REF! | #REF! | 118.1862 | 181.6068 | 237.8484 | 0 | #REF! | 752.3636 | 804.61 | 752.3636 | 804.61 | 1.0694
#REF! | #REF! | #REF! | #REF! | #REF! | 122.2526 | 188.4019 | 247.2922 | 0 | #REF! | 727.6364 | 754.42 | 727.6364 | 754.42 | 1.0368
#REF! | #REF! | #REF! | #REF! | #REF! | 126.4835 | 195.5069 | 257.1963 | 0 | #REF! | 682.1818 | 686.35 | 682.1818 | 686.35 | 1.0061
#REF! | #REF! | #REF! | #REF! | #REF! | 130.8886 | 202.941 | 267.5902 | 0 | #REF! | 656 | 681.38 | 656 | 681.38 | 1.0387
#REF! | #REF! | #REF! | #REF! | #REF! | 135.4767 | 210.7229 | 278.502 | 0 | #REF! | 703.0909 | 670.83 | 703.0909 | 670.83 | 0.95412
#REF! | #REF! | #REF! | #REF! | #REF! | 140.2573 | 218.8726 | 289.9625 | 0 | #REF! | 956.7273 | 747.62 | 956.7273 | 747.62 | 0.78143
#REF! | #REF! | #REF! | #REF! | #REF! | 145.2401 | 227.41 | 302.0025 | 0 | #REF! | 1390.7273 | 1098.9 | 1390.7273 | 1098.9 | 0.79016
#REF! | #REF! | #REF! | #REF! | #REF! | 150.4375 | 236.3612 | 314.661 | 0 | #REF! | 1839.0909 | 1519.29 | 1839.0909 | 1519.29 | 0.82611
#REF! | #REF! | #REF! | #REF! | #REF! | 155.8611 | 245.7499 | 327.9739 | 0 | #REF! | 1957.0909 | 1843.1999 | 1957.0909 | 1843.1999 | 0.94181
#REF! | #REF! | #REF! | #REF! | #REF! | 161.5214 | 255.5987 | 341.976 | 0 | #REF! | 1802.7273 | 1919.92 | 1802.7273 | 1919.92 | 1.065
#REF! | #REF! | #REF! | #REF! | #REF! | 167.4319 | 265.9359 | 356.7095 | 0 | #REF! | 1524.9091 | 1143.8001 | 1524.9091 | 1143.8001 | 0.75008
#REF! | #REF! | #REF! | #REF! | #REF! | 173.609 | 276.7946 | 372.2247 | 0 | #REF! | 1122 | 469 | 1122 | 469 | 0.418
#REF! | #REF! | #REF! | #REF! | #REF! | 180.065 | 288.2008 | 388.5604 | 0 | #REF! | 719.4545 | 203.04 | 719.4545 | 203.04 | 0.28221
Perform Charge Correction
Initialize Charge Correction

## Aerosol_Props
Calibration Aerosol Properties:
Van't Hoff Factor Parameterization - Pitzer Model
Solute - Ammonium Sulfate |  |  |  | Temperature = 303 K
A |  | 2.0311e-09 | m
Density |  | 1770 | kg/m3 | %SS | Dp,crit | AS_VHF
Molecular Weight |  | 0.13214 | kg/mol | 0.1 | 125.41 | 2.7
Van't Hoff factor: |  | 2.5 | (nominal) | 0.2 | 79.94 | 2.609
| a0 | 1.8853 | VHF(Dp) = a0 + a1*Dp | 0.3 | 61.5 | 2.549
| a1 | 0.024642 | + a2*Dp^2 + | 0.4 | 51.09 | 2.504
| a2 | -0.00036942 | a3*Dp^3 + | 0.5 | 44.26 | 2.467
| a3 | 2.8686e-06 | a4*Dp^4 | 0.6 | 39.37 | 2.437
| a4 | -8.5855e-09 |  | 0.7 | 35.67 | 2.41
Dynamic Shape Factor |  | 1 |  | 0.8 | 32.75 | 2.387
Solvent - Water |  |  |  | 0.9 | 30.37 | 2.367
Density |  | 994.323 | kg/m3 | 1 | 28.39 | 2.348
Molecular Weight |  | 0.01802 | kg/mol | 1.1 | 26.72 | 2.331
Surface Tension |  | 0.070581 | N/m | 1.2 | 25.28 | 2.315
Ideal Gas Constant |  | 8.314 | J/mol*K | 1.3 | 24.02 | 2.301
Temperature |  | 303 | K | 1.4 | 22.92 | 2.287
Solute - Sodium Chloride |  |  |  | Van't Hoff Factor Parameterization - Pitzer Model
A |  | 2.0311e-09 | m | Temperature = 303 K
Density |  | 2165 | kg/m3
Molecular Weight |  | 0.05844 | kg/mol | %SS | Dp,crit | SC_VHF
Van't Hoff factor: |  | 2 | (nominal) | 0.1 | 101.082 | 1.9315
| a0 | 1.8356 | VHF(Dp) = a0 + a1*Dp | 0.2 | 64.1318 | 1.912
| a1 | 0.001154 | + a2*Dp^2 + | 0.3 | 49.1686 | 1.8997
| a2 | 1.8045e-05 | a3*Dp^3 + | 0.4 | 40.7265 | 1.891
| a3 | -4.0105e-07 | a4*Dp^4 | 0.5 | 35.1912 | 1.8843
| a4 | 2.0033e-09 |  | 0.6 | 31.2316 | 1.8791
Dynamic Shape Factor |  | 1.08 |  | 0.7 | 28.2329 | 1.8749
Solvent - Water |  |  |  | 0.8 | 25.8685 | 1.8715
Density |  | 994.323 | kg/m3 | 0.9 | 23.9471 | 1.8686
Molecular Weight |  | 0.01802 | kg/mol | 1 | 22.3491 | 1.8662
Surface Tension |  | 0.070581 | N/m | 1.1 | 20.995 | 1.8643
Ideal Gas Constant |  | 8.314 | J/mol*K | 1.2 | 19.8301 | 1.8626
Temperature |  | 303 | K | 1.3 | 18.8152 | 1.8612
|  |  |  | 1.4 | 17.9216 | 1.8601
|  |  |  | 1.5 | 17.1274 | 1.8591