## CTF_updated
CTF Simulation |  | © Henning Stahlberg, Henning.Stahlberg@epfl.ch, March 2022, http://lbem.epfl.ch
This version includes defocus spread, contributed by Victor Y. Zenou, Materials Eng. Dept, Ben-Gurion University, Israel
Name | Symbol | Value | Unit | Value | Unit | modify the blue areas only
High Voltage | U | 300 | kV | 300000 | V
Wavelength | Lambda | 1.9688 | pm | 1.9688e-12 | m | Jong and van Dyck, Ultramicr. 49 (1993) 66-80
Spherical Aberation Constant | Cs | 2.7 | mm | 0.0027 | m | ai(therm. el. source) = 0.7 mrad
Size of Electron Source | Qo | 0.023056 | mrad | 2.3056e-05 | rad | ai(field emission source) = 0.1 mrad, but much smaller us typically used today, such as 0.002 mrad
Opening Angle of Source, Beam convergence | ai | 0.01 | mrad | 1e-05 | rad | FEG is 103 to 104 times brighter
Defocus (underfocus positive) | ∆z | 300 | nm | -3e-07 | m
Defocus generalized (Scherzer: ∆z*=1) | ∆z* | -4.1147 |  | -4.1147
Defocus Scherzer (underfocus positive) | ∆zs | 72.9094 | nm | 7.2909e-08 | m | Reimer, p. 219 :
Chromatic Aberation Constant | Cc | 2.7 | mm | 0.0027 | m | ∆E(therm. el. source) = 1.0 … 2.0 eV
Energy Spread of Electron Source | ∆E | 0.7 | eV | 0.7 | eV | ∆E(field emission gun) = 0.2 … 0.7 eV
Point resolution | 1/qs | 2.4627 | Å | 2.4627e-10 | m
Apply Scattering Profile | 1=yes,0=no | 0
Apply Envelope Function | 1=yes,0=no | 1
Fraction of Amplitude Contrast | fraction | 0.07 |  | 0.99755 | phase contrast
Theta increment | increment | 0.0002
Additional PhaseShift from PhasePlate | ∆W | 0 | deg | 0 | rad
Focal spread | σ∆z | 0 | nm | 0 | m
0
Name | Symbol | Value | Unit | Value | Unit
Energy Spead of Electron Source | ∆E | 0.7 | eV
kinetic energy of electron | E | 300 | keV | 300000 | eV
Energy of electron at rest | Eo | 511 | keV | 511000 | eV
relativistic correction factor | fR | 0.77352
FWHM of Energy spread of Electrons | H | 4.8732 | nm | 4.8732e-09 | m
Planck's quantum | h | 6.6256e-34 | N m s
Light speed | c | 299790000 | m/s
Scat.Profile Carbon Amplitude 1 | a1 | 1.494 |  | 1.494
Scat.Profile Carbon Amplitude 2 | a2 | 0.937 |  | 0.937
Scat.Profile Carbon Halfwidth 1 | b1 | 23.22 | 1/Å2 | 2.322e-19 | 1/m2
Scat.Profile Carbon Halfwidth 2 | b2 | 3.79 | 1/Å2 | 3.79e-20 | 1/m2
[6 empty] | Th.Prep. | Theta | Resolution | Resolution | Resolution | ThetaStar | Phaseshift | sin(W) | Es | Ed | fq | ET | Ec | Envelope | CTF-1 | PowerSpec
[7 empty] | T | q | q | 1/q | T* | W(q) | sin(W) | Ks(q) | Kc(q) | f(q) | Kt(q) | Kc(q) | Ks*Kc*Kt | CTF-1 | PS
[7 empty] | rad | 1/m | 1/nm | nm |  | rad
[7 empty] | Theta | Resolution | Resolution | Resolution | ThetaStar | Phaseshift | sin(W) | Es | Ed | fq | ET | EC | Envelope | CTF | PowerSpec | CTF-Abs.
[5 empty] | 0.0002 | 0.0002 | 0.00010638 | 54031419.2703 | 0.054031 | 18.5077 | 0.020471 | 0.0054174 | 0.075403 | 1 | 1 | 0.99954 | 1 | 1 | 1 | 0.075403 | 0.0056856 | 0.075403
[6 empty] | 0.0004 | 0.00021276 | 108062838.5406 | 0.10806 | 9.2539 | 0.040942 | 0.021673 | 0.091602 | 1 | 1 | 0.99817 | 1 | 1 | 1 | 0.091601 | 0.0083908 | 0.091601
[6 empty] | 0.0006 | 0.00031913 | 162094257.8109 | 0.16209 | 6.1692 | 0.061413 | 0.048776 | 0.11855 | 1 | 1 | 0.99588 | 1 | 1 | 1 | 0.11855 | 0.014055 | 0.11855
[6 empty] | 0.0008 | 0.00042551 | 216125677.0812 | 0.21613 | 4.6269 | 0.081884 | 0.086744 | 0.15616 | 0.99999 | 1 | 0.99269 | 1 | 1 | 0.99999 | 0.15616 | 0.024386 | 0.15616
[6 empty] | 0.001 | 0.00053189 | 270157096.3515 | 0.27016 | 3.7016 | 0.10236 | 0.1356 | 0.20421 | 0.99999 | 1 | 0.98861 | 1 | 1 | 0.99999 | 0.20421 | 0.041701 | 0.20421
[6 empty] | 0.0012 | 0.00063827 | 324188515.6218 | 0.32419 | 3.0846 | 0.12283 | 0.19537 | 0.26232 | 0.99999 | 1 | 0.98365 | 1 | 1 | 0.99998 | 0.26232 | 0.068812 | 0.26232
[6 empty] | 0.0014 | 0.00074464 | 378219934.8921 | 0.37822 | 2.644 | 0.1433 | 0.2661 | 0.32986 | 0.99998 | 1 | 0.97784 | 1 | 1 | 0.99998 | 0.32986 | 0.1088 | 0.32986
[6 empty] | 0.0016 | 0.00085102 | 432251354.1624 | 0.43225 | 2.3135 | 0.16377 | 0.34782 | 0.40583 | 0.99998 | 1 | 0.97119 | 1 | 1 | 0.99997 | 0.40581 | 0.16468 | 0.40581
[6 empty] | 0.0018 | 0.0009574 | 486282773.4328 | 0.48628 | 2.0564 | 0.18424 | 0.4406 | 0.48875 | 0.99997 | 0.99999 | 0.96373 | 1 | 0.99999 | 0.99996 | 0.48873 | 0.23886 | 0.48873
[6 empty] | 0.002 | 0.0010638 | 540314192.7031 | 0.54031 | 1.8508 | 0.20471 | 0.54447 | 0.57657 | 0.99996 | 0.99999 | 0.95548 | 1 | 0.99999 | 0.99995 | 0.57654 | 0.3324 | 0.57654
[6 empty] | 0.0022 | 0.0011702 | 594345611.9734 | 0.59435 | 1.6825 | 0.22518 | 0.65951 | 0.66655 | 0.99995 | 0.99998 | 0.94648 | 1 | 0.99998 | 0.99994 | 0.6665 | 0.44423 | 0.6665
[6 empty] | 0.0024 | 0.0012765 | 648377031.2437 | 0.64838 | 1.5423 | 0.24565 | 0.78578 | 0.75512 | 0.99994 | 0.99998 | 0.93675 | 1 | 0.99998 | 0.99992 | 0.75506 | 0.57012 | 0.75506
[6 empty] | 0.0026 | 0.0013829 | 702408450.514 | 0.70241 | 1.4237 | 0.26612 | 0.92337 | 0.8379 | 0.99993 | 0.99997 | 0.92634 | 1 | 0.99997 | 0.9999 | 0.83782 | 0.70194 | 0.83782
[6 empty] | 0.0028 | 0.0014893 | 756439869.7843 | 0.75644 | 1.322 | 0.28659 | 1.0724 | 0.90964 | 0.99992 | 0.99996 | 0.91527 | 1 | 0.99996 | 0.99988 | 0.90953 | 0.82724 | 0.90953
[6 empty] | 0.003 | 0.0015957 | 810471289.0546 | 0.81047 | 1.2338 | 0.30707 | 1.2328 | 0.96432 | 0.99991 | 0.99994 | 0.90359 | 1 | 0.99994 | 0.99985 | 0.96418 | 0.92964 | 0.96418
[6 empty] | 0.0032 | 0.001702 | 864502708.3249 | 0.8645 | 1.1567 | 0.32754 | 1.4049 | 0.99541 | 0.9999 | 0.99992 | 0.89133 | 1 | 0.99992 | 0.99982 | 0.99523 | 0.99048 | 0.99523
[6 empty] | 0.0034 | 0.0018084 | 918534127.5952 | 0.91853 | 1.0887 | 0.34801 | 1.5886 | 0.99614 | 0.99989 | 0.9999 | 0.87853 | 1 | 0.9999 | 0.99979 | 0.99593 | 0.99188 | 0.99593
[6 empty] | 0.0036 | 0.0019148 | 972565546.8655 | 0.97257 | 1.0282 | 0.36848 | 1.7841 | 0.96012 | 0.99987 | 0.99988 | 0.86524 | 1 | 0.99988 | 0.99975 | 0.95988 | 0.92137 | 0.95988
[6 empty] | 0.0038 | 0.0020212 | 1026596966.1358 | 1.0266 | 0.97409 | 0.38895 | 1.9915 | 0.88196 | 0.99985 | 0.99985 | 0.8515 | 1 | 0.99985 | 0.9997 | 0.8817 | 0.7774 | 0.8817
[6 empty] | 0.004 | 0.0021276 | 1080628385.4061 | 1.0806 | 0.92539 | 0.40942 | 2.211 | 0.75821 | 0.99984 | 0.99981 | 0.83735 | 1 | 0.99981 | 0.99965 | 0.75794 | 0.57447 | 0.75794
[6 empty] | 0.0042 | 0.0022339 | 1134659804.6764 | 1.1347 | 0.88132 | 0.42989 | 2.4426 | 0.58829 | 0.99982 | 0.99977 | 0.82284 | 1 | 0.99977 | 0.99959 | 0.58805 | 0.3458 | 0.58805
[6 empty] | 0.0044 | 0.0023403 | 1188691223.9467 | 1.1887 | 0.84126 | 0.45036 | 2.6865 | 0.37559 | 0.9998 | 0.99973 | 0.80801 | 1 | 0.99973 | 0.99953 | 0.37541 | 0.14094 | 0.37541
[6 empty] | 0.0046 | 0.0024467 | 1242722643.217 | 1.2427 | 0.80468 | 0.47083 | 2.9428 | 0.12834 | 0.99978 | 0.99967 | 0.79289 | 1 | 0.99967 | 0.99945 | 0.12827 | 0.016452 | 0.12827
[6 empty] | 0.0048 | 0.0025531 | 1296754062.4873 | 1.2968 | 0.77116 | 0.4913 | 3.2118 | -0.13978 | 0.99976 | 0.99961 | 0.77754 | 1 | 0.99961 | 0.99937 | -0.13969 | 0.019514 | 0.13969
[6 empty] | 0.005 | 0.0026594 | 1350785481.7576 | 1.3508 | 0.74031 | 0.51178 | 3.4934 | -0.40951 | 0.99974 | 0.99954 | 0.762 | 1 | 0.99954 | 0.99928 | -0.40921 | 0.16745 | 0.40921
[6 empty] | 0.0052 | 0.0027658 | 1404816901.0279 | 1.4048 | 0.71184 | 0.53225 | 3.788 | -0.65674 | 0.99971 | 0.99947 | 0.7463 | 1 | 0.99947 | 0.99918 | -0.6562 | 0.4306 | 0.6562
[6 empty] | 0.0054 | 0.0028722 | 1458848320.2983 | 1.4588 | 0.68547 | 0.55272 | 4.0957 | -0.85427 | 0.99969 | 0.99938 | 0.7305 | 1 | 0.99938 | 0.99907 | -0.85347 | 0.72841 | 0.85347
[6 empty] | 0.0056 | 0.0029786 | 1512879739.5686 | 1.5129 | 0.66099 | 0.57319 | 4.4166 | -0.97462 | 0.99966 | 0.99928 | 0.71462 | 1 | 0.99928 | 0.99894 | -0.97359 | 0.94788 | 0.97359
[6 empty] | 0.0058 | 0.0030849 | 1566911158.8389 | 1.5669 | 0.6382 | 0.59366 | 4.7509 | -0.99411 | 0.99963 | 0.99918 | 0.69871 | 1 | 0.99918 | 0.9988 | -0.99292 | 0.9859 | 0.99292
[6 empty] | 0.006 | 0.0031913 | 1620942578.1092 | 1.6209 | 0.61693 | 0.61413 | 5.0988 | -0.8976 | 0.9996 | 0.99906 | 0.6828 | 1 | 0.99906 | 0.99865 | -0.89639 | 0.80352 | 0.89639
[6 empty] | 0.0062 | 0.0032977 | 1674973997.3795 | 1.675 | 0.59702 | 0.6346 | 5.4606 | -0.68349 | 0.99957 | 0.99892 | 0.66692 | 1 | 0.99892 | 0.99849 | -0.68246 | 0.46575 | 0.68246
[6 empty] | 0.0064 | 0.0034041 | 1729005416.6498 | 1.729 | 0.57837 | 0.65507 | 5.8364 | -0.36791 | 0.99953 | 0.99878 | 0.65112 | 1 | 0.99878 | 0.99831 | -0.36729 | 0.1349 | 0.36729
[6 empty] | 0.0066 | 0.0035105 | 1783036835.9201 | 1.783 | 0.56084 | 0.67554 | 6.2264 | 0.013242 | 0.9995 | 0.99862 | 0.63542 | 1 | 0.99862 | 0.99812 | 0.013217 | 0.00017468 | 0.013217
[6 empty] | 0.0068 | 0.0036168 | 1837068255.1904 | 1.8371 | 0.54435 | 0.69602 | 6.6308 | 0.40564 | 0.99946 | 0.99844 | 0.61986 | 1 | 0.99844 | 0.9979 | 0.40479 | 0.16386 | 0.40479
[6 empty] | 0.007 | 0.0037232 | 1891099674.4607 | 1.8911 | 0.52879 | 0.71649 | 7.0499 | 0.74249 | 0.99942 | 0.99825 | 0.60445 | 1 | 0.99825 | 0.99767 | 0.74076 | 0.54873 | 0.74076
[6 empty] | 0.0072 | 0.0038296 | 1945131093.731 | 1.9451 | 0.5141 | 0.73696 | 7.4839 | 0.95533 | 0.99938 | 0.99804 | 0.58923 | 1 | 0.99804 | 0.99742 | 0.95287 | 0.90795 | 0.95287
[6 empty] | 0.0074 | 0.003936 | 1999162513.0013 | 1.9992 | 0.50021 | 0.75743 | 7.933 | 0.98891 | 0.99934 | 0.99782 | 0.57422 | 1 | 0.99782 | 0.99715 | 0.98609 | 0.97237 | 0.98609
[6 empty] | 0.0076 | 0.0040423 | 2053193932.2716 | 2.0532 | 0.48705 | 0.7779 | 8.3975 | 0.8176 | 0.99929 | 0.99757 | 0.55944 | 1 | 0.99757 | 0.99686 | 0.81503 | 0.66428 | 0.81503
[6 empty] | 0.0078 | 0.0041487 | 2107225351.5419 | 2.1072 | 0.47456 | 0.79837 | 8.8776 | 0.45923 | 0.99924 | 0.9973 | 0.54491 | 1 | 0.9973 | 0.99655 | 0.45765 | 0.20944 | 0.45765
[6 empty] | 0.008 | 0.0042551 | 2161256770.8122 | 2.1613 | 0.46269 | 0.81884 | 9.3736 | -0.018835 | 0.99919 | 0.99702 | 0.53064 | 1 | 0.99702 | 0.99621 | -0.018764 | 0.00035208 | 0.018764
[6 empty] | 0.0082 | 0.0043615 | 2215288190.0825 | 2.2153 | 0.45141 | 0.83931 | 9.8857 | -0.50634 | 0.99914 | 0.99671 | 0.51666 | 1 | 0.99671 | 0.99585 | -0.50424 | 0.25425 | 0.50424
[6 empty] | 0.0084 | 0.0044679 | 2269319609.3528 | 2.2693 | 0.44066 | 0.85978 | 10.4141 | -0.87208 | 0.99908 | 0.99638 | 0.50298 | 1 | 0.99638 | 0.99546 | -0.86812 | 0.75363 | 0.86812
[6 empty] | 0.0086 | 0.0045742 | 2323351028.6231 | 2.3234 | 0.43041 | 0.88025 | 10.9593 | -0.99943 | 0.99902 | 0.99602 | 0.4896 | 1 | 0.99602 | 0.99505 | -0.99448 | 0.98899 | 0.99448
[6 empty] | 0.0088 | 0.0046806 | 2377382447.8934 | 2.3774 | 0.42063 | 0.90073 | 11.5214 | -0.82763 | 0.99896 | 0.99564 | 0.47654 | 1 | 0.99564 | 0.9946 | -0.82316 | 0.67759 | 0.82316
[6 empty] | 0.009 | 0.004787 | 2431413867.1638 | 2.4314 | 0.41128 | 0.9212 | 12.1008 | -0.38527 | 0.9989 | 0.99523 | 0.46381 | 1 | 0.99523 | 0.99413 | -0.38301 | 0.14669 | 0.38301
[6 empty] | 0.0092 | 0.0048934 | 2485445286.4341 | 2.4854 | 0.40234 | 0.94167 | 12.6977 | 0.20005 | 0.99883 | 0.99479 | 0.45141 | 1 | 0.99479 | 0.99363 | 0.19877 | 0.039511 | 0.19877
[6 empty] | 0.0094 | 0.0049997 | 2539476705.7044 | 2.5395 | 0.39378 | 0.96214 | 13.3125 | 0.72852 | 0.99876 | 0.99432 | 0.43934 | 1 | 0.99432 | 0.99309 | 0.72348 | 0.52343 | 0.72348
[6 empty] | 0.0096 | 0.0051061 | 2593508124.9747 | 2.5935 | 0.38558 | 0.98261 | 13.9454 | 0.9926 | 0.99869 | 0.99382 | 0.42761 | 1 | 0.99382 | 0.99252 | 0.98517 | 0.97057 | 0.98517
[6 empty] | 0.0098 | 0.0052125 | 2647539544.245 | 2.6475 | 0.37771 | 1.0031 | 14.5967 | 0.86301 | 0.99861 | 0.9933 | 0.41622 | 1 | 0.9933 | 0.99191 | 0.85603 | 0.73278 | 0.85603
[6 empty] | 0.01 | 0.0053189 | 2701570963.5153 | 2.7016 | 0.37015 | 1.0236 | 15.2669 | 0.3626 | 0.99853 | 0.99273 | 0.40516 | 1 | 0.99273 | 0.99127 | 0.35943 | 0.12919 | 0.35943
[6 empty] | 0.0102 | 0.0054253 | 2755602382.7856 | 2.7556 | 0.3629 | 1.044 | 15.9561 | -0.31285 | 0.99844 | 0.99214 | 0.39445 | 1 | 0.99214 | 0.99059 | -0.30991 | 0.096045 | 0.30991
[6 empty] | 0.0104 | 0.0055316 | 2809633802.0559 | 2.8096 | 0.35592 | 1.0645 | 16.6648 | -0.85569 | 0.99836 | 0.9915 | 0.38406 | 1 | 0.9915 | 0.98987 | -0.84703 | 0.71746 | 0.84703
[6 empty] | 0.0106 | 0.005638 | 2863665221.3262 | 2.8637 | 0.3492 | 1.085 | 17.3933 | -0.98301 | 0.99826 | 0.99083 | 0.37401 | 1 | 0.99083 | 0.98911 | -0.97231 | 0.94539 | 0.97231
[6 empty] | 0.0108 | 0.0057444 | 2917696640.5965 | 2.9177 | 0.34274 | 1.1054 | 18.1419 | -0.59528 | 0.99817 | 0.99013 | 0.36428 | 1 | 0.99013 | 0.98831 | -0.58832 | 0.34612 | 0.58832
[6 empty] | 0.011 | 0.0058508 | 2971728059.8668 | 2.9717 | 0.3365 | 1.1259 | 18.911 | 0.13111 | 0.99807 | 0.98938 | 0.35487 | 1 | 0.98938 | 0.98747 | 0.12946 | 0.016761 | 0.12946
[6 empty] | 0.0112 | 0.0059571 | 3025759479.1371 | 3.0258 | 0.3305 | 1.1464 | 19.7009 | 0.79647 | 0.99796 | 0.98859 | 0.34577 | 1 | 0.98859 | 0.98657 | 0.78578 | 0.61744 | 0.78578
[6 empty] | 0.0114 | 0.0060635 | 3079790898.4074 | 3.0798 | 0.3247 | 1.1668 | 20.5121 | 0.98694 | 0.99785 | 0.98776 | 0.33698 | 1 | 0.98776 | 0.98563 | 0.97276 | 0.94626 | 0.97276
[6 empty] | 0.0116 | 0.0061699 | 3133822317.6777 | 3.1338 | 0.3191 | 1.1873 | 21.3448 | 0.54488 | 0.99773 | 0.98688 | 0.32848 | 1 | 0.98688 | 0.98465 | 0.53651 | 0.28784 | 0.53651
[6 empty] | 0.0118 | 0.0062763 | 3187853736.948 | 3.1879 | 0.31369 | 1.2078 | 22.1996 | -0.2749 | 0.99761 | 0.98596 | 0.32027 | 1 | 0.98596 | 0.98361 | -0.27039 | 0.07311 | 0.27039
[6 empty] | 0.012 | 0.0063827 | 3241885156.2183 | 3.2419 | 0.30846 | 1.2283 | 23.0767 | -0.91502 | 0.99749 | 0.98499 | 0.31235 | 1 | 0.98499 | 0.98252 | -0.89902 | 0.80824 | 0.89902
[6 empty] | 0.0122 | 0.006489 | 3295916575.4886 | 3.2959 | 0.30341 | 1.2487 | 23.9765 | -0.88486 | 0.99736 | 0.98397 | 0.30469 | 1 | 0.98397 | 0.98137 | -0.86838 | 0.75408 | 0.86838
[6 empty] | 0.0124 | 0.0065954 | 3349947994.759 | 3.3499 | 0.29851 | 1.2692 | 24.8995 | -0.16249 | 0.99722 | 0.98291 | 0.29729 | 1 | 0.98291 | 0.98017 | -0.15927 | 0.025366 | 0.15927
[6 empty] | 0.0126 | 0.0067018 | 3403979414.0293 | 3.404 | 0.29377 | 1.2897 | 25.846 | 0.70564 | 0.99708 | 0.98179 | 0.29014 | 1 | 0.98179 | 0.97892 | 0.69077 | 0.47716 | 0.69077
[6 empty] | 0.0128 | 0.0068082 | 3458010833.2996 | 3.458 | 0.28918 | 1.3101 | 26.8165 | 0.98329 | 0.99693 | 0.98061 | 0.28324 | 1 | 0.98061 | 0.9776 | 0.96127 | 0.92403 | 0.96127
[6 empty] | 0.013 | 0.0069145 | 3512042252.5699 | 3.512 | 0.28473 | 1.3306 | 27.8114 | 0.38281 | 0.99677 | 0.97939 | 0.27656 | 1 | 0.97939 | 0.97622 | 0.37371 | 0.13966 | 0.37371
[6 empty] | 0.0132 | 0.0070209 | 3566073671.8402 | 3.5661 | 0.28042 | 1.3511 | 28.8312 | -0.58664 | 0.99661 | 0.9781 | 0.27011 | 1 | 0.9781 | 0.97479 | -0.57184 | 0.32701 | 0.57184
[6 empty] | 0.0134 | 0.0071273 | 3620105091.1105 | 3.6201 | 0.27624 | 1.3716 | 29.8762 | -0.99489 | 0.99644 | 0.97676 | 0.26387 | 1 | 0.97676 | 0.97328 | -0.96831 | 0.93763 | 0.96831
[6 empty] | 0.0136 | 0.0072337 | 3674136510.3808 | 3.6741 | 0.27217 | 1.392 | 30.9468 | -0.38852 | 0.99626 | 0.97536 | 0.25782 | 1 | 0.97536 | 0.97172 | -0.37753 | 0.14253 | 0.37753
[6 empty] | 0.0138 | 0.00734 | 3728167929.6511 | 3.7282 | 0.26823 | 1.4125 | 32.0437 | 0.64253 | 0.99608 | 0.9739 | 0.25198 | 1 | 0.9739 | 0.97008 | 0.6233 | 0.38851 | 0.6233
[6 empty] | 0.014 | 0.0074464 | 3782199348.9214 | 3.7822 | 0.2644 | 1.433 | 33.1671 | 0.96881 | 0.99589 | 0.97237 | 0.24631 | 1 | 0.97237 | 0.96837 | 0.93817 | 0.88016 | 0.93817
[6 empty] | 0.0142 | 0.0075528 | 3836230768.1917 | 3.8362 | 0.26067 | 1.4534 | 34.3176 | 0.16909 | 0.99569 | 0.97078 | 0.24082 | 1 | 0.97078 | 0.9666 | 0.16345 | 0.026714 | 0.16345
[6 empty] | 0.0144 | 0.0076592 | 3890262187.462 | 3.8903 | 0.25705 | 1.4739 | 35.4955 | -0.84581 | 0.99548 | 0.96913 | 0.23549 | 1 | 0.96913 | 0.96475 | -0.816 | 0.66585 | 0.816
[6 empty] | 0.0146 | 0.0077656 | 3944293606.7323 | 3.9443 | 0.25353 | 1.4944 | 36.7015 | -0.80014 | 0.99527 | 0.96741 | 0.23032 | 1 | 0.96741 | 0.96283 | -0.77039 | 0.5935 | 0.77039
[6 empty] | 0.0148 | 0.0078719 | 3998325026.0026 | 3.9983 | 0.2501 | 1.5149 | 37.936 | 0.30215 | 0.99504 | 0.96561 | 0.2253 | 1 | 0.96561 | 0.96083 | 0.29031 | 0.08428 | 0.29031
[6 empty] | 0.015 | 0.0079783 | 4052356445.2729 | 4.0524 | 0.24677 | 1.5353 | 39.1994 | 1 | 0.99481 | 0.96375 | 0.22042 | 1 | 0.96375 | 0.95875 | 0.95875 | 0.9192 | 0.95875
[6 empty] | 0.0152 | 0.0080847 | 4106387864.5432 | 4.1064 | 0.24352 | 1.5558 | 40.4923 | 0.27476 | 0.99456 | 0.96182 | 0.21567 | 1 | 0.96182 | 0.95659 | 0.26284 | 0.069083 | 0.26284
[6 empty] | 0.0154 | 0.0081911 | 4160419283.8135 | 4.1604 | 0.24036 | 1.5763 | 41.8151 | -0.86467 | 0.99431 | 0.95981 | 0.21105 | 1 | 0.95981 | 0.95435 | -0.8252 | 0.68095 | 0.8252
[6 empty] | 0.0156 | 0.0082974 | 4214450703.0838 | 4.2145 | 0.23728 | 1.5967 | 43.1684 | -0.67709 | 0.99405 | 0.95773 | 0.20655 | 1 | 0.95773 | 0.95203 | -0.64461 | 0.41552 | 0.64461
[6 empty] | 0.0158 | 0.0084038 | 4268482122.3542 | 4.2685 | 0.23428 | 1.6172 | 44.5527 | 0.59757 | 0.99377 | 0.95557 | 0.20216 | 1 | 0.95557 | 0.94962 | 0.56746 | 0.32201 | 0.56746
[6 empty] | 0.016 | 0.0085102 | 4322513541.6245 | 4.3225 | 0.23135 | 1.6377 | 45.9685 | 0.88447 | 0.99349 | 0.95333 | 0.19788 | 1 | 0.95333 | 0.94712 | 0.8377 | 0.70175 | 0.8377
[6 empty] | 0.0162 | 0.0086166 | 4376544960.8948 | 4.3765 | 0.22849 | 1.6582 | 47.4163 | -0.35454 | 0.99319 | 0.95101 | 0.1937 | 1 | 0.95101 | 0.94454 | -0.33487 | 0.11214 | 0.33487
[6 empty] | 0.0164 | 0.008723 | 4430576380.1651 | 4.4306 | 0.2257 | 1.6786 | 48.8966 | -0.96325 | 0.99288 | 0.94861 | 0.18961 | 1 | 0.94861 | 0.94186 | -0.90724 | 0.82309 | 0.90724
[6 empty] | 0.0166 | 0.0088293 | 4484607799.4354 | 4.4846 | 0.22298 | 1.6991 | 50.41 | 0.21295 | 0.99257 | 0.94613 | 0.18561 | 1 | 0.94613 | 0.93909 | 0.19998 | 0.039993 | 0.19998
[6 empty] | 0.0168 | 0.0089357 | 4538639218.7057 | 4.5386 | 0.22033 | 1.7196 | 51.9571 | 0.98184 | 0.99223 | 0.94356 | 0.18171 | 1 | 0.94356 | 0.93623 | 0.91923 | 0.84498 | 0.91923
[6 empty] | 0.017 | 0.0090421 | 4592670637.976 | 4.5927 | 0.21774 | 1.74 | 53.5384 | -0.19998 | 0.99189 | 0.94091 | 0.17788 | 1 | 0.94091 | 0.93327 | -0.18664 | 0.034833 | 0.18664
[6 empty] | 0.0172 | 0.0091485 | 4646702057.2463 | 4.6467 | 0.21521 | 1.7605 | 55.1544 | -0.96976 | 0.99153 | 0.93816 | 0.17413 | 1 | 0.93816 | 0.93022 | -0.90209 | 0.81376 | 0.90209
[6 empty] | 0.0174 | 0.0092548 | 4700733476.5166 | 4.7007 | 0.21273 | 1.781 | 56.8057 | 0.32129 | 0.99116 | 0.93533 | 0.17045 | 1 | 0.93533 | 0.92707 | 0.29785 | 0.088717 | 0.29785
[6 empty] | 0.0176 | 0.0093612 | 4754764895.7869 | 4.7548 | 0.21032 | 1.8015 | 58.4929 | 0.90325 | 0.99077 | 0.93241 | 0.16685 | 1 | 0.93241 | 0.92381 | 0.83443 | 0.69628 | 0.83443
[6 empty] | 0.0178 | 0.0094676 | 4808796315.0572 | 4.8088 | 0.20795 | 1.8219 | 60.2166 | -0.56167 | 0.99037 | 0.9294 | 0.16331 | 1 | 0.9294 | 0.92046 | -0.517 | 0.26728 | 0.517
[6 empty] | 0.018 | 0.009574 | 4862827734.3275 | 4.8628 | 0.20564 | 1.8424 | 61.9774 | -0.70643 | 0.98996 | 0.9263 | 0.15983 | 1 | 0.9263 | 0.917 | -0.64779 | 0.41964 | 0.64779
[6 empty] | 0.0182 | 0.0096804 | 4916859153.5978 | 4.9169 | 0.20338 | 1.8629 | 63.7757 | 0.84893 | 0.98953 | 0.9231 | 0.15642 | 1 | 0.9231 | 0.91343 | 0.77544 | 0.6013 | 0.77544
[6 empty] | 0.0184 | 0.0097867 | 4970890572.8681 | 4.9709 | 0.20117 | 1.8833 | 65.6124 | 0.28691 | 0.98908 | 0.9198 | 0.15307 | 1 | 0.9198 | 0.90976 | 0.26102 | 0.068131 | 0.26102
[6 empty] | 0.0186 | 0.0098931 | 5024921992.1384 | 5.0249 | 0.19901 | 1.9038 | 67.4879 | -0.99991 | 0.98862 | 0.91641 | 0.14978 | 1 | 0.91641 | 0.90598 | -0.9059 | 0.82065 | 0.9059
[6 empty] | 0.0188 | 0.0099995 | 5078953411.4087 | 5.079 | 0.19689 | 1.9243 | 69.4028 | 0.35027 | 0.98814 | 0.91292 | 0.14654 | 1 | 0.91292 | 0.90209 | 0.31598 | 0.099841 | 0.31598
[6 empty] | 0.019 | 0.010106 | 5132984830.679 | 5.133 | 0.19482 | 1.9447 | 71.3579 | 0.73705 | 0.98764 | 0.90933 | 0.14335 | 1 | 0.90933 | 0.8981 | 0.66194 | 0.43816 | 0.66194
[6 empty] | 0.0192 | 0.010212 | 5187016249.9493 | 5.187 | 0.19279 | 1.9652 | 73.3536 | -0.9196 | 0.98713 | 0.90564 | 0.14022 | 1 | 0.90564 | 0.89398 | -0.82211 | 0.67586 | 0.82211
[6 empty] | 0.0194 | 0.010319 | 5241047669.2197 | 5.241 | 0.1908 | 1.9857 | 75.3907 | 0.06252 | 0.98659 | 0.90185 | 0.13713 | 1 | 0.90185 | 0.88976 | 0.055628 | 0.0030945 | 0.055628
[6 empty] | 0.0196 | 0.010425 | 5295079088.49 | 5.2951 | 0.18885 | 2.0062 | 77.4698 | 0.84145 | 0.98604 | 0.89796 | 0.1341 | 1 | 0.89796 | 0.88542 | 0.74504 | 0.55508 | 0.74504
[6 empty] | 0.0198 | 0.010531 | 5349110507.7603 | 5.3491 | 0.18695 | 2.0266 | 79.5915 | -0.90086 | 0.98547 | 0.89396 | 0.13111 | 1 | 0.89396 | 0.88097 | -0.79363 | 0.62985 | 0.79363
[6 empty] | 0.02 | 0.010638 | 5403141927.0306 | 5.4031 | 0.18508 | 2.0471 | 81.7565 | 0.14462 | 0.98487 | 0.88986 | 0.12818 | 1 | 0.88986 | 0.8764 | 0.12675 | 0.016065 | 0.12675
[6 empty] | 0.0202 | 0.010744 | 5457173346.3009 | 5.4572 | 0.18325 | 2.0676 | 83.9654 | 0.70862 | 0.98426 | 0.88566 | 0.12528 | 1 | 0.88566 | 0.87172 | 0.61772 | 0.38157 | 0.61772
[6 empty] | 0.0204 | 0.010851 | 5511204765.5712 | 5.5112 | 0.18145 | 2.088 | 86.2189 | -0.99451 | 0.98362 | 0.88134 | 0.12244 | 1 | 0.88134 | 0.86691 | -0.86215 | 0.74331 | 0.86215
[6 empty] | 0.0206 | 0.010957 | 5565236184.8415 | 5.5652 | 0.17969 | 2.1085 | 88.5177 | 0.58362 | 0.98297 | 0.87692 | 0.11963 | 1 | 0.87692 | 0.86199 | 0.50307 | 0.25308 | 0.50307
[6 empty] | 0.0208 | 0.011063 | 5619267604.1118 | 5.6193 | 0.17796 | 2.129 | 90.8625 | 0.1728 | 0.98229 | 0.8724 | 0.11687 | 1 | 0.8724 | 0.85695 | 0.14808 | 0.021927 | 0.14808
[6 empty] | 0.021 | 0.01117 | 5673299023.3821 | 5.6733 | 0.17626 | 2.1495 | 93.2538 | -0.79794 | 0.98158 | 0.86776 | 0.11416 | 1 | 0.86776 | 0.85178 | -0.67967 | 0.46196 | 0.67967
[6 empty] | 0.0212 | 0.011276 | 5727330442.6524 | 5.7273 | 0.1746 | 2.1699 | 95.6926 | 0.99843 | 0.98086 | 0.86302 | 0.11148 | 1 | 0.86302 | 0.8465 | 0.84517 | 0.71432 | 0.84517
[6 empty] | 0.0214 | 0.011382 | 5781361861.9227 | 5.7814 | 0.17297 | 2.1904 | 98.1793 | -0.75784 | 0.9801 | 0.85817 | 0.10885 | 1 | 0.85817 | 0.8411 | -0.63741 | 0.4063 | 0.63741
[6 empty] | 0.0216 | 0.011489 | 5835393281.193 | 5.8354 | 0.17137 | 2.2109 | 100.7148 | 0.25116 | 0.97933 | 0.85321 | 0.10627 | 1 | 0.85321 | 0.83557 | 0.20986 | 0.044042 | 0.20986
[6 empty] | 0.0218 | 0.011595 | 5889424700.4633 | 5.8894 | 0.1698 | 2.2313 | 103.2997 | 0.29818 | 0.97852 | 0.84814 | 0.10372 | 1 | 0.84814 | 0.82992 | 0.24747 | 0.06124 | 0.24747
[6 empty] | 0.022 | 0.011702 | 5943456119.7336 | 5.9435 | 0.16825 | 2.2518 | 105.9348 | -0.7238 | 0.97769 | 0.84295 | 0.10121 | 1 | 0.84295 | 0.82415 | -0.59652 | 0.35584 | 0.59652
[6 empty] | 0.0222 | 0.011808 | 5997487539.0039 | 5.9975 | 0.16674 | 2.2723 | 108.6208 | 0.95358 | 0.97684 | 0.83766 | 0.09875 | 1 | 0.83766 | 0.81826 | 0.78027 | 0.60883 | 0.78027
[6 empty] | 0.0224 | 0.011914 | 6051518958.2742 | 6.0515 | 0.16525 | 2.2928 | 111.3584 | -0.99519 | 0.97595 | 0.83226 | 0.096327 | 1 | 0.83226 | 0.81224 | -0.80834 | 0.65341 | 0.80834
[6 empty] | 0.0226 | 0.012021 | 6105550377.5445 | 6.1056 | 0.16379 | 2.3132 | 114.1483 | 0.90056 | 0.97504 | 0.82675 | 0.093944 | 1 | 0.82675 | 0.80611 | 0.72595 | 0.527 | 0.72595
[6 empty] | 0.0228 | 0.012127 | 6159581796.8148 | 6.1596 | 0.16235 | 2.3337 | 116.9914 | -0.73286 | 0.97409 | 0.82112 | 0.091601 | 1 | 0.82112 | 0.79985 | -0.58618 | 0.34361 | 0.58618
[6 empty] | 0.023 | 0.012233 | 6213613216.0851 | 6.2136 | 0.16094 | 2.3542 | 119.8883 | 0.54622 | 0.97312 | 0.81539 | 0.089298 | 1 | 0.81539 | 0.79347 | 0.43341 | 0.18784 | 0.43341
[6 empty] | 0.0232 | 0.01234 | 6267644635.3555 | 6.2676 | 0.15955 | 2.3746 | 122.8398 | -0.37813 | 0.97211 | 0.80955 | 0.087035 | 1 | 0.80955 | 0.78697 | -0.29758 | 0.088551 | 0.29758
[6 empty] | 0.0234 | 0.012446 | 6321676054.6258 | 6.3217 | 0.15819 | 2.3951 | 125.8467 | 0.25039 | 0.97107 | 0.80359 | 0.084812 | 1 | 0.80359 | 0.78035 | 0.19539 | 0.038178 | 0.19539
[6 empty] | 0.0236 | 0.012553 | 6375707473.8961 | 6.3757 | 0.15685 | 2.4156 | 128.9098 | -0.17366 | 0.97 | 0.79753 | 0.082628 | 1 | 0.79753 | 0.77361 | -0.13435 | 0.018049 | 0.13435
[6 empty] | 0.0238 | 0.012659 | 6429738893.1664 | 6.4297 | 0.15553 | 2.4361 | 132.0298 | 0.15234 | 0.96889 | 0.79136 | 0.080484 | 1 | 0.79136 | 0.76675 | 0.1168 | 0.013643 | 0.1168
[6 empty] | 0.024 | 0.012765 | 6483770312.4367 | 6.4838 | 0.15423 | 2.4565 | 135.2075 | -0.18791 | 0.96775 | 0.78508 | 0.078378 | 1 | 0.78508 | 0.75977 | -0.14277 | 0.020382 | 0.14277
[6 empty] | 0.0242 | 0.012872 | 6537801731.707 | 6.5378 | 0.15296 | 2.477 | 138.4437 | 0.27984 | 0.96658 | 0.7787 | 0.076311 | 1 | 0.7787 | 0.75267 | 0.21063 | 0.044366 | 0.21063
[6 empty] | 0.0244 | 0.012978 | 6591833150.9773 | 6.5918 | 0.1517 | 2.4975 | 141.7391 | -0.42371 | 0.96537 | 0.77221 | 0.074283 | 1 | 0.77221 | 0.74546 | -0.31586 | 0.099766 | 0.31586
[6 empty] | 0.0246 | 0.013084 | 6645864570.2476 | 6.6459 | 0.15047 | 2.5179 | 145.0947 | 0.60641 | 0.96412 | 0.76561 | 0.072293 | 1 | 0.76561 | 0.73814 | 0.44761 | 0.20036 | 0.44761
[6 empty] | 0.0248 | 0.013191 | 6699895989.5179 | 6.6999 | 0.14926 | 2.5384 | 148.5113 | -0.7995 | 0.96283 | 0.75891 | 0.070342 | 1 | 0.75891 | 0.7307 | -0.58419 | 0.34128 | 0.58419
[6 empty] | 0.025 | 0.013297 | 6753927408.7882 | 6.7539 | 0.14806 | 2.5589 | 151.9895 | 0.95304 | 0.96151 | 0.7521 | 0.068428 | 1 | 0.7521 | 0.72315 | 0.68919 | 0.47498 | 0.68919
[6 empty] | 0.0252 | 0.013404 | 6807958828.0585 | 6.808 | 0.14689 | 2.5794 | 155.5303 | -0.99581 | 0.96014 | 0.74519 | 0.066551 | 1 | 0.74519 | 0.71549 | -0.7125 | 0.50765 | 0.7125
[6 empty] | 0.0254 | 0.01351 | 6861990247.3288 | 6.862 | 0.14573 | 2.5998 | 159.1345 | 0.85036 | 0.95873 | 0.73819 | 0.064713 | 1 | 0.73819 | 0.70773 | 0.60182 | 0.36219 | 0.60182
[6 empty] | 0.0256 | 0.013616 | 6916021666.5991 | 6.916 | 0.14459 | 2.6203 | 162.8029 | -0.47049 | 0.95729 | 0.73108 | 0.062911 | 1 | 0.73108 | 0.69985 | -0.32927 | 0.10842 | 0.32927
[6 empty] | 0.0258 | 0.013723 | 6970053085.8694 | 6.9701 | 0.14347 | 2.6408 | 166.5364 | -0.10188 | 0.9558 | 0.72387 | 0.061146 | 1 | 0.72387 | 0.69188 | -0.070486 | 0.0049682 | 0.070486
[6 empty] | 0.026 | 0.013829 | 7024084505.1397 | 7.0241 | 0.14237 | 2.6612 | 170.3358 | 0.68883 | 0.95426 | 0.71657 | 0.059418 | 1 | 0.71657 | 0.6838 | 0.47102 | 0.22186 | 0.47102
[6 empty] | 0.0262 | 0.013935 | 7078115924.41 | 7.0781 | 0.14128 | 2.6817 | 174.202 | -0.99628 | 0.95268 | 0.70918 | 0.057725 | 1 | 0.70918 | 0.67562 | -0.67311 | 0.45307 | 0.67311
[6 empty] | 0.0264 | 0.014042 | 7132147343.6803 | 7.1321 | 0.14021 | 2.7022 | 178.1358 | 0.76103 | 0.95106 | 0.70169 | 0.056069 | 1 | 0.70169 | 0.66735 | 0.50787 | 0.25793 | 0.50787
[6 empty] | 0.0266 | 0.014148 | 7186178762.9506 | 7.1862 | 0.13916 | 2.7226 | 182.1382 | -0.0041446 | 0.94939 | 0.69411 | 0.054448 | 1 | 0.69411 | 0.65898 | -0.0027312 | 7.4595e-06 | 0.0027312
[6 empty] | 0.0268 | 0.014255 | 7240210182.2209 | 7.2402 | 0.13812 | 2.7431 | 186.2099 | -0.79922 | 0.94767 | 0.68645 | 0.052863 | 1 | 0.68645 | 0.65052 | -0.51991 | 0.27031 | 0.51991
[6 empty] | 0.027 | 0.014361 | 7294241601.4913 | 7.2942 | 0.13709 | 2.7636 | 190.3519 | 0.93743 | 0.9459 | 0.6787 | 0.051312 | 1 | 0.6787 | 0.64198 | 0.60181 | 0.36218 | 0.60181
[6 empty] | 0.0272 | 0.014467 | 7348273020.7616 | 7.3483 | 0.13609 | 2.7841 | 194.5651 | -0.14312 | 0.94408 | 0.67086 | 0.049796 | 1 | 0.67086 | 0.63335 | -0.090643 | 0.0082161 | 0.090643
[6 empty] | 0.0274 | 0.014574 | 7402304440.0319 | 7.4023 | 0.13509 | 2.8045 | 198.8503 | -0.84149 | 0.94221 | 0.66294 | 0.048314 | 1 | 0.66294 | 0.62463 | -0.52562 | 0.27628 | 0.52562
[6 empty] | 0.0276 | 0.01468 | 7456335859.3022 | 7.4563 | 0.13411 | 2.825 | 203.2085 | 0.7986 | 0.94029 | 0.65495 | 0.046866 | 1 | 0.65495 | 0.61584 | 0.49181 | 0.24188 | 0.49181
[6 empty] | 0.0278 | 0.014786 | 7510367278.5725 | 7.5104 | 0.13315 | 2.8455 | 207.6406 | 0.35743 | 0.93832 | 0.64688 | 0.045451 | 1 | 0.64688 | 0.60698 | 0.21695 | 0.047068 | 0.21695
[6 empty] | 0.028 | 0.014893 | 7564398697.8428 | 7.5644 | 0.1322 | 2.8659 | 212.1474 | -0.98723 | 0.93629 | 0.63873 | 0.044069 | 1 | 0.63873 | 0.59804 | -0.5904 | 0.34858 | 0.5904
[6 empty] | 0.0282 | 0.014999 | 7618430117.1131 | 7.6184 | 0.13126 | 2.8864 | 216.73 | -0.030171 | 0.93421 | 0.63052 | 0.04272 | 1 | 0.63052 | 0.58904 | -0.017772 | 0.00031584 | 0.017772
[6 empty] | 0.0284 | 0.015106 | 7672461536.3834 | 7.6725 | 0.13034 | 2.9069 | 221.3892 | 0.99974 | 0.93208 | 0.62223 | 0.041403 | 1 | 0.62223 | 0.57997 | 0.57981 | 0.33619 | 0.57981
[6 empty] | 0.0286 | 0.015212 | 7726492955.6537 | 7.7265 | 0.12942 | 2.9274 | 226.126 | 0.0014276 | 0.92988 | 0.61389 | 0.040118 | 1 | 0.61389 | 0.57084 | 0.00081494 | 6.6412e-07 | 0.00081494
[6 empty] | 0.0288 | 0.015318 | 7780524374.924 | 7.7805 | 0.12853 | 2.9478 | 230.9414 | -0.99456 | 0.92763 | 0.60548 | 0.038864 | 1 | 0.60548 | 0.56166 | -0.5586 | 0.31204 | 0.5586
[6 empty] | 0.029 | 0.015425 | 7834555794.1943 | 7.8346 | 0.12764 | 2.9683 | 235.8362 | -0.28288 | 0.92532 | 0.59701 | 0.03764 | 1 | 0.59701 | 0.55242 | -0.15627 | 0.024419 | 0.15627
[6 empty] | 0.0292 | 0.015531 | 7888587213.4646 | 7.8886 | 0.12677 | 2.9888 | 240.8114 | 0.85271 | 0.92295 | 0.58849 | 0.036448 | 1 | 0.58849 | 0.54314 | 0.46314 | 0.2145 | 0.46314
[6 empty] | 0.0294 | 0.015637 | 7942618632.7349 | 7.9426 | 0.1259 | 3.0092 | 245.8681 | 0.77951 | 0.92051 | 0.57991 | 0.035285 | 1 | 0.57991 | 0.53382 | 0.41612 | 0.17316 | 0.41612
[6 empty] | 0.0296 | 0.015744 | 7996650052.0052 | 7.9967 | 0.12505 | 3.0297 | 251.0071 | -0.24769 | 0.91802 | 0.57129 | 0.034152 | 1 | 0.57129 | 0.52446 | -0.1299 | 0.016874 | 0.1299
[6 empty] | 0.0298 | 0.01585 | 8050681471.2755 | 8.0507 | 0.12421 | 3.0502 | 256.2294 | -0.96648 | 0.91546 | 0.56263 | 0.033048 | 1 | 0.56263 | 0.51506 | -0.49779 | 0.2478 | 0.49779
[6 empty] | 0.03 | 0.015957 | 8104712890.5458 | 8.1047 | 0.12339 | 3.0707 | 261.5361 | -0.75388 | 0.91284 | 0.55392 | 0.031972 | 1 | 0.55392 | 0.50564 | -0.38119 | 0.14531 | 0.38119
[6 empty] | 0.0302 | 0.016063 | 8158744309.8161 | 8.1587 | 0.12257 | 3.0911 | 266.9281 | 0.037241 | 0.91015 | 0.54518 | 0.030925 | 1 | 0.54518 | 0.49619 | 0.018479 | 0.00034147 | 0.018479
[6 empty] | 0.0304 | 0.016169 | 8212775729.0864 | 8.2128 | 0.12176 | 3.1116 | 272.4064 | 0.74607 | 0.90739 | 0.5364 | 0.029905 | 1 | 0.5364 | 0.48672 | 0.36313 | 0.13186 | 0.36313
[6 empty] | 0.0306 | 0.016276 | 8266807148.3567 | 8.2668 | 0.12097 | 3.1321 | 277.972 | 0.99994 | 0.90457 | 0.52759 | 0.028913 | 1 | 0.52759 | 0.47724 | 0.47722 | 0.22773 | 0.47722
[6 empty] | 0.0308 | 0.016382 | 8320838567.6271 | 8.3208 | 0.12018 | 3.1525 | 283.6259 | 0.81495 | 0.90168 | 0.51876 | 0.027947 | 1 | 0.51876 | 0.46776 | 0.3812 | 0.14531 | 0.3812
[6 empty] | 0.031 | 0.016489 | 8374869986.8974 | 8.3749 | 0.1194 | 3.173 | 289.3692 | 0.40111 | 0.89872 | 0.50991 | 0.027008 | 1 | 0.50991 | 0.45826 | 0.18381 | 0.033788 | 0.18381
[6 empty] | 0.0312 | 0.016595 | 8428901406.1677 | 8.4289 | 0.11864 | 3.1935 | 295.2028 | -0.036805 | 0.89569 | 0.50104 | 0.026094 | 1 | 0.50104 | 0.44877 | -0.016517 | 0.00027282 | 0.016517
[6 empty] | 0.0314 | 0.016701 | 8482932825.438 | 8.4829 | 0.11788 | 3.214 | 301.1279 | -0.38477 | 0.89258 | 0.49215 | 0.025206 | 1 | 0.49215 | 0.43929 | -0.16903 | 0.02857 | 0.16903
[6 empty] | 0.0316 | 0.016808 | 8536964244.7083 | 8.537 | 0.11714 | 3.2344 | 307.1453 | -0.61366 | 0.88941 | 0.48326 | 0.024343 | 1 | 0.48326 | 0.42981 | -0.26376 | 0.069568 | 0.26376
[6 empty] | 0.0318 | 0.016914 | 8590995663.9786 | 8.591 | 0.1164 | 3.2549 | 313.2563 | -0.73991 | 0.88615 | 0.47436 | 0.023504 | 1 | 0.47436 | 0.42036 | -0.31102 | 0.096736 | 0.31102
[6 empty] | 0.032 | 0.01702 | 8645027083.2489 | 8.645 | 0.11567 | 3.2754 | 319.4617 | -0.78991 | 0.88283 | 0.46546 | 0.022689 | 1 | 0.46546 | 0.41092 | -0.32459 | 0.10536 | 0.32459
[6 empty] | 0.0322 | 0.017127 | 8699058502.5192 | 8.6991 | 0.11495 | 3.2958 | 325.7628 | -0.77884 | 0.87943 | 0.45656 | 0.021897 | 1 | 0.45656 | 0.40151 | -0.31272 | 0.097791 | 0.31272
[6 empty] | 0.0324 | 0.017233 | 8753089921.7895 | 8.7531 | 0.11425 | 3.3163 | 332.1604 | -0.70209 | 0.87595 | 0.44767 | 0.021128 | 1 | 0.44767 | 0.39214 | -0.27532 | 0.075799 | 0.27532
[6 empty] | 0.0326 | 0.01734 | 8807121341.0598 | 8.8071 | 0.11354 | 3.3368 | 338.6558 | -0.53637 | 0.87239 | 0.4388 | 0.020382 | 1 | 0.4388 | 0.3828 | -0.20532 | 0.042158 | 0.20532
[6 empty] | 0.0328 | 0.017446 | 8861152760.3301 | 8.8612 | 0.11285 | 3.3573 | 345.25 | -0.25238 | 0.86876 | 0.42994 | 0.019658 | 1 | 0.42994 | 0.37351 | -0.094269 | 0.0088866 | 0.094269
[6 empty] | 0.033 | 0.017552 | 8915184179.6004 | 8.9152 | 0.11217 | 3.3777 | 351.944 | 0.15508 | 0.86504 | 0.4211 | 0.018955 | 1 | 0.4211 | 0.36427 | 0.056489 | 0.0031911 | 0.056489
[6 empty] | 0.0332 | 0.017659 | 8969215598.8707 | 8.9692 | 0.11149 | 3.3982 | 358.739 | 0.61903 | 0.86125 | 0.41228 | 0.018273 | 1 | 0.41228 | 0.35508 | 0.2198 | 0.048314 | 0.2198
[6 empty] | 0.0334 | 0.017765 | 9023247018.141 | 9.0232 | 0.11082 | 3.4187 | 365.6361 | 0.95841 | 0.85737 | 0.40349 | 0.017612 | 1 | 0.40349 | 0.34594 | 0.33155 | 0.10993 | 0.33155
[6 empty] | 0.0336 | 0.017871 | 9077278437.4113 | 9.0773 | 0.11017 | 3.4391 | 372.6363 | 0.90997 | 0.85341 | 0.39474 | 0.016971 | 1 | 0.39474 | 0.33687 | 0.30655 | 0.09397 | 0.30655
[6 empty] | 0.0338 | 0.017978 | 9131309856.6816 | 9.1313 | 0.10951 | 3.4596 | 379.7407 | 0.3164 | 0.84937 | 0.38602 | 0.01635 | 1 | 0.38602 | 0.32788 | 0.10374 | 0.010762 | 0.10374
[6 empty] | 0.034 | 0.018084 | 9185341275.9519 | 9.1853 | 0.10887 | 3.4801 | 386.9506 | -0.56853 | 0.84525 | 0.37735 | 0.015748 | 1 | 0.37735 | 0.31895 | -0.18133 | 0.032882 | 0.18133
[6 empty] | 0.0342 | 0.018191 | 9239372695.2222 | 9.2394 | 0.10823 | 3.5005 | 394.2669 | -0.99775 | 0.84104 | 0.36872 | 0.015165 | 1 | 0.36872 | 0.3101 | -0.30941 | 0.095732 | 0.30941
[6 empty] | 0.0344 | 0.018297 | 9293404114.4926 | 9.2934 | 0.1076 | 3.521 | 401.6909 | -0.35502 | 0.83674 | 0.36014 | 0.014601 | 1 | 0.36014 | 0.30134 | -0.10698 | 0.011445 | 0.10698
[6 empty] | 0.0346 | 0.018403 | 9347435533.7629 | 9.3474 | 0.10698 | 3.5415 | 409.2236 | 0.77494 | 0.83236 | 0.35161 | 0.014054 | 1 | 0.35161 | 0.29267 | 0.2268 | 0.051438 | 0.2268
[6 empty] | 0.0348 | 0.01851 | 9401466953.0332 | 9.4015 | 0.10637 | 3.562 | 416.8663 | 0.78053 | 0.82789 | 0.34314 | 0.013524 | 1 | 0.34314 | 0.28409 | 0.22174 | 0.049167 | 0.22174
[6 empty] | 0.035 | 0.018616 | 9455498372.3035 | 9.4555 | 0.10576 | 3.5824 | 424.6199 | -0.54383 | 0.82334 | 0.33474 | 0.013012 | 1 | 0.33474 | 0.2756 | -0.14988 | 0.022464 | 0.14988
[6 empty] | 0.0352 | 0.018722 | 9509529791.5738 | 9.5095 | 0.10516 | 3.6029 | 432.4859 | -0.83265 | 0.81869 | 0.3264 | 0.012516 | 1 | 0.3264 | 0.26722 | -0.2225 | 0.049506 | 0.2225
[6 empty] | 0.0354 | 0.018829 | 9563561210.8441 | 9.5636 | 0.10456 | 3.6234 | 440.4651 | 0.65352 | 0.81396 | 0.31813 | 0.012037 | 1 | 0.31813 | 0.25894 | 0.16923 | 0.028637 | 0.16923
[6 empty] | 0.0356 | 0.018935 | 9617592630.1144 | 9.6176 | 0.10398 | 3.6438 | 448.559 | 0.58 | 0.80914 | 0.30993 | 0.011573 | 1 | 0.30993 | 0.25078 | 0.14545 | 0.021156 | 0.14545
[6 empty] | 0.0358 | 0.019042 | 9671624049.3847 | 9.6716 | 0.1034 | 3.6643 | 456.7685 | -0.96557 | 0.80423 | 0.30181 | 0.011125 | 1 | 0.30181 | 0.24273 | -0.23437 | 0.054929 | 0.23437
[6 empty] | 0.036 | 0.019148 | 9725655468.655 | 9.7257 | 0.10282 | 3.6848 | 465.095 | 0.20777 | 0.79923 | 0.29377 | 0.010692 | 1 | 0.29377 | 0.23479 | 0.048783 | 0.0023798 | 0.048783
[6 empty] | 0.0362 | 0.019254 | 9779686887.9253 | 9.7797 | 0.10225 | 3.7053 | 473.5395 | 0.69682 | 0.79414 | 0.28582 | 0.010273 | 1 | 0.28582 | 0.22698 | 0.15816 | 0.025016 | 0.15816
[6 empty] | 0.0364 | 0.019361 | 9833718307.1956 | 9.8337 | 0.10169 | 3.7257 | 482.1033 | -0.99813 | 0.78896 | 0.27795 | 0.0098688 | 1 | 0.27795 | 0.21929 | -0.21888 | 0.04791 | 0.21888
[6 empty] | 0.0366 | 0.019467 | 9887749726.4659 | 9.8877 | 0.10114 | 3.7462 | 490.7876 | 0.69555 | 0.78369 | 0.27018 | 0.0094781 | 1 | 0.27018 | 0.21174 | 0.14727 | 0.021689 | 0.14727
[6 empty] | 0.0368 | 0.019573 | 9941781145.7362 | 9.9418 | 0.10059 | 3.7667 | 499.5936 | -0.14982 | 0.77833 | 0.2625 | 0.009101 | 1 | 0.2625 | 0.20431 | -0.03061 | 0.00093698 | 0.03061
[6 empty] | 0.037 | 0.01968 | 9995812565.0065 | 9.9958 | 0.10004 | 3.7871 | 508.5225 | -0.33867 | 0.77288 | 0.25491 | 0.0087369 | 1 | 0.25491 | 0.19702 | -0.066723 | 0.0044519 | 0.066723
[6 empty] | 0.0372 | 0.019786 | 10049843984.2768 | 10.0498 | 0.099504 | 3.8076 | 517.5755 | 0.65732 | 0.76733 | 0.24743 | 0.0083855 | 1 | 0.24743 | 0.18986 | 0.1248 | 0.015575 | 0.1248
[6 empty] | 0.0374 | 0.019893 | 10103875403.5471 | 10.1039 | 0.098972 | 3.8281 | 526.7538 | -0.82129 | 0.7617 | 0.24005 | 0.0080465 | 1 | 0.24005 | 0.18285 | -0.15017 | 0.022551 | 0.15017
[6 empty] | 0.0376 | 0.019999 | 10157906822.8174 | 10.1579 | 0.098445 | 3.8486 | 536.0588 | 0.88359 | 0.75597 | 0.23278 | 0.0077194 | 1 | 0.23278 | 0.17597 | 0.15549 | 0.024177 | 0.15549
[6 empty] | 0.0378 | 0.020105 | 10211938242.0877 | 10.2119 | 0.097925 | 3.869 | 545.4916 | -0.87981 | 0.75016 | 0.22561 | 0.0074041 | 1 | 0.22561 | 0.16925 | -0.1489 | 0.022172 | 0.1489
[6 empty] | 0.038 | 0.020212 | 10265969661.358 | 10.266 | 0.097409 | 3.8895 | 555.0535 | 0.80658 | 0.74425 | 0.21856 | 0.0071 | 1 | 0.21856 | 0.16266 | 0.1312 | 0.017214 | 0.1312
[6 empty] | 0.0382 | 0.020318 | 10320001080.6284 | 10.32 | 0.096899 | 3.91 | 564.7457 | -0.6217 | 0.73825 | 0.21162 | 0.0068069 | 1 | 0.21162 | 0.15623 | -0.097128 | 0.0094338 | 0.097128
[6 empty] | 0.0384 | 0.020424 | 10374032499.8987 | 10.374 | 0.096395 | 3.9304 | 574.5695 | 0.26855 | 0.73216 | 0.2048 | 0.0065245 | 1 | 0.2048 | 0.14994 | 0.040267 | 0.0016214 | 0.040267
[6 empty] | 0.0386 | 0.020531 | 10428063919.169 | 10.4281 | 0.095895 | 3.9509 | 584.5262 | 0.25709 | 0.72599 | 0.19809 | 0.0062524 | 1 | 0.19809 | 0.14381 | 0.036972 | 0.0013669 | 0.036972
[6 empty] | 0.0388 | 0.020637 | 10482095338.4393 | 10.4821 | 0.095401 | 3.9714 | 594.617 | -0.79925 | 0.71972 | 0.1915 | 0.0059903 | 1 | 0.1915 | 0.13783 | -0.11016 | 0.012135 | 0.11016
[6 empty] | 0.039 | 0.020744 | 10536126757.7096 | 10.5361 | 0.094912 | 3.9919 | 604.8432 | 0.98775 | 0.71336 | 0.18504 | 0.005738 | 1 | 0.18504 | 0.132 | 0.13038 | 0.016999 | 0.13038
[6 empty] | 0.0392 | 0.02085 | 10590158176.9799 | 10.5902 | 0.094427 | 4.0123 | 615.2062 | -0.45817 | 0.70692 | 0.17869 | 0.005495 | 1 | 0.17869 | 0.12632 | -0.057877 | 0.0033498 | 0.057877
[6 empty] | 0.0394 | 0.020956 | 10644189596.2502 | 10.6442 | 0.093948 | 4.0328 | 625.7071 | -0.56483 | 0.70039 | 0.17248 | 0.0052612 | 1 | 0.17248 | 0.1208 | -0.068232 | 0.0046556 | 0.068232
[6 empty] | 0.0396 | 0.021063 | 10698221015.5205 | 10.6982 | 0.093473 | 4.0533 | 636.3473 | 0.97016 | 0.69377 | 0.16639 | 0.0050362 | 1 | 0.16639 | 0.11543 | 0.11199 | 0.012542 | 0.11199
[6 empty] | 0.0398 | 0.021169 | 10752252434.7908 | 10.7523 | 0.093004 | 4.0737 | 647.1282 | 0.030165 | 0.68707 | 0.16042 | 0.0048198 | 1 | 0.16042 | 0.11022 | 0.0033248 | 1.1054e-05 | 0.0033248
[6 empty] | 0.04 | 0.021276 | 10806283854.0611 | 10.8063 | 0.092539 | 4.0942 | 658.051 | -0.99909 | 0.68028 | 0.15458 | 0.0046116 | 1 | 0.15458 | 0.10516 | -0.10507 | 0.011039 | 0.10507
[6 empty] | 0.0402 | 0.021382 | 10860315273.3314 | 10.8603 | 0.092078 | 4.1147 | 669.117 | -0.027812 | 0.67341 | 0.14888 | 0.0044115 | 1 | 0.14888 | 0.10026 | -0.0027883 | 7.7749e-06 | 0.0027883
[6 empty] | 0.0404 | 0.021488 | 10914346692.6017 | 10.9143 | 0.091623 | 4.1352 | 680.3276 | 0.97066 | 0.66646 | 0.1433 | 0.0042191 | 1 | 0.1433 | 0.095503 | 0.092701 | 0.0085934 | 0.092701
[6 empty] | 0.0406 | 0.021595 | 10968378111.872 | 10.9684 | 0.091171 | 4.1556 | 691.6841 | 0.56774 | 0.65942 | 0.13785 | 0.0040342 | 1 | 0.13785 | 0.090902 | 0.051609 | 0.0026635 | 0.051609
[6 empty] | 0.0408 | 0.021701 | 11022409531.1423 | 11.0224 | 0.090724 | 4.1761 | 703.1878 | -0.44293 | 0.65231 | 0.13253 | 0.0038566 | 1 | 0.13253 | 0.086453 | -0.038293 | 0.0014663 | 0.038293
[6 empty] | 0.041 | 0.021807 | 11076440950.4126 | 11.0764 | 0.090282 | 4.1966 | 714.8402 | -0.98048 | 0.64512 | 0.12735 | 0.0036859 | 1 | 0.12735 | 0.082154 | -0.080551 | 0.0064884 | 0.080551
[6 empty] | 0.0412 | 0.021914 | 11130472369.6829 | 11.1305 | 0.089843 | 4.217 | 726.6425 | -0.84395 | 0.63785 | 0.12229 | 0.0035221 | 1 | 0.12229 | 0.078004 | -0.065832 | 0.0043338 | 0.065832
[6 empty] | 0.0414 | 0.02202 | 11184503788.9532 | 11.1845 | 0.089409 | 4.2375 | 738.5961 | -0.38189 | 0.6305 | 0.11737 | 0.0033647 | 1 | 0.11737 | 0.074001 | -0.02826 | 0.00079862 | 0.02826
[6 empty] | 0.0416 | 0.022127 | 11238535208.2235 | 11.2385 | 0.08898 | 4.258 | 750.7023 | 0.068209 | 0.62308 | 0.11257 | 0.0032137 | 1 | 0.11257 | 0.070143 | 0.0047844 | 2.2891e-05 | 0.0047844
[6 empty] | 0.0418 | 0.022233 | 11292566627.4939 | 11.2926 | 0.088554 | 4.2784 | 762.9626 | 0.36563 | 0.61559 | 0.10791 | 0.0030688 | 1 | 0.10791 | 0.066429 | 0.024289 | 0.00058994 | 0.024289
[6 empty] | 0.042 | 0.022339 | 11346598046.7642 | 11.3466 | 0.088132 | 4.2989 | 775.3784 | 0.50115 | 0.60804 | 0.10338 | 0.0029298 | 1 | 0.10338 | 0.062857 | 0.031501 | 0.00099232 | 0.031501
[6 empty] | 0.0422 | 0.022446 | 11400629466.0345 | 11.4006 | 0.087714 | 4.3194 | 787.951 | 0.49578 | 0.60041 | 0.098973 | 0.0027965 | 1 | 0.098973 | 0.059424 | 0.029461 | 0.00086798 | 0.029461
[6 empty] | 0.0424 | 0.022552 | 11454660885.3048 | 11.4547 | 0.087301 | 4.3399 | 800.6817 | 0.34696 | 0.59271 | 0.094698 | 0.0026686 | 1 | 0.094698 | 0.056129 | 0.019474 | 0.00037925 | 0.019474
[6 empty] | 0.0426 | 0.022658 | 11508692304.5751 | 11.5087 | 0.086891 | 4.3603 | 813.5721 | 0.030304 | 0.58496 | 0.09055 | 0.002546 | 1 | 0.09055 | 0.052968 | 0.0016051 | 2.5764e-06 | 0.0016051
[6 empty] | 0.0428 | 0.022765 | 11562723723.8454 | 11.5627 | 0.086485 | 4.3808 | 826.6235 | -0.43922 | 0.57714 | 0.086529 | 0.0024285 | 1 | 0.086529 | 0.049939 | -0.021934 | 0.00048111 | 0.021934
[6 empty] | 0.043 | 0.022871 | 11616755143.1157 | 11.6168 | 0.086083 | 4.4013 | 839.8374 | -0.8922 | 0.56926 | 0.082633 | 0.0023159 | 1 | 0.082633 | 0.04704 | -0.041969 | 0.0017614 | 0.041969
[6 empty] | 0.0432 | 0.022978 | 11670786562.386 | 11.6708 | 0.085684 | 4.4217 | 853.2151 | -0.94185 | 0.56133 | 0.078862 | 0.0022081 | 1 | 0.078862 | 0.044268 | -0.041693 | 0.0017383 | 0.041693
[6 empty] | 0.0434 | 0.023084 | 11724817981.6563 | 11.7248 | 0.085289 | 4.4422 | 866.758 | -0.24884 | 0.55334 | 0.075214 | 0.0021048 | 1 | 0.075214 | 0.041619 | -0.010357 | 0.00010726 | 0.010357
[6 empty] | 0.0436 | 0.02319 | 11778849400.9266 | 11.7788 | 0.084898 | 4.4627 | 880.4677 | 0.77822 | 0.54531 | 0.071687 | 0.0020059 | 1 | 0.071687 | 0.039091 | 0.030422 | 0.00092548 | 0.030422
[6 empty] | 0.0438 | 0.023297 | 11832880820.1969 | 11.8329 | 0.08451 | 4.4832 | 894.3455 | 0.80656 | 0.53722 | 0.06828 | 0.0019112 | 1 | 0.06828 | 0.036682 | 0.029586 | 0.00087533 | 0.029586
[6 empty] | 0.044 | 0.023403 | 11886912239.4672 | 11.8869 | 0.084126 | 4.5036 | 908.393 | -0.51649 | 0.52909 | 0.064992 | 0.0018206 | 1 | 0.064992 | 0.034386 | -0.01776 | 0.00031542 | 0.01776
[6 empty] | 0.0442 | 0.023509 | 11940943658.7375 | 11.9409 | 0.083745 | 4.5241 | 922.6114 | -0.81152 | 0.52092 | 0.061819 | 0.0017339 | 1 | 0.061819 | 0.032203 | -0.026133 | 0.00068294 | 0.026133
[6 empty] | 0.0444 | 0.023616 | 11994975078.0078 | 11.995 | 0.083368 | 4.5446 | 937.0024 | 0.76938 | 0.5127 | 0.058762 | 0.0016509 | 1 | 0.058762 | 0.030127 | 0.023179 | 0.00053729 | 0.023179
[6 empty] | 0.0446 | 0.023722 | 12049006497.2781 | 12.049 | 0.082994 | 4.565 | 951.5674 | 0.26203 | 0.50445 | 0.055817 | 0.0015716 | 1 | 0.055817 | 0.028157 | 0.007378 | 5.4435e-05 | 0.007378
[6 empty] | 0.0448 | 0.023829 | 12103037916.5484 | 12.103 | 0.082624 | 4.5855 | 966.3078 | -0.94338 | 0.49617 | 0.052983 | 0.0014958 | 1 | 0.052983 | 0.026289 | -0.0248 | 0.00061504 | 0.0248
[6 empty] | 0.045 | 0.023935 | 12157069335.8187 | 12.1571 | 0.082257 | 4.606 | 981.2252 | 0.89936 | 0.48786 | 0.050257 | 0.0014233 | 1 | 0.050257 | 0.024518 | 0.022051 | 0.00048624 | 0.022051
[6 empty] | 0.0452 | 0.024041 | 12211100755.089 | 12.2111 | 0.081893 | 4.6265 | 996.3209 | -0.48477 | 0.47952 | 0.047638 | 0.001354 | 1 | 0.047638 | 0.022843 | -0.011074 | 0.00012263 | 0.011074
[6 empty] | 0.0454 | 0.024148 | 12265132174.3593 | 12.2651 | 0.081532 | 4.6469 | 1011.5966 | 0.07374 | 0.47116 | 0.045123 | 0.0012878 | 1 | 0.045123 | 0.02126 | 0.0015677 | 2.4577e-06 | 0.0015677
[6 empty] | 0.0456 | 0.024254 | 12319163593.6297 | 12.3192 | 0.081174 | 4.6674 | 1027.0537 | 0.17615 | 0.46278 | 0.04271 | 0.0012246 | 1 | 0.04271 | 0.019765 | 0.0034817 | 1.2122e-05 | 0.0034817
[6 empty] | 0.0458 | 0.02436 | 12373195012.9 | 12.3732 | 0.08082 | 4.6879 | 1042.6937 | -0.2426 | 0.45438 | 0.040397 | 0.0011642 | 1 | 0.040397 | 0.018356 | -0.0044531 | 1.983e-05 | 0.0044531
[6 empty] | 0.046 | 0.024467 | 12427226432.1703 | 12.4272 | 0.080468 | 4.7083 | 1058.5181 | 0.12823 | 0.44597 | 0.038181 | 0.0011066 | 1 | 0.038181 | 0.017027 | 0.0021835 | 4.7678e-06 | 0.0021835
[6 empty] | 0.0462 | 0.024573 | 12481257851.4406 | 12.4813 | 0.08012 | 4.7288 | 1074.5284 | 0.17294 | 0.43755 | 0.03606 | 0.0010515 | 1 | 0.03606 | 0.015778 | 0.0027287 | 7.4457e-06 | 0.0027287
[6 empty] | 0.0464 | 0.02468 | 12535289270.7109 | 12.5353 | 0.079775 | 4.7493 | 1090.7263 | -0.61603 | 0.42913 | 0.034031 | 0.00099902 | 1 | 0.034031 | 0.014604 | -0.0089963 | 8.0933e-05 | 0.0089963
[6 empty] | 0.0466 | 0.024786 | 12589320689.9812 | 12.5893 | 0.079432 | 4.7698 | 1107.1131 | 0.97407 | 0.42071 | 0.032092 | 0.00094892 | 1 | 0.032092 | 0.013501 | 0.013151 | 0.00017295 | 0.013151
[6 empty] | 0.0468 | 0.024892 | 12643352109.2515 | 12.6434 | 0.079093 | 4.7902 | 1123.6905 | -0.80137 | 0.41229 | 0.03024 | 0.00090114 | 1 | 0.03024 | 0.012468 | -0.0099913 | 9.9826e-05 | 0.0099913
[6 empty] | 0.047 | 0.024999 | 12697383528.5218 | 12.6974 | 0.078756 | 4.8107 | 1140.46 | -0.13155 | 0.40387 | 0.028474 | 0.00085557 | 1 | 0.028474 | 0.0115 | -0.0015128 | 2.2887e-06 | 0.0015128
[6 empty] | 0.0472 | 0.025105 | 12751414947.7921 | 12.7514 | 0.078423 | 4.8312 | 1157.4231 | 0.98317 | 0.39547 | 0.02679 | 0.00081212 | 1 | 0.02679 | 0.010595 | 0.010416 | 0.0001085 | 0.010416
[6 empty] | 0.0474 | 0.025211 | 12805446367.0624 | 12.8054 | 0.078092 | 4.8516 | 1174.5815 | -0.29946 | 0.38708 | 0.025186 | 0.00077071 | 1 | 0.025186 | 0.0097491 | -0.0029195 | 8.5234e-06 | 0.0029195
[6 empty] | 0.0476 | 0.025318 | 12859477786.3327 | 12.8595 | 0.077764 | 4.8721 | 1191.9366 | -0.97417 | 0.37872 | 0.023659 | 0.00073125 | 1 | 0.023659 | 0.0089602 | -0.0087287 | 7.619e-05 | 0.0087287
[6 empty] | 0.0478 | 0.025424 | 12913509205.603 | 12.9135 | 0.077438 | 4.8926 | 1209.49 | -0.046868 | 0.37037 | 0.022208 | 0.00069366 | 1 | 0.022208 | 0.0082251 | -0.00038549 | 1.4861e-07 | 0.00038549
[6 empty] | 0.048 | 0.025531 | 12967540624.8733 | 12.9675 | 0.077116 | 4.913 | 1227.2434 | 0.86708 | 0.36205 | 0.020828 | 0.00065785 | 1 | 0.020828 | 0.007541 | 0.0065386 | 4.2754e-05 | 0.0065386
[6 empty] | 0.0482 | 0.025637 | 13021572044.1436 | 13.0216 | 0.076796 | 4.9335 | 1245.1982 | 0.93116 | 0.35377 | 0.019519 | 0.00062376 | 1 | 0.019519 | 0.0069051 | 0.0064298 | 4.1342e-05 | 0.0064298
[6 empty] | 0.0484 | 0.025743 | 13075603463.4139 | 13.0756 | 0.076478 | 4.954 | 1263.3562 | 0.48467 | 0.34552 | 0.018277 | 0.0005913 | 1 | 0.018277 | 0.006315 | 0.0030607 | 9.3679e-06 | 0.0030607
[6 empty] | 0.0486 | 0.02585 | 13129634882.6842 | 13.1296 | 0.076164 | 4.9745 | 1281.7188 | 0.019088 | 0.3373 | 0.0171 | 0.0005604 | 1 | 0.0171 | 0.0057679 | 0.0001101 | 1.2122e-08 | 0.0001101
[6 empty] | 0.0488 | 0.025956 | 13183666301.9545 | 13.1837 | 0.075851 | 4.9949 | 1300.2878 | -0.25854 | 0.32913 | 0.015986 | 0.00053101 | 1 | 0.015986 | 0.0052614 | -0.0013603 | 1.8504e-06 | 0.0013603
[6 empty] | 0.049 | 0.026062 | 13237697721.2248 | 13.2377 | 0.075542 | 5.0154 | 1319.0647 | -0.32801 | 0.32101 | 0.014931 | 0.00050304 | 1 | 0.014931 | 0.0047932 | -0.0015722 | 2.4718e-06 | 0.0015722
[6 empty] | 0.0492 | 0.026169 | 13291729140.4951 | 13.2917 | 0.075235 | 5.0359 | 1338.0511 | -0.19607 | 0.31294 | 0.013935 | 0.00047644 | 1 | 0.013935 | 0.0043609 | -0.00085503 | 7.3107e-07 | 0.00085503
[6 empty] | 0.0494 | 0.026275 | 13345760559.7655 | 13.3458 | 0.07493 | 5.0563 | 1357.2486 | 0.15009 | 0.30493 | 0.012994 | 0.00045115 | 1 | 0.012994 | 0.0039623 | 0.00059471 | 3.5368e-07 | 0.00059471
[6 empty] | 0.0496 | 0.026382 | 13399791979.0358 | 13.3998 | 0.074628 | 5.0768 | 1376.659 | 0.65294 | 0.29698 | 0.012106 | 0.0004271 | 1 | 0.012106 | 0.0035953 | 0.0023475 | 5.5109e-06 | 0.0023475
[6 empty] | 0.0498 | 0.026488 | 13453823398.3061 | 13.4538 | 0.074328 | 5.0973 | 1396.2838 | 0.99646 | 0.28909 | 0.01127 | 0.00040425 | 1 | 0.01127 | 0.0032579 | 0.0032464 | 1.0539e-05 | 0.0032464
[6 empty] | 0.05 | 0.026594 | 13507854817.5764 | 13.5079 | 0.074031 | 5.1178 | 1416.1246 | 0.61597 | 0.28126 | 0.010482 | 0.00038254 | 1 | 0.010482 | 0.0029481 | 0.0018159 | 3.2976e-06 | 0.0018159
[6 empty] | 0.0502 | 0.026701 | 13561886236.8467 | 13.5619 | 0.073736 | 5.1382 | 1436.1832 | -0.51877 | 0.27351 | 0.00974 | 0.00036191 | 1 | 0.00974 | 0.002664 | -0.001382 | 1.91e-06 | 0.001382
[6 empty] | 0.0504 | 0.026807 | 13615917656.117 | 13.6159 | 0.073443 | 5.1587 | 1456.4612 | -0.91988 | 0.26584 | 0.009043 | 0.00034232 | 1 | 0.009043 | 0.002404 | -0.0022114 | 4.8902e-06 | 0.0022114
[6 empty] | 0.0506 | 0.026914 | 13669949075.3873 | 13.6699 | 0.073153 | 5.1792 | 1476.9602 | 0.46327 | 0.25824 | 0.0083884 | 0.00032371 | 1 | 0.0083884 | 0.0021662 | 0.0010035 | 1.0071e-06 | 0.0010035
[6 empty] | 0.0508 | 0.02702 | 13723980494.6576 | 13.724 | 0.072865 | 5.1996 | 1497.6819 | 0.70876 | 0.25073 | 0.0077742 | 0.00030605 | 1 | 0.0077742 | 0.0019492 | 0.0013815 | 1.9086e-06 | 0.0013815
[6 empty] | 0.051 | 0.027126 | 13778011913.9279 | 13.778 | 0.072579 | 5.2201 | 1518.628 | -0.96589 | 0.2433 | 0.0071985 | 0.00028929 | 1 | 0.0071985 | 0.0017514 | -0.0016916 | 2.8616e-06 | 0.0016916
[6 empty] | 0.0512 | 0.027233 | 13832043333.1982 | 13.832 | 0.072296 | 5.2406 | 1539.8003 | 0.47054 | 0.23596 | 0.0066594 | 0.00027339 | 1 | 0.0066594 | 0.0015713 | 0.00073939 | 5.4669e-07 | 0.00073939
[6 empty] | 0.0514 | 0.027339 | 13886074752.4685 | 13.8861 | 0.072015 | 5.2611 | 1561.2002 | 0.10108 | 0.22871 | 0.0061549 | 0.00025831 | 1 | 0.0061549 | 0.0014077 | 0.00014229 | 2.0246e-08 | 0.00014229
[6 empty] | 0.0516 | 0.027445 | 13940106171.7388 | 13.9401 | 0.071735 | 5.2815 | 1582.8297 | -0.44658 | 0.22157 | 0.0056835 | 0.000244 | 1 | 0.0056835 | 0.0012593 | -0.00056236 | 3.1624e-07 | 0.00056236
[6 empty] | 0.0518 | 0.027552 | 13994137591.0091 | 13.9941 | 0.071458 | 5.302 | 1604.6903 | 0.55922 | 0.21452 | 0.0052432 | 0.00023043 | 1 | 0.0052432 | 0.0011248 | 0.00062898 | 3.9562e-07 | 0.00062898
[6 empty] | 0.052 | 0.027658 | 14048169010.2794 | 14.0482 | 0.071184 | 5.3225 | 1626.7839 | -0.47156 | 0.20757 | 0.0048325 | 0.00021757 | 1 | 0.0048325 | 0.0010031 | -0.00047301 | 2.2374e-07 | 0.00047301
[6 empty] | 0.0522 | 0.027765 | 14102200429.5497 | 14.1022 | 0.070911 | 5.3429 | 1649.112 | 0.15346 | 0.20073 | 0.0044498 | 0.00020538 | 1 | 0.0044498 | 0.0008932 | 0.00013707 | 1.8788e-08 | 0.00013707
[6 empty] | 0.0524 | 0.027871 | 14156231848.82 | 14.1562 | 0.07064 | 5.3634 | 1671.6765 | 0.40708 | 0.19399 | 0.0040935 | 0.00019384 | 1 | 0.0040935 | 0.00079411 | 0.00032327 | 1.045e-07 | 0.00032327
[6 empty] | 0.0526 | 0.027977 | 14210263268.0903 | 14.2103 | 0.070372 | 5.3839 | 1694.479 | -0.9427 | 0.18737 | 0.0037621 | 0.0001829 | 1 | 0.0037621 | 0.00070491 | -0.00066452 | 4.4159e-07 | 0.00066452
[6 empty] | 0.0528 | 0.028084 | 14264294687.3606 | 14.2643 | 0.070105 | 5.4044 | 1717.5214 | 0.75769 | 0.18086 | 0.0034541 | 0.00017254 | 1 | 0.0034541 | 0.00062473 | 0.00047335 | 2.2406e-07 | 0.00047335
[6 empty] | 0.053 | 0.02819 | 14318326106.631 | 14.3183 | 0.069841 | 5.4248 | 1740.8052 | 0.41957 | 0.17447 | 0.0031683 | 0.00016273 | 1 | 0.0031683 | 0.00055279 | 0.00023193 | 5.3794e-08 | 0.00023193
[6 empty] | 0.0532 | 0.028296 | 14372357525.9013 | 14.3724 | 0.069578 | 5.4453 | 1764.3324 | -0.92176 | 0.1682 | 0.0029033 | 0.00015344 | 1 | 0.0029033 | 0.00048833 | -0.00045013 | 2.0261e-07 | 0.00045013
[6 empty] | 0.0534 | 0.028403 | 14426388945.1716 | 14.4264 | 0.069317 | 5.4658 | 1788.1047 | -0.57164 | 0.16205 | 0.0026578 | 0.00014466 | 1 | 0.0026578 | 0.00043069 | -0.0002462 | 6.0613e-08 | 0.0002462
[6 empty] | 0.0536 | 0.028509 | 14480420364.4419 | 14.4804 | 0.069059 | 5.4862 | 1812.1237 | 0.48397 | 0.15602 | 0.0024307 | 0.00013634 | 1 | 0.0024307 | 0.00037922 | 0.00018353 | 3.3684e-08 | 0.00018353
[6 empty] | 0.0538 | 0.028616 | 14534451783.7122 | 14.5345 | 0.068802 | 5.5067 | 1836.3914 | 0.97996 | 0.15011 | 0.0022207 | 0.00012848 | 1 | 0.0022207 | 0.00033334 | 0.00032667 | 1.0671e-07 | 0.00032667
[6 empty] | 0.054 | 0.028722 | 14588483202.9825 | 14.5885 | 0.068547 | 5.5272 | 1860.9094 | 0.91543 | 0.14433 | 0.0020268 | 0.00012104 | 1 | 0.0020268 | 0.00029252 | 0.00026779 | 7.1709e-08 | 0.00026779
[6 empty] | 0.0542 | 0.028828 | 14642514622.2528 | 14.6425 | 0.068294 | 5.5477 | 1885.6795 | 0.71316 | 0.13868 | 0.0018479 | 0.00011401 | 1 | 0.0018479 | 0.00025626 | 0.00018276 | 3.3401e-08 | 0.00018276
[6 empty] | 0.0544 | 0.028935 | 14696546041.5231 | 14.6965 | 0.068043 | 5.5681 | 1910.7036 | 0.63296 | 0.13315 | 0.0016831 | 0.00010736 | 1 | 0.0016831 | 0.00022411 | 0.00014185 | 2.0122e-08 | 0.00014185
[6 empty] | 0.0546 | 0.029041 | 14750577460.7934 | 14.7506 | 0.067794 | 5.5886 | 1935.9835 | 0.7396 | 0.12776 | 0.0015314 | 0.00010108 | 1 | 0.0015314 | 0.00019565 | 0.0001447 | 2.0939e-08 | 0.0001447
[6 empty] | 0.0548 | 0.029147 | 14804608880.0637 | 14.8046 | 0.067547 | 5.6091 | 1961.5209 | 0.94485 | 0.12249 | 0.001392 | 9.514e-05 | 1 | 0.001392 | 0.00017051 | 0.0001611 | 2.5954e-08 | 0.0001611
[6 empty] | 0.055 | 0.029254 | 14858640299.334 | 14.8586 | 0.067301 | 5.6295 | 1987.3176 | 0.94593 | 0.11736 | 0.0012639 | 8.9533e-05 | 1 | 0.0012639 | 0.00014833 | 0.00014031 | 1.9687e-08 | 0.00014031
[6 empty] | 0.0552 | 0.02936 | 14912671718.6043 | 14.9127 | 0.067057 | 5.65 | 2013.3756 | 0.31009 | 0.11236 | 0.0011463 | 8.4237e-05 | 1 | 0.0011463 | 0.00012881 | 3.9941e-05 | 1.5953e-09 | 3.9941e-05
[6 empty] | 0.0554 | 0.029467 | 14966703137.8746 | 14.9667 | 0.066815 | 5.6705 | 2039.6965 | -0.76619 | 0.10749 | 0.0010386 | 7.9238e-05 | 1 | 0.0010386 | 0.00011165 | -8.5544e-05 | 7.3178e-09 | 8.5544e-05
[6 empty] | 0.0556 | 0.029573 | 15020734557.1449 | 15.0207 | 0.066575 | 5.6909 | 2066.2822 | -0.72821 | 0.10276 | 0.00094004 | 7.4518e-05 | 1 | 0.00094004 | 9.66e-05 | -7.0344e-05 | 4.9483e-09 | 7.0344e-05
[6 empty] | 0.0558 | 0.029679 | 15074765976.4152 | 15.0748 | 0.066336 | 5.7114 | 2093.1346 | 0.78579 | 0.098161 | 0.00084988 | 7.0064e-05 | 1 | 0.00084988 | 8.3425e-05 | 6.5554e-05 | 4.2974e-09 | 6.5554e-05
[6 empty] | 0.056 | 0.029786 | 15128797395.6855 | 15.1288 | 0.066099 | 5.7319 | 2120.2555 | 0.24687 | 0.093693 | 0.00076754 | 6.5862e-05 | 1 | 0.00076754 | 7.1913e-05 | 1.7753e-05 | 3.1518e-10 | 1.7753e-05
[6 empty] | 0.0562 | 0.029892 | 15182828814.9558 | 15.1828 | 0.065864 | 5.7524 | 2147.6468 | -0.90548 | 0.089359 | 0.00069241 | 6.1898e-05 | 1 | 0.00069241 | 6.1872e-05 | -5.6024e-05 | 3.1387e-09 | 5.6024e-05
[6 empty] | 0.0564 | 0.029998 | 15236860234.2261 | 15.2369 | 0.06563 | 5.7728 | 2175.3103 | 0.98515 | 0.085156 | 0.00062394 | 5.816e-05 | 1 | 0.00062394 | 5.3133e-05 | 5.2343e-05 | 2.7398e-09 | 5.2343e-05
[6 empty] | 0.0566 | 0.030105 | 15290891653.4964 | 15.2909 | 0.065398 | 5.7933 | 2203.2479 | -0.87307 | 0.081086 | 0.00056162 | 5.4636e-05 | 1 | 0.00056162 | 4.554e-05 | -3.9759e-05 | 1.5808e-09 | 3.9759e-05
[6 empty] | 0.0568 | 0.030211 | 15344923072.7668 | 15.3449 | 0.065168 | 5.8138 | 2231.4614 | 0.84182 | 0.077147 | 0.00050496 | 5.1313e-05 | 1 | 0.00050496 | 3.8956e-05 | 3.2794e-05 | 1.0754e-09 | 3.2794e-05
[6 empty] | 0.057 | 0.030318 | 15398954492.0371 | 15.399 | 0.064939 | 5.8342 | 2259.9527 | -0.93829 | 0.073339 | 0.0004535 | 4.8182e-05 | 1 | 0.0004535 | 3.3259e-05 | -3.1207e-05 | 9.7385e-10 | 3.1207e-05
[6 empty] | 0.0572 | 0.030424 | 15452985911.3074 | 15.453 | 0.064712 | 5.8547 | 2288.7238 | 0.98971 | 0.069659 | 0.00040682 | 4.5232e-05 | 1 | 0.00040682 | 2.8339e-05 | 2.8047e-05 | 7.8664e-10 | 2.8047e-05
[6 empty] | 0.0574 | 0.03053 | 15507017330.5777 | 15.507 | 0.064487 | 5.8752 | 2317.7764 | -0.6043 | 0.066108 | 0.00036453 | 4.2454e-05 | 1 | 0.00036453 | 2.4098e-05 | -1.4563e-05 | 2.1207e-10 | 1.4563e-05
[6 empty] | 0.0576 | 0.030637 | 15561048749.848 | 15.561 | 0.064263 | 5.8957 | 2347.1126 | -0.40129 | 0.062683 | 0.00032626 | 3.9837e-05 | 1 | 0.00032626 | 2.0451e-05 | -8.2067e-06 | 6.735e-11 | 8.2067e-06
[6 empty] | 0.0578 | 0.030743 | 15615080169.1183 | 15.6151 | 0.064041 | 5.9161 | 2376.7341 | 0.98213 | 0.059384 | 0.00029167 | 3.7373e-05 | 1 | 0.00029167 | 1.732e-05 | 1.7011e-05 | 2.8936e-10 | 1.7011e-05
[6 empty] | 0.058 | 0.030849 | 15669111588.3886 | 15.6691 | 0.06382 | 5.9366 | 2406.643 | 0.25038 | 0.056208 | 0.00026044 | 3.5054e-05 | 1 | 0.00026044 | 1.4639e-05 | 3.6653e-06 | 1.3434e-11 | 3.6653e-06
[6 empty] | 0.0582 | 0.030956 | 15723143007.6589 | 15.7231 | 0.063601 | 5.9571 | 2436.8411 | -0.82191 | 0.053155 | 0.00023228 | 3.2871e-05 | 1 | 0.00023228 | 1.2347e-05 | -1.0148e-05 | 1.0298e-10 | 1.0148e-05
[6 empty] | 0.0584 | 0.031062 | 15777174426.9292 | 15.7772 | 0.063383 | 5.9775 | 2467.3303 | -0.94903 | 0.050221 | 0.00020692 | 3.0818e-05 | 1 | 0.00020692 | 1.0392e-05 | -9.8619e-06 | 9.7257e-11 | 9.8619e-06
[6 empty] | 0.0586 | 0.031169 | 15831205846.1995 | 15.8312 | 0.063166 | 5.998 | 2498.1126 | -0.57821 | 0.047405 | 0.00018411 | 2.8886e-05 | 1 | 0.00018411 | 8.7276e-06 | -5.0464e-06 | 2.5466e-11 | 5.0464e-06
[6 empty] | 0.0588 | 0.031275 | 15885237265.4698 | 15.8852 | 0.062952 | 6.0185 | 2529.19 | -0.27437 | 0.044705 | 0.00016361 | 2.707e-05 | 1 | 0.00016361 | 7.3144e-06 | -2.0069e-06 | 4.0276e-12 | 2.0069e-06
[6 empty] | 0.059 | 0.031381 | 15939268684.7401 | 15.9393 | 0.062738 | 6.039 | 2560.5643 | -0.23411 | 0.04212 | 0.00014523 | 2.5362e-05 | 1 | 0.00014523 | 6.1168e-06 | -1.432e-06 | 2.0507e-12 | 1.432e-06
[6 empty] | 0.0592 | 0.031488 | 15993300104.0104 | 15.9933 | 0.062526 | 6.0594 | 2592.2375 | -0.47378 | 0.039645 | 0.00012875 | 2.3757e-05 | 1 | 0.00012875 | 5.1042e-06 | -2.4183e-06 | 5.848e-12 | 2.4183e-06
[6 empty] | 0.0594 | 0.031594 | 16047331523.2807 | 16.0473 | 0.062316 | 6.0799 | 2624.2115 | -0.86828 | 0.037279 | 0.000114 | 2.2248e-05 | 1 | 0.000114 | 4.2498e-06 | -3.69e-06 | 1.3616e-11 | 3.69e-06
[6 empty] | 0.0596 | 0.031701 | 16101362942.551 | 16.1014 | 0.062107 | 6.1004 | 2656.4884 | -0.94212 | 0.03502 | 0.00010081 | 2.083e-05 | 1 | 0.00010081 | 3.5306e-06 | -3.3262e-06 | 1.1064e-11 | 3.3262e-06
[6 empty] | 0.0598 | 0.031807 | 16155394361.8213 | 16.1554 | 0.061899 | 6.1208 | 2689.0701 | -0.063088 | 0.032865 | 8.9045e-05 | 1.9499e-05 | 1 | 8.9045e-05 | 2.9265e-06 | -1.8463e-07 | 3.4087e-14 | 1.8463e-07
[6 empty] | 0.06 | 0.031913 | 16209425781.0916 | 16.2094 | 0.061693 | 6.1413 | 2721.9586 | 0.98701 | 0.030812 | 7.8551e-05 | 1.8248e-05 | 1 | 7.8551e-05 | 2.4203e-06 | 2.3889e-06 | 5.7066e-12 | 2.3889e-06
[6 empty] | 0.0602 | 0.03202 | 16263457200.3619 | 16.2635 | 0.061488 | 6.1618 | 2755.1558 | -0.049102 | 0.028857 | 6.9206e-05 | 1.7074e-05 | 1 | 6.9206e-05 | 1.9971e-06 | -9.806e-08 | 9.6158e-15 | 9.806e-08
[6 empty] | 0.0604 | 0.032126 | 16317488619.6322 | 16.3175 | 0.061284 | 6.1823 | 2788.6638 | -0.8417 | 0.026998 | 6.0896e-05 | 1.5972e-05 | 1 | 6.0896e-05 | 1.6441e-06 | -1.3838e-06 | 1.9149e-12 | 1.3838e-06
[6 empty] | 0.0606 | 0.032232 | 16371520038.9026 | 16.3715 | 0.061082 | 6.2027 | 2822.4845 | 0.98618 | 0.025232 | 5.3515e-05 | 1.4938e-05 | 1 | 5.3515e-05 | 1.3503e-06 | 1.3317e-06 | 1.7733e-12 | 1.3317e-06
[6 empty] | 0.0608 | 0.032339 | 16425551458.1729 | 16.4256 | 0.060881 | 6.2232 | 2856.62 | -0.83178 | 0.023557 | 4.6969e-05 | 1.3968e-05 | 1 | 4.6969e-05 | 1.1064e-06 | -9.2031e-07 | 8.4698e-13 | 9.2031e-07
[6 empty] | 0.061 | 0.032445 | 16479582877.4432 | 16.4796 | 0.060681 | 6.2437 | 2891.0722 | 0.76883 | 0.021969 | 4.117e-05 | 1.3058e-05 | 1 | 4.117e-05 | 9.0446e-07 | 6.9537e-07 | 4.8354e-13 | 6.9537e-07
[6 empty] | 0.0612 | 0.032552 | 16533614296.7135 | 16.5336 | 0.060483 | 6.2641 | 2925.8432 | -0.88685 | 0.020466 | 3.604e-05 | 1.2204e-05 | 1 | 3.604e-05 | 7.3758e-07 | -6.5412e-07 | 4.2787e-13 | 6.5412e-07
[6 empty] | 0.0614 | 0.032658 | 16587645715.9838 | 16.5876 | 0.060286 | 6.2846 | 2960.935 | 0.99855 | 0.019044 | 3.1507e-05 | 1.1404e-05 | 1 | 3.1507e-05 | 6.0004e-07 | 5.9916e-07 | 3.59e-13 | 5.9916e-07
[6 empty] | 0.0616 | 0.032764 | 16641677135.2541 | 16.6417 | 0.06009 | 6.3051 | 2996.3496 | -0.61293 | 0.017702 | 2.7509e-05 | 1.0654e-05 | 1 | 2.7509e-05 | 4.8695e-07 | -2.9847e-07 | 8.9085e-14 | 2.9847e-07
[6 empty] | 0.0618 | 0.032871 | 16695708554.5244 | 16.6957 | 0.059896 | 6.3256 | 3032.089 | -0.49877 | 0.016435 | 2.3986e-05 | 9.9506e-06 | 1 | 2.3986e-05 | 3.9421e-07 | -1.9662e-07 | 3.8659e-14 | 1.9662e-07
[6 empty] | 0.062 | 0.032977 | 16749739973.7947 | 16.7497 | 0.059702 | 6.346 | 3068.1554 | 0.89595 | 0.015242 | 2.0886e-05 | 9.2919e-06 | 1 | 2.0886e-05 | 3.1834e-07 | 2.8522e-07 | 8.1349e-14 | 2.8522e-07
[6 empty] | 0.0622 | 0.033083 | 16803771393.065 | 16.8038 | 0.05951 | 6.3665 | 3104.5507 | 0.66479 | 0.014118 | 1.8163e-05 | 8.6749e-06 | 1 | 1.8163e-05 | 2.5642e-07 | 1.7047e-07 | 2.906e-14 | 1.7047e-07
[6 empty] | 0.0624 | 0.03319 | 16857802812.3353 | 16.8578 | 0.05932 | 6.387 | 3141.277 | -0.24311 | 0.013062 | 1.5773e-05 | 8.0971e-06 | 1 | 1.5773e-05 | 2.0603e-07 | -5.0089e-08 | 2.5089e-15 | 5.0089e-08
[6 empty] | 0.0626 | 0.033296 | 16911834231.6056 | 16.9118 | 0.05913 | 6.4074 | 3178.3364 | -0.77411 | 0.012071 | 1.3679e-05 | 7.5561e-06 | 1 | 1.3679e-05 | 1.6512e-07 | -1.2782e-07 | 1.6337e-14 | 1.2782e-07
[6 empty] | 0.0628 | 0.033403 | 16965865650.8759 | 16.9659 | 0.058942 | 6.4279 | 3215.7309 | -0.92833 | 0.011142 | 1.1846e-05 | 7.0496e-06 | 1 | 1.1846e-05 | 1.3199e-07 | -1.2253e-07 | 1.5013e-14 | 1.2253e-07
[6 empty] | 0.063 | 0.033509 | 17019897070.1462 | 17.0199 | 0.058755 | 6.4484 | 3253.4627 | -0.91571 | 0.010271 | 1.0245e-05 | 6.5757e-06 | 1 | 1.0245e-05 | 1.0523e-07 | -9.636e-08 | 9.2853e-15 | 9.636e-08
[6 empty] | 0.0632 | 0.033615 | 17073928489.4165 | 17.0739 | 0.058569 | 6.4688 | 3291.5337 | -0.70707 | 0.0094567 | 8.8485e-06 | 6.1323e-06 | 1 | 8.8485e-06 | 8.3678e-08 | -5.9166e-08 | 3.5007e-15 | 5.9166e-08
[6 empty] | 0.0634 | 0.033722 | 17127959908.6868 | 17.128 | 0.058384 | 6.4893 | 3329.9461 | -0.071986 | 0.0086961 | 7.6314e-06 | 5.7175e-06 | 1 | 7.6314e-06 | 6.6363e-08 | -4.7772e-09 | 2.2822e-17 | 4.7772e-09
[6 empty] | 0.0636 | 0.033828 | 17181991327.9571 | 17.182 | 0.0582 | 6.5098 | 3368.702 | 0.83312 | 0.0079866 | 6.5724e-06 | 5.3295e-06 | 1 | 6.5724e-06 | 5.2491e-08 | 4.3731e-08 | 1.9124e-15 | 4.3731e-08
[6 empty] | 0.0638 | 0.033934 | 17236022747.2275 | 17.236 | 0.058018 | 6.5303 | 3407.8035 | 0.68493 | 0.0073255 | 5.6523e-06 | 4.9668e-06 | 1 | 5.6523e-06 | 4.1406e-08 | 2.836e-08 | 8.0431e-16 | 2.836e-08
[6 empty] | 0.064 | 0.034041 | 17290054166.4978 | 17.2901 | 0.057837 | 6.5507 | 3447.2526 | -0.83906 | 0.0067104 | 4.8542e-06 | 4.6278e-06 | 1 | 4.8542e-06 | 3.2574e-08 | -2.7331e-08 | 7.4699e-16 | 2.7331e-08
[6 empty] | 0.0642 | 0.034147 | 17344085585.7681 | 17.3441 | 0.057657 | 6.5712 | 3487.0516 | -0.046176 | 0.0061389 | 4.1627e-06 | 4.311e-06 | 1 | 4.1627e-06 | 2.5555e-08 | -1.18e-09 | 1.3925e-18 | 1.18e-09
[6 empty] | 0.0644 | 0.034254 | 17398117005.0384 | 17.3981 | 0.057477 | 6.5917 | 3527.2025 | 0.67133 | 0.0056086 | 3.5647e-06 | 4.0149e-06 | 1 | 3.5647e-06 | 1.9993e-08 | 1.3422e-08 | 1.8015e-16 | 1.3422e-08
[6 empty] | 0.0646 | 0.03436 | 17452148424.3087 | 17.4521 | 0.0573 | 6.6121 | 3567.7074 | -0.87805 | 0.0051172 | 3.0481e-06 | 3.7384e-06 | 1 | 3.0481e-06 | 1.5598e-08 | -1.3696e-08 | 1.8757e-16 | 1.3696e-08
[6 empty] | 0.0648 | 0.034466 | 17506179843.579 | 17.5062 | 0.057123 | 6.6326 | 3608.5686 | 0.86809 | 0.0046625 | 2.6025e-06 | 3.4801e-06 | 1 | 2.6025e-06 | 1.2134e-08 | 1.0534e-08 | 1.1096e-16 | 1.0534e-08
[6 empty] | 0.065 | 0.034573 | 17560211262.8493 | 17.5602 | 0.056947 | 6.6531 | 3649.788 | -0.623 | 0.0042424 | 2.2189e-06 | 3.239e-06 | 1 | 2.2189e-06 | 9.4133e-09 | -5.8645e-09 | 3.4392e-17 | 5.8645e-09
[6 empty] | 0.0652 | 0.034679 | 17614242682.1196 | 17.6142 | 0.056772 | 6.6736 | 3691.368 | -0.066624 | 0.0038547 | 1.889e-06 | 3.0139e-06 | 1 | 1.889e-06 | 7.2814e-09 | -4.8512e-10 | 2.3534e-19 | 4.8512e-10
[6 empty] | 0.0654 | 0.034785 | 17668274101.3899 | 17.6683 | 0.056599 | 6.694 | 3733.3106 | 0.92018 | 0.0034975 | 1.6057e-06 | 2.8038e-06 | 1 | 1.6057e-06 | 5.616e-09 | 5.1678e-09 | 2.6706e-17 | 5.1678e-09
[6 empty] | 0.0656 | 0.034892 | 17722305520.6602 | 17.7223 | 0.056426 | 6.7145 | 3775.6179 | -0.485 | 0.0031689 | 1.3629e-06 | 2.6078e-06 | 1 | 1.3629e-06 | 4.3189e-09 | -2.0947e-09 | 4.3876e-18 | 2.0947e-09
[6 empty] | 0.0658 | 0.034998 | 17776336939.9305 | 17.7763 | 0.056255 | 6.735 | 3818.2923 | -0.97049 | 0.002867 | 1.1551e-06 | 2.4249e-06 | 1 | 1.1551e-06 | 3.3115e-09 | -3.2138e-09 | 1.0329e-17 | 3.2138e-09
[6 empty] | 0.066 | 0.035105 | 17830368359.2008 | 17.8304 | 0.056084 | 6.7554 | 3861.3358 | -0.37876 | 0.00259 | 9.7744e-07 | 2.2544e-06 | 1 | 9.7744e-07 | 2.5316e-09 | -9.5888e-10 | 9.1945e-19 | 9.5888e-10
[6 empty] | 0.0662 | 0.035211 | 17884399778.4711 | 17.8844 | 0.055915 | 6.7759 | 3904.7506 | 0.17808 | 0.0023364 | 8.2587e-07 | 2.0954e-06 | 1 | 8.2587e-07 | 1.9296e-09 | 3.4361e-10 | 1.1807e-19 | 3.4361e-10
[6 empty] | 0.0664 | 0.035317 | 17938431197.7414 | 17.9384 | 0.055746 | 6.7964 | 3948.5389 | 0.36443 | 0.0021044 | 6.9673e-07 | 1.9472e-06 | 1 | 6.9673e-07 | 1.4662e-09 | 5.3433e-10 | 2.8551e-19 | 5.3433e-10
[6 empty] | 0.0666 | 0.035424 | 17992462617.0118 | 17.9925 | 0.055579 | 6.8169 | 3992.7029 | 0.19013 | 0.0018926 | 5.8688e-07 | 1.8091e-06 | 1 | 5.8688e-07 | 1.1107e-09 | 2.1119e-10 | 4.4601e-20 | 2.1119e-10
[6 empty] | 0.0668 | 0.03553 | 18046494036.2821 | 18.0465 | 0.055412 | 6.8373 | 4037.2448 | -0.36006 | 0.0016995 | 4.9358e-07 | 1.6804e-06 | 1 | 4.9358e-07 | 8.3886e-10 | -3.0204e-10 | 9.1226e-20 | 3.0204e-10
[6 empty] | 0.067 | 0.035636 | 18100525455.5524 | 18.1005 | 0.055247 | 6.8578 | 4082.1669 | -0.96568 | 0.0015238 | 4.1447e-07 | 1.5605e-06 | 1 | 4.1447e-07 | 6.3155e-10 | -6.0988e-10 | 3.7196e-19 | 6.0988e-10
[6 empty] | 0.0672 | 0.035743 | 18154556874.8227 | 18.1546 | 0.055083 | 6.8783 | 4127.4712 | -0.48945 | 0.001364 | 3.4749e-07 | 1.4488e-06 | 1 | 3.4749e-07 | 4.7399e-10 | -2.3199e-10 | 5.3821e-20 | 2.3199e-10
[6 empty] | 0.0674 | 0.035849 | 18208588294.093 | 18.2086 | 0.054919 | 6.8987 | 4173.1601 | 0.93027 | 0.0012191 | 2.9087e-07 | 1.3449e-06 | 1 | 2.9087e-07 | 3.5461e-10 | 3.2988e-10 | 1.0882e-19 | 3.2988e-10
[6 empty] | 0.0676 | 0.035956 | 18262619713.3633 | 18.2626 | 0.054757 | 6.9192 | 4219.2358 | -0.14643 | 0.0010879 | 2.4309e-07 | 1.2481e-06 | 1 | 2.4309e-07 | 2.6445e-10 | -3.8722e-11 | 1.4994e-21 | 3.8722e-11
[6 empty] | 0.0678 | 0.036062 | 18316651132.6336 | 18.3167 | 0.054595 | 6.9397 | 4265.7005 | -0.49013 | 0.00096914 | 2.0284e-07 | 1.158e-06 | 1 | 2.0284e-07 | 1.9658e-10 | -9.6349e-11 | 9.2831e-21 | 9.6349e-11
[6 empty] | 0.068 | 0.036168 | 18370682551.9039 | 18.3707 | 0.054435 | 6.9602 | 4312.5565 | 0.7034 | 0.00086196 | 1.6898e-07 | 1.0742e-06 | 1 | 1.6898e-07 | 1.4565e-10 | 1.0245e-10 | 1.0497e-20 | 1.0245e-10
[6 empty] | 0.0682 | 0.036275 | 18424713971.1742 | 18.4247 | 0.054275 | 6.9806 | 4359.8059 | -0.60885 | 0.00076536 | 1.4054e-07 | 9.9623e-07 | 1 | 1.4054e-07 | 1.0757e-10 | -6.5492e-11 | 4.2892e-21 | 6.5492e-11
[6 empty] | 0.0684 | 0.036381 | 18478745390.4445 | 18.4787 | 0.054116 | 7.0011 | 4407.4511 | 0.1329 | 0.00067845 | 1.167e-07 | 9.2372e-07 | 1 | 1.167e-07 | 7.9176e-11 | 1.0523e-11 | 1.1073e-22 | 1.0523e-11
[6 empty] | 0.0686 | 0.036487 | 18532776809.7148 | 18.5328 | 0.053958 | 7.0216 | 4455.4944 | 0.70756 | 0.00060039 | 9.6745e-08 | 8.5629e-07 | 1 | 9.6745e-08 | 5.8085e-11 | 4.1098e-11 | 1.6891e-21 | 4.1098e-11
[6 empty] | 0.0688 | 0.036594 | 18586808228.9851 | 18.5868 | 0.053802 | 7.042 | 4503.9379 | -0.86035 | 0.00053039 | 8.007e-08 | 7.9362e-07 | 1 | 8.007e-08 | 4.2469e-11 | -3.6538e-11 | 1.335e-21 | 3.6538e-11
[6 empty] | 0.069 | 0.0367 | 18640839648.2554 | 18.6408 | 0.053646 | 7.0625 | 4552.7839 | -0.63362 | 0.00046774 | 6.616e-08 | 7.3536e-07 | 1 | 6.616e-08 | 3.0946e-11 | -1.9608e-11 | 3.8446e-22 | 1.9608e-11
[6 empty] | 0.0692 | 0.036807 | 18694871067.5257 | 18.6949 | 0.053491 | 7.083 | 4602.0348 | 0.32249 | 0.00041177 | 5.4575e-08 | 6.8124e-07 | 1 | 5.4575e-08 | 2.2472e-11 | 7.2471e-12 | 5.2521e-23 | 7.2471e-12
[6 empty] | 0.0694 | 0.036913 | 18748902486.7961 | 18.7489 | 0.053336 | 7.1035 | 4651.6928 | 0.80509 | 0.00036185 | 4.4943e-08 | 6.3095e-07 | 1 | 4.4943e-08 | 1.6262e-11 | 1.3093e-11 | 1.7142e-22 | 1.3093e-11
[6 empty] | 0.0696 | 0.037019 | 18802933906.0664 | 18.8029 | 0.053183 | 7.1239 | 4701.7603 | 0.90606 | 0.00031741 | 3.6949e-08 | 5.8425e-07 | 1 | 3.6949e-08 | 1.1728e-11 | 1.0626e-11 | 1.1291e-22 | 1.0626e-11
[6 empty] | 0.0698 | 0.037126 | 18856965325.3367 | 18.857 | 0.053031 | 7.1444 | 4752.2394 | 0.79572 | 0.00027792 | 3.0325e-08 | 5.4089e-07 | 1 | 3.0325e-08 | 8.4279e-12 | 6.7062e-12 | 4.4974e-23 | 6.7062e-12
[6 empty] | 0.07 | 0.037232 | 18910996744.607 | 18.911 | 0.052879 | 7.1649 | 4803.1326 | 0.28833 | 0.0002429 | 2.4846e-08 | 5.0063e-07 | 1 | 2.4846e-08 | 6.0351e-12 | 1.7401e-12 | 3.028e-24 | 1.7401e-12
[6 empty] | 0.0702 | 0.037339 | 18965028163.8773 | 18.965 | 0.052729 | 7.1853 | 4854.4421 | -0.68278 | 0.00021189 | 2.0322e-08 | 4.6327e-07 | 1 | 2.0322e-08 | 4.3062e-12 | -2.9402e-12 | 8.6447e-24 | 2.9402e-12
[6 empty] | 0.0704 | 0.037445 | 19019059583.1476 | 19.0191 | 0.052579 | 7.2058 | 4906.1703 | -0.80002 | 0.0001845 | 1.6594e-08 | 4.286e-07 | 1 | 1.6594e-08 | 3.0615e-12 | -2.4493e-12 | 5.9991e-24 | 2.4493e-12
[6 empty] | 0.0706 | 0.037551 | 19073091002.4179 | 19.0731 | 0.05243 | 7.2263 | 4958.3195 | 0.8171 | 0.00016034 | 1.3526e-08 | 3.9644e-07 | 1 | 1.3526e-08 | 2.1687e-12 | 1.7721e-12 | 3.1402e-24 | 1.7721e-12
[6 empty] | 0.0708 | 0.037658 | 19127122421.6882 | 19.1271 | 0.052282 | 7.2467 | 5010.892 | -0.12148 | 0.00013908 | 1.1006e-08 | 3.6661e-07 | 1 | 1.1006e-08 | 1.5307e-12 | -1.8595e-13 | 3.4577e-26 | 1.8595e-13
[6 empty] | 0.071 | 0.037764 | 19181153840.9585 | 19.1812 | 0.052135 | 7.2672 | 5063.8902 | -0.2832 | 0.00012041 | 8.9393e-09 | 3.3895e-07 | 1 | 8.9393e-09 | 1.0763e-12 | -3.0482e-13 | 9.2917e-26 | 3.0482e-13
[6 empty] | 0.0712 | 0.03787 | 19235185260.2288 | 19.2352 | 0.051988 | 7.2877 | 5117.3163 | 0.26484 | 0.00010403 | 7.2482e-09 | 3.1331e-07 | 1 | 7.2482e-09 | 7.5404e-13 | 1.997e-13 | 3.9881e-26 | 1.997e-13
[6 empty] | 0.0714 | 0.037977 | 19289216679.4991 | 19.2892 | 0.051842 | 7.3082 | 5171.1729 | 0.18042 | 8.9707e-05 | 5.8666e-09 | 2.8954e-07 | 1 | 5.8666e-09 | 5.2627e-13 | 9.4951e-14 | 9.0157e-27 | 9.4951e-14
[6 empty] | 0.0716 | 0.038083 | 19343248098.7694 | 19.3432 | 0.051698 | 7.3286 | 5225.4621 | -0.87411 | 7.7199e-05 | 4.7399e-09 | 2.6752e-07 | 1 | 4.7399e-09 | 3.6591e-13 | -3.1985e-13 | 1.023e-25 | 3.1985e-13
[6 empty] | 0.0718 | 0.03819 | 19397279518.0397 | 19.3973 | 0.051554 | 7.3491 | 5280.1865 | 0.68945 | 6.63e-05 | 3.8227e-09 | 2.4711e-07 | 1 | 3.8227e-09 | 2.5344e-13 | 1.7474e-13 | 3.0533e-26 | 1.7474e-13
[6 empty] | 0.072 | 0.038296 | 19451310937.31 | 19.4513 | 0.05141 | 7.3696 | 5335.3483 | 0.83821 | 5.6824e-05 | 3.0774e-09 | 2.2822e-07 | 1 | 3.0774e-09 | 1.7487e-13 | 1.4658e-13 | 2.1485e-26 | 1.4658e-13
[6 empty] | 0.0722 | 0.038402 | 19505342356.5804 | 19.5053 | 0.051268 | 7.39 | 5390.9499 | 0.046966 | 4.8601e-05 | 2.4729e-09 | 2.1072e-07 | 1 | 2.4729e-09 | 1.2019e-13 | 5.6447e-15 | 3.1863e-29 | 5.6447e-15
[6 empty] | 0.0724 | 0.038509 | 19559373775.8507 | 19.5594 | 0.051126 | 7.4105 | 5446.9938 | -0.44198 | 4.148e-05 | 1.9836e-09 | 1.9452e-07 | 1 | 1.9836e-09 | 8.228e-14 | -3.6366e-14 | 1.3225e-27 | 3.6366e-14
[6 empty] | 0.0726 | 0.038615 | 19613405195.121 | 19.6134 | 0.050986 | 7.431 | 5503.4823 | -0.49512 | 3.5328e-05 | 1.5881e-09 | 1.7952e-07 | 1 | 1.5881e-09 | 5.6107e-14 | -2.7779e-14 | 7.7169e-28 | 2.7779e-14
[6 empty] | 0.0728 | 0.038721 | 19667436614.3913 | 19.6674 | 0.050845 | 7.4515 | 5560.4178 | -0.13074 | 3.0024e-05 | 1.2692e-09 | 1.6565e-07 | 1 | 1.2692e-09 | 3.8107e-14 | -4.9819e-15 | 2.482e-29 | 4.9819e-15
[6 empty] | 0.073 | 0.038828 | 19721468033.6616 | 19.7215 | 0.050706 | 7.4719 | 5617.8028 | 0.64817 | 2.5461e-05 | 1.0124e-09 | 1.5281e-07 | 1 | 1.0124e-09 | 2.5778e-14 | 1.6708e-14 | 2.7917e-28 | 1.6708e-14
[6 empty] | 0.0732 | 0.038934 | 19775499452.9319 | 19.7755 | 0.050568 | 7.4924 | 5675.6396 | 0.91206 | 2.1545e-05 | 8.0609e-10 | 1.4094e-07 | 1 | 8.0609e-10 | 1.7367e-14 | 1.584e-14 | 2.509e-28 | 1.584e-14
[6 empty] | 0.0734 | 0.039041 | 19829530872.2022 | 19.8295 | 0.05043 | 7.5129 | 5733.9306 | -0.55977 | 1.819e-05 | 6.4061e-10 | 1.2996e-07 | 1 | 6.4061e-10 | 1.1653e-14 | -6.523e-15 | 4.2549e-29 | 6.523e-15
[6 empty] | 0.0736 | 0.039147 | 19883562291.4725 | 19.8836 | 0.050293 | 7.5333 | 5792.6784 | -0.3414 | 1.5324e-05 | 5.0814e-10 | 1.1981e-07 | 1 | 5.0814e-10 | 7.7867e-15 | -2.6583e-15 | 7.0667e-30 | 2.6583e-15
[6 empty] | 0.0738 | 0.039253 | 19937593710.7428 | 19.9376 | 0.050157 | 7.5538 | 5851.8853 | 0.73914 | 1.288e-05 | 4.023e-10 | 1.1043e-07 | 1 | 4.023e-10 | 5.1816e-15 | 3.8299e-15 | 1.4668e-29 | 3.8299e-15
[6 empty] | 0.074 | 0.03936 | 19991625130.0131 | 19.9916 | 0.050021 | 7.5743 | 5911.5537 | -0.75366 | 1.0801e-05 | 3.179e-10 | 1.0176e-07 | 1 | 3.179e-10 | 3.4336e-15 | -2.5878e-15 | 6.6967e-30 | 2.5878e-15
[6 empty] | 0.0742 | 0.039466 | 20045656549.2834 | 20.0457 | 0.049886 | 7.5948 | 5971.6861 | 0.39992 | 9.0369e-06 | 2.5072e-10 | 9.3747e-08 | 1 | 2.5072e-10 | 2.2657e-15 | 9.0613e-16 | 8.2107e-31 | 9.0613e-16
[6 empty] | 0.0744 | 0.039572 | 20099687968.5537 | 20.0997 | 0.049752 | 7.6152 | 6032.285 | 0.47697 | 7.5433e-06 | 1.9736e-10 | 8.6348e-08 | 1 | 1.9736e-10 | 1.4887e-15 | 7.1009e-16 | 5.0423e-31 | 7.1009e-16
[6 empty] | 0.0746 | 0.039679 | 20153719387.824 | 20.1537 | 0.049619 | 7.6357 | 6093.3529 | -0.95415 | 6.2817e-06 | 1.5506e-10 | 7.9516e-08 | 1 | 1.5506e-10 | 9.74e-16 | -9.2935e-16 | 8.6369e-31 | 9.2935e-16
[6 empty] | 0.0748 | 0.039785 | 20207750807.0943 | 20.2078 | 0.049486 | 7.6562 | 6154.8921 | -0.5498 | 5.2186e-06 | 1.2158e-10 | 7.3208e-08 | 1 | 1.2158e-10 | 6.3449e-16 | -3.4884e-16 | 1.2169e-31 | 3.4884e-16
[6 empty] | 0.075 | 0.039892 | 20261782226.3647 | 20.2618 | 0.049354 | 7.6766 | 6216.9051 | 0.23445 | 4.3251e-06 | 9.5147e-11 | 6.7386e-08 | 1 | 9.5147e-11 | 4.1152e-16 | 9.6481e-17 | 9.3086e-33 | 9.6481e-17
[6 empty] | 0.0752 | 0.039998 | 20315813645.635 | 20.3158 | 0.049223 | 7.6971 | 6279.3946 | 0.54727 | 3.5758e-06 | 7.4314e-11 | 6.2013e-08 | 1 | 7.4314e-11 | 2.6573e-16 | 1.4543e-16 | 2.1149e-32 | 1.4543e-16
[6 empty] | 0.0754 | 0.040104 | 20369845064.9053 | 20.3698 | 0.049092 | 7.7176 | 6342.3628 | 0.4284 | 2.9491e-06 | 5.7928e-11 | 5.7055e-08 | 1 | 5.7928e-11 | 1.7083e-16 | 7.3184e-17 | 5.356e-33 | 7.3184e-17
[6 empty] | 0.0756 | 0.040211 | 20423876484.1756 | 20.4239 | 0.048962 | 7.7381 | 6405.8123 | -0.17409 | 2.4262e-06 | 4.5065e-11 | 5.2483e-08 | 1 | 4.5065e-11 | 1.0933e-16 | -1.9034e-17 | 3.6228e-34 | 1.9034e-17
[6 empty] | 0.0758 | 0.040317 | 20477907903.4459 | 20.4779 | 0.048833 | 7.7585 | 6469.7457 | -0.957 | 1.991e-06 | 3.4988e-11 | 4.8266e-08 | 1 | 3.4988e-11 | 6.966e-17 | -6.6665e-17 | 4.4442e-33 | 6.6665e-17
[6 empty] | 0.076 | 0.040423 | 20531939322.7162 | 20.5319 | 0.048705 | 7.779 | 6534.1654 | -0.27373 | 1.6297e-06 | 2.711e-11 | 4.4378e-08 | 1 | 2.711e-11 | 4.4181e-17 | -1.2094e-17 | 1.4625e-34 | 1.2094e-17
[6 empty] | 0.0762 | 0.04053 | 20585970741.9865 | 20.586 | 0.048577 | 7.7995 | 6599.0739 | 0.97398 | 1.3306e-06 | 2.0963e-11 | 4.0794e-08 | 1 | 2.0963e-11 | 2.7893e-17 | 2.7168e-17 | 7.3808e-34 | 2.7168e-17
[6 empty] | 0.0764 | 0.040636 | 20640002161.2568 | 20.64 | 0.04845 | 7.8199 | 6664.4738 | -0.94109 | 1.0836e-06 | 1.6177e-11 | 3.7492e-08 | 1 | 1.6177e-11 | 1.7529e-17 | -1.6496e-17 | 2.7213e-34 | 1.6496e-17
[6 empty] | 0.0766 | 0.040743 | 20694033580.5271 | 20.694 | 0.048323 | 7.8404 | 6730.3676 | 0.91119 | 8.801e-07 | 1.2458e-11 | 3.4449e-08 | 1 | 1.2458e-11 | 1.0965e-17 | 9.9908e-18 | 9.9816e-35 | 9.9908e-18
[6 empty] | 0.0768 | 0.040849 | 20748064999.7974 | 20.7481 | 0.048197 | 7.8609 | 6796.7577 | -0.99997 | 7.1295e-07 | 9.5748e-12 | 3.1646e-08 | 1 | 9.5748e-12 | 6.8264e-18 | -6.8262e-18 | 4.6596e-35 | 6.8262e-18
[6 empty] | 0.077 | 0.040955 | 20802096419.0677 | 20.8021 | 0.048072 | 7.8814 | 6863.6469 | 0.61552 | 5.7601e-07 | 7.3435e-12 | 2.9065e-08 | 1 | 7.3435e-12 | 4.2299e-18 | 2.6036e-18 | 6.7786e-36 | 2.6036e-18
[6 empty] | 0.0772 | 0.041062 | 20856127838.338 | 20.8561 | 0.047948 | 7.9018 | 6931.0375 | 0.68466 | 4.6412e-07 | 5.6204e-12 | 2.6688e-08 | 1 | 5.6204e-12 | 2.6086e-18 | 1.786e-18 | 3.1897e-36 | 1.786e-18
[6 empty] | 0.0774 | 0.041168 | 20910159257.6083 | 20.9102 | 0.047824 | 7.9223 | 6998.9321 | -0.44953 | 3.7295e-07 | 4.2927e-12 | 2.4501e-08 | 1 | 4.2927e-12 | 1.601e-18 | -7.1969e-19 | 5.1795e-37 | 7.1969e-19
[6 empty] | 0.0776 | 0.041274 | 20964190676.8787 | 20.9642 | 0.0477 | 7.9428 | 7067.3334 | -0.92461 | 2.9886e-07 | 3.2718e-12 | 2.2487e-08 | 1 | 3.2718e-12 | 9.7782e-19 | -9.041e-19 | 8.174e-37 | 9.041e-19
[6 empty] | 0.0778 | 0.041381 | 21018222096.149 | 21.0182 | 0.047578 | 7.9632 | 7136.2439 | -0.98271 | 2.3883e-07 | 2.4884e-12 | 2.0635e-08 | 1 | 2.4884e-12 | 5.9431e-19 | -5.8404e-19 | 3.411e-37 | 5.8404e-19
[6 empty] | 0.078 | 0.041487 | 21072253515.4193 | 21.0723 | 0.047456 | 7.9837 | 7205.666 | -0.88075 | 1.9032e-07 | 1.8886e-12 | 1.8931e-08 | 1 | 1.8886e-12 | 3.5945e-19 | -3.1658e-19 | 1.0022e-37 | 3.1658e-19
[6 empty] | 0.0782 | 0.041594 | 21126284934.6896 | 21.1263 | 0.047334 | 8.0042 | 7275.6026 | -0.25318 | 1.5124e-07 | 1.4303e-12 | 1.7363e-08 | 1 | 1.4303e-12 | 2.1632e-19 | -5.4767e-20 | 2.9994e-39 | 5.4767e-20
[6 empty] | 0.0784 | 0.0417 | 21180316353.9599 | 21.1803 | 0.047214 | 8.0246 | 7346.056 | 0.8831 | 1.1984e-07 | 1.0809e-12 | 1.5922e-08 | 1 | 1.0809e-12 | 1.2953e-19 | 1.1439e-19 | 1.3085e-38 | 1.1439e-19
[6 empty] | 0.0786 | 0.041806 | 21234347773.2302 | 21.2343 | 0.047094 | 8.0451 | 7417.0289 | 0.19992 | 9.4681e-08 | 8.1512e-13 | 1.4598e-08 | 1 | 8.1512e-13 | 7.7177e-20 | 1.5429e-20 | 2.3807e-40 | 1.5429e-20
[6 empty] | 0.0788 | 0.041913 | 21288379192.5005 | 21.2884 | 0.046974 | 8.0656 | 7488.5239 | -0.82084 | 7.4588e-08 | 6.1335e-13 | 1.338e-08 | 1 | 6.1335e-13 | 4.5749e-20 | -3.7552e-20 | 1.4102e-39 | 3.7552e-20
[6 empty] | 0.079 | 0.042019 | 21342410611.7708 | 21.3424 | 0.046855 | 8.0861 | 7560.5437 | 0.93196 | 5.8586e-08 | 4.6053e-13 | 1.2262e-08 | 1 | 4.6053e-13 | 2.698e-20 | 2.5145e-20 | 6.3225e-40 | 2.5145e-20
[6 empty] | 0.0792 | 0.042125 | 21396442031.0411 | 21.3964 | 0.046737 | 8.1065 | 7633.0907 | -0.78912 | 4.588e-08 | 3.4503e-13 | 1.1234e-08 | 1 | 3.4503e-13 | 1.583e-20 | -1.2492e-20 | 1.5604e-40 | 1.2492e-20
[6 empty] | 0.0794 | 0.042232 | 21450473450.3114 | 21.4505 | 0.046619 | 8.127 | 7706.1677 | 0.088898 | 3.5822e-08 | 2.5793e-13 | 1.0291e-08 | 1 | 2.5793e-13 | 9.2394e-21 | 8.2136e-22 | 6.7463e-43 | 8.2136e-22
[6 empty] | 0.0796 | 0.042338 | 21504504869.5817 | 21.5045 | 0.046502 | 8.1475 | 7779.7772 | 0.95327 | 2.7884e-08 | 1.9239e-13 | 9.4241e-09 | 1 | 1.9239e-13 | 5.3645e-21 | 5.1139e-21 | 2.6152e-41 | 5.1139e-21
[6 empty] | 0.0798 | 0.042445 | 21558536288.852 | 21.5585 | 0.046385 | 8.1679 | 7853.922 | 0.010407 | 2.1639e-08 | 1.4319e-13 | 8.6287e-09 | 1 | 1.4319e-13 | 3.0983e-21 | 3.2244e-23 | 1.0396e-45 | 3.2244e-23
[6 empty] | 0.08 | 0.042551 | 21612567708.1223 | 21.6126 | 0.046269 | 8.1884 | 7928.6045 | -0.64823 | 1.674e-08 | 1.0633e-13 | 7.8986e-09 | 1 | 1.0633e-13 | 1.78e-21 | -1.1538e-21 | 1.3313e-42 | 1.1538e-21
[6 empty] | 0.0802 | 0.042657 | 21666599127.3926 | 21.6666 | 0.046154 | 8.2089 | 8003.8276 | -0.77102 | 1.291e-08 | 7.8782e-14 | 7.2287e-09 | 1 | 7.8782e-14 | 1.0171e-21 | -7.8419e-22 | 6.1496e-43 | 7.8419e-22
[6 empty] | 0.0804 | 0.042764 | 21720630546.663 | 21.7206 | 0.046039 | 8.2294 | 8079.5938 | -0.49035 | 9.925e-09 | 5.8241e-14 | 6.6141e-09 | 1 | 5.8241e-14 | 5.7804e-22 | -2.8344e-22 | 8.0341e-44 | 2.8344e-22
[6 empty] | 0.0806 | 0.04287 | 21774661965.9333 | 21.7747 | 0.045925 | 8.2498 | 8155.9057 | 0.39055 | 7.6058e-09 | 4.2958e-14 | 6.0505e-09 | 1 | 4.2958e-14 | 3.2673e-22 | 1.2761e-22 | 1.6283e-44 | 1.2761e-22
[6 empty] | 0.0808 | 0.042977 | 21828693385.2036 | 21.8287 | 0.045811 | 8.2703 | 8232.7661 | 0.95749 | 5.8098e-09 | 3.1614e-14 | 5.5337e-09 | 1 | 3.1614e-14 | 1.8367e-22 | 1.7586e-22 | 3.0928e-44 | 1.7586e-22
[6 empty] | 0.081 | 0.043083 | 21882724804.4739 | 21.8827 | 0.045698 | 8.2908 | 8310.1776 | -0.67067 | 4.4235e-09 | 2.3212e-14 | 5.0599e-09 | 1 | 2.3212e-14 | 1.0268e-22 | -6.8864e-23 | 4.7422e-45 | 6.8864e-23
[6 empty] | 0.0812 | 0.043189 | 21936756223.7442 | 21.9368 | 0.045586 | 8.3112 | 8388.1429 | 0.15992 | 3.3569e-09 | 1.7004e-14 | 4.6256e-09 | 1 | 1.7004e-14 | 5.7082e-23 | 9.1288e-24 | 8.3335e-47 | 9.1288e-24
[6 empty] | 0.0814 | 0.043296 | 21990787643.0145 | 21.9908 | 0.045474 | 8.3317 | 8466.6647 | -0.14211 | 2.5391e-09 | 1.2428e-14 | 4.2277e-09 | 1 | 1.2428e-14 | 3.1556e-23 | -4.4846e-24 | 2.0111e-47 | 4.4846e-24
[6 empty] | 0.0816 | 0.043402 | 22044819062.2848 | 22.0448 | 0.045362 | 8.3522 | 8545.7457 | 0.63171 | 1.9141e-09 | 9.0621e-15 | 3.8631e-09 | 1 | 9.0621e-15 | 1.7346e-23 | 1.0958e-23 | 1.2007e-46 | 1.0958e-23
[6 empty] | 0.0818 | 0.043508 | 22098850481.5551 | 22.0989 | 0.045251 | 8.3727 | 8625.3886 | -0.97676 | 1.4381e-09 | 6.5925e-15 | 3.5292e-09 | 1 | 6.5925e-15 | 9.4806e-24 | -9.2603e-24 | 8.5753e-47 | 9.2603e-24
[6 empty] | 0.082 | 0.043615 | 22152881900.8254 | 22.1529 | 0.045141 | 8.3931 | 8705.596 | -0.30773 | 1.0768e-09 | 4.7846e-15 | 3.2235e-09 | 1 | 4.7846e-15 | 5.152e-24 | -1.5854e-24 | 2.5135e-48 | 1.5854e-24
[6 empty] | 0.0822 | 0.043721 | 22206913320.0957 | 22.2069 | 0.045031 | 8.4136 | 8786.3707 | 0.55959 | 8.0346e-10 | 3.4644e-15 | 2.9436e-09 | 1 | 3.4644e-15 | 2.7835e-24 | 1.5576e-24 | 2.4262e-48 | 1.5576e-24
[6 empty] | 0.0824 | 0.043828 | 22260944739.366 | 22.2609 | 0.044922 | 8.4341 | 8867.7154 | 0.80197 | 5.9744e-10 | 2.5025e-15 | 2.6874e-09 | 1 | 2.5025e-15 | 1.4951e-24 | 1.199e-24 | 1.4377e-48 | 1.199e-24
[6 empty] | 0.0826 | 0.043934 | 22314976158.6363 | 22.315 | 0.044813 | 8.4545 | 8949.6329 | 0.64003 | 4.4269e-10 | 1.8034e-15 | 2.4529e-09 | 1 | 1.8034e-15 | 7.9836e-25 | 5.1098e-25 | 2.611e-49 | 5.1098e-25
[6 empty] | 0.0828 | 0.04404 | 22369007577.9066 | 22.369 | 0.044705 | 8.475 | 9032.1258 | -0.11669 | 3.2686e-10 | 1.2965e-15 | 2.2384e-09 | 1 | 1.2965e-15 | 4.2379e-25 | -4.9453e-26 | 2.4456e-51 | 4.9453e-26
[6 empty] | 0.083 | 0.044147 | 22423038997.1769 | 22.423 | 0.044597 | 8.4955 | 9115.1969 | -0.99794 | 2.4048e-10 | 9.2987e-16 | 2.0422e-09 | 1 | 9.2987e-16 | 2.2362e-25 | -2.2316e-25 | 4.9799e-50 | 2.2316e-25
[6 empty] | 0.0832 | 0.044253 | 22477070416.4473 | 22.4771 | 0.04449 | 8.516 | 9198.8489 | 0.32944 | 1.7629e-10 | 6.653e-16 | 1.8628e-09 | 1 | 6.653e-16 | 1.1728e-25 | 3.8638e-26 | 1.4929e-51 | 3.8638e-26
[6 empty] | 0.0834 | 0.044359 | 22531101835.7176 | 22.5311 | 0.044383 | 8.5364 | 9283.0847 | 0.24891 | 1.2876e-10 | 4.7485e-16 | 1.6988e-09 | 1 | 4.7485e-16 | 6.1143e-26 | 1.5219e-26 | 2.3161e-52 | 1.5219e-26
[6 empty] | 0.0836 | 0.044466 | 22585133254.9879 | 22.5851 | 0.044277 | 8.5569 | 9367.9069 | -0.24969 | 9.3703e-11 | 3.3809e-16 | 1.5489e-09 | 1 | 3.3809e-16 | 3.168e-26 | -7.9102e-27 | 6.2572e-53 | 7.9102e-27
[6 empty] | 0.0838 | 0.044572 | 22639164674.2582 | 22.6392 | 0.044171 | 8.5774 | 9453.3183 | -0.32976 | 6.7937e-11 | 2.4014e-16 | 1.4119e-09 | 1 | 2.4014e-16 | 1.6314e-26 | -5.3797e-27 | 2.8942e-53 | 5.3797e-27
[6 empty] | 0.084 | 0.044679 | 22693196093.5285 | 22.6932 | 0.044066 | 8.5978 | 9539.3217 | 0.99852 | 4.9071e-11 | 1.7014e-16 | 1.2867e-09 | 1 | 1.7014e-16 | 8.3491e-27 | 8.3367e-27 | 6.9501e-53 | 8.3367e-27
[6 empty] | 0.0842 | 0.044785 | 22747227512.7988 | 22.7472 | 0.043961 | 8.6183 | 9625.9198 | 0.14945 | 3.5311e-11 | 1.2025e-16 | 1.1724e-09 | 1 | 1.2025e-16 | 4.2462e-27 | 6.3457e-28 | 4.0268e-55 | 6.3457e-28
[6 empty] | 0.0844 | 0.044891 | 22801258932.0691 | 22.8013 | 0.043857 | 8.6388 | 9713.1155 | -0.58012 | 2.5312e-11 | 8.478e-17 | 1.068e-09 | 1 | 8.478e-17 | 2.1459e-27 | -1.2449e-27 | 1.5498e-54 | 1.2449e-27
[6 empty] | 0.0846 | 0.044998 | 22855290351.3394 | 22.8553 | 0.043754 | 8.6593 | 9800.9116 | -0.70854 | 1.8075e-11 | 5.9623e-17 | 9.7266e-10 | 1 | 5.9623e-17 | 1.0777e-27 | -7.6357e-28 | 5.8304e-55 | 7.6357e-28
[6 empty] | 0.0848 | 0.045104 | 22909321770.6097 | 22.9093 | 0.04365 | 8.6797 | 9889.3108 | -0.34557 | 1.2857e-11 | 4.1826e-17 | 8.8565e-10 | 1 | 4.1826e-17 | 5.3774e-28 | -1.8583e-28 | 3.4532e-56 | 1.8583e-28
[6 empty] | 0.085 | 0.04521 | 22963353189.88 | 22.9634 | 0.043548 | 8.7002 | 9978.3159 | 0.63474 | 9.1091e-12 | 2.9267e-17 | 8.0624e-10 | 1 | 2.9267e-17 | 2.666e-28 | 1.6922e-28 | 2.8635e-56 | 1.6922e-28
[6 empty] | 0.0852 | 0.045317 | 23017384609.1503 | 23.0174 | 0.043445 | 8.7207 | 10067.9297 | 0.72059 | 6.4284e-12 | 2.0428e-17 | 7.3379e-10 | 1 | 2.0428e-17 | 1.3132e-28 | 9.4626e-29 | 8.9541e-57 | 9.4626e-29
[6 empty] | 0.0854 | 0.045423 | 23071416028.4206 | 23.0714 | 0.043344 | 8.7411 | 10158.1552 | -0.99344 | 4.5185e-12 | 1.4222e-17 | 6.677e-10 | 1 | 1.4222e-17 | 6.4261e-29 | -6.384e-29 | 4.0755e-57 | 6.384e-29
[6 empty] | 0.0856 | 0.04553 | 23125447447.6909 | 23.1254 | 0.043242 | 8.7616 | 10248.995 | 0.9283 | 3.1632e-12 | 9.8761e-18 | 6.0743e-10 | 1 | 9.8761e-18 | 3.124e-29 | 2.9e-29 | 8.4101e-58 | 2.9e-29
[6 empty] | 0.0858 | 0.045636 | 23179478866.9612 | 23.1795 | 0.043142 | 8.7821 | 10340.452 | -0.99955 | 2.2054e-12 | 6.8407e-18 | 5.5248e-10 | 1 | 6.8407e-18 | 1.5087e-29 | -1.508e-29 | 2.274e-58 | 1.508e-29
[6 empty] | 0.086 | 0.045742 | 23233510286.2316 | 23.2335 | 0.043041 | 8.8025 | 10432.5291 | 0.58916 | 1.5314e-12 | 4.726e-18 | 5.0239e-10 | 1 | 4.726e-18 | 7.2371e-30 | 4.2638e-30 | 1.818e-59 | 4.2638e-30
[6 empty] | 0.0862 | 0.045849 | 23287541705.5019 | 23.2875 | 0.042941 | 8.823 | 10525.2291 | 0.82136 | 1.0589e-12 | 3.2566e-18 | 4.5674e-10 | 1 | 3.2566e-18 | 3.4484e-30 | 2.8324e-30 | 8.0223e-60 | 2.8324e-30
[6 empty] | 0.0864 | 0.045955 | 23341573124.7722 | 23.3416 | 0.042842 | 8.8435 | 10618.5549 | 0.041747 | 7.2915e-13 | 2.2382e-18 | 4.1514e-10 | 1 | 2.2382e-18 | 1.632e-30 | 6.813e-32 | 4.6417e-63 | 6.813e-32
[6 empty] | 0.0866 | 0.046061 | 23395604544.0425 | 23.3956 | 0.042743 | 8.864 | 10712.5092 | -0.24902 | 4.9996e-13 | 1.5343e-18 | 3.7725e-10 | 1 | 1.5343e-18 | 7.6709e-31 | -1.9102e-31 | 3.6489e-62 | 1.9102e-31
[6 empty] | 0.0868 | 0.046168 | 23449635963.3128 | 23.4496 | 0.042645 | 8.8844 | 10807.095 | 0.086268 | 3.4136e-13 | 1.049e-18 | 3.4275e-10 | 1 | 1.049e-18 | 3.5808e-31 | 3.0891e-32 | 9.5425e-64 | 3.0891e-32
[6 empty] | 0.087 | 0.046274 | 23503667382.5831 | 23.5037 | 0.042547 | 8.9049 | 10902.3152 | 0.87175 | 2.3207e-13 | 7.153e-19 | 3.1133e-10 | 1 | 7.153e-19 | 1.66e-31 | 1.4471e-31 | 2.0941e-62 | 1.4471e-31
[6 empty] | 0.0872 | 0.046381 | 23557698801.8534 | 23.5577 | 0.042449 | 8.9254 | 10998.1726 | 0.45577 | 1.5708e-13 | 4.8647e-19 | 2.8272e-10 | 1 | 4.8647e-19 | 7.6417e-32 | 3.4828e-32 | 1.213e-63 | 3.4828e-32
[6 empty] | 0.0874 | 0.046487 | 23611730221.1237 | 23.6117 | 0.042352 | 8.9458 | 11094.6701 | -0.97892 | 1.0586e-13 | 3.2997e-19 | 2.5669e-10 | 1 | 3.2997e-19 | 3.4932e-32 | -3.4195e-32 | 1.1693e-63 | 3.4195e-32
[6 empty] | 0.0876 | 0.046593 | 23665761640.394 | 23.6658 | 0.042255 | 8.9663 | 11191.8105 | 0.99907 | 7.103e-14 | 2.2321e-19 | 2.33e-10 | 1 | 2.2321e-19 | 1.5855e-32 | 1.584e-32 | 2.5091e-64 | 1.584e-32
[6 empty] | 0.0878 | 0.0467 | 23719793059.6643 | 23.7198 | 0.042159 | 8.9868 | 11289.5969 | -0.93808 | 4.7446e-14 | 1.5059e-19 | 2.1146e-10 | 1 | 1.5059e-19 | 7.1451e-33 | -6.7027e-33 | 4.4926e-65 | 6.7027e-33
[6 empty] | 0.088 | 0.046806 | 23773824478.9346 | 23.7738 | 0.042063 | 9.0073 | 11388.032 | 0.17043 | 3.155e-14 | 1.0133e-19 | 1.9186e-10 | 1 | 1.0133e-19 | 3.1969e-33 | 5.4485e-34 | 2.9686e-67 | 5.4485e-34
[6 empty] | 0.0882 | 0.046912 | 23827855898.2049 | 23.8279 | 0.041968 | 9.0277 | 11487.1189 | 0.99901 | 2.0885e-14 | 6.7991e-20 | 1.7404e-10 | 1 | 6.7991e-20 | 1.42e-33 | 1.4186e-33 | 2.0124e-66 | 1.4186e-33
[6 empty] | 0.0884 | 0.047019 | 23881887317.4752 | 23.8819 | 0.041873 | 9.0482 | 11586.8604 | 0.67185 | 1.3762e-14 | 4.5499e-20 | 1.5784e-10 | 1 | 4.5499e-20 | 6.2616e-34 | 4.2068e-34 | 1.7697e-67 | 4.2068e-34
[6 empty] | 0.0886 | 0.047125 | 23935918736.7455 | 23.9359 | 0.041778 | 9.0687 | 11687.2594 | 0.56855 | 9.0263e-15 | 3.0365e-20 | 1.4312e-10 | 1 | 3.0365e-20 | 2.7408e-34 | 1.5583e-34 | 2.4282e-68 | 1.5583e-34
[6 empty] | 0.0888 | 0.047232 | 23989950156.0159 | 23.99 | 0.041684 | 9.0891 | 11788.3188 | 0.90579 | 5.8927e-15 | 2.0209e-20 | 1.2974e-10 | 1 | 2.0209e-20 | 1.1909e-34 | 1.0787e-34 | 1.1635e-68 | 1.0787e-34
[6 empty] | 0.089 | 0.047338 | 24043981575.2862 | 24.044 | 0.04159 | 9.1096 | 11890.0417 | 0.72872 | 3.829e-15 | 1.3413e-20 | 1.1759e-10 | 1 | 1.3413e-20 | 5.1356e-35 | 3.7424e-35 | 1.4006e-69 | 3.7424e-35
[6 empty] | 0.0892 | 0.047444 | 24098012994.5565 | 24.098 | 0.041497 | 9.1301 | 11992.4309 | -0.86332 | 2.4762e-15 | 8.8773e-21 | 1.0655e-10 | 1 | 8.8773e-21 | 2.1982e-35 | -1.8977e-35 | 3.6014e-70 | 1.8977e-35
[6 empty] | 0.0894 | 0.047551 | 24152044413.8268 | 24.152 | 0.041404 | 9.1506 | 12095.4894 | 0.41484 | 1.5937e-15 | 5.8592e-21 | 9.6522e-11 | 1 | 5.8592e-21 | 9.3379e-36 | 3.8737e-36 | 1.5006e-71 | 3.8737e-36
[6 empty] | 0.0896 | 0.047657 | 24206075833.0971 | 24.2061 | 0.041312 | 9.171 | 12199.2202 | -0.46707 | 1.0208e-15 | 3.8564e-21 | 8.7422e-11 | 1 | 3.8564e-21 | 3.9365e-36 | -1.8387e-36 | 3.3806e-72 | 1.8387e-36
[6 empty] | 0.0898 | 0.047764 | 24260107252.3674 | 24.2601 | 0.04122 | 9.1915 | 12303.6262 | 0.93889 | 6.5064e-16 | 2.5311e-21 | 7.9162e-11 | 1 | 2.5311e-21 | 1.6468e-36 | 1.5462e-36 | 2.3907e-72 | 1.5462e-36
[6 empty] | 0.09 | 0.04787 | 24314138671.6377 | 24.3141 | 0.041128 | 9.212 | 12408.7103 | -0.48868 | 4.1268e-16 | 1.6565e-21 | 7.1667e-11 | 1 | 1.6565e-21 | 6.8362e-37 | -3.3407e-37 | 1.116e-73 | 3.3407e-37
[6 empty] | 0.0902 | 0.047976 | 24368170090.908 | 24.3682 | 0.041037 | 9.2324 | 12514.4757 | -0.99994 | 2.6045e-16 | 1.0811e-21 | 6.4867e-11 | 1 | 1.0811e-21 | 2.8158e-37 | -2.8156e-37 | 7.9276e-74 | 2.8156e-37
[6 empty] | 0.0904 | 0.048083 | 24422201510.1783 | 24.4222 | 0.040946 | 9.2529 | 12620.9251 | -0.93824 | 1.6356e-16 | 7.0355e-22 | 5.8699e-11 | 1 | 7.0355e-22 | 1.1507e-37 | -1.0797e-37 | 1.1657e-74 | 1.0797e-37
[6 empty] | 0.0906 | 0.048189 | 24476232929.4486 | 24.4762 | 0.040856 | 9.2734 | 12728.0617 | -0.99952 | 1.022e-16 | 4.5654e-22 | 5.3106e-11 | 1 | 4.5654e-22 | 4.6657e-38 | -4.6634e-38 | 2.1748e-75 | 4.6634e-38
[6 empty] | 0.0908 | 0.048295 | 24530264348.7189 | 24.5303 | 0.040766 | 9.2939 | 12835.8885 | -0.55557 | 6.3529e-17 | 2.9541e-22 | 4.8035e-11 | 1 | 2.9541e-22 | 1.8767e-38 | -1.0426e-38 | 1.0871e-76 | 1.0426e-38
[6 empty] | 0.091 | 0.048402 | 24584295767.9892 | 24.5843 | 0.040676 | 9.3143 | 12944.4083 | 0.89864 | 3.9291e-17 | 1.9059e-22 | 4.3439e-11 | 1 | 1.9059e-22 | 7.4884e-39 | 6.7294e-39 | 4.5285e-77 | 6.7294e-39
[6 empty] | 0.0912 | 0.048508 | 24638327187.2595 | 24.6383 | 0.040587 | 9.3348 | 13053.6243 | -0.36806 | 2.4174e-17 | 1.2261e-22 | 3.9274e-11 | 1 | 1.2261e-22 | 2.964e-39 | -1.0909e-39 | 1.1902e-78 | 1.0909e-39
[6 empty] | 0.0914 | 0.048615 | 24692358606.5299 | 24.6924 | 0.040498 | 9.3553 | 13163.5395 | 0.33003 | 1.4796e-17 | 7.8648e-23 | 3.55e-11 | 1 | 7.8648e-23 | 1.1637e-39 | 3.8406e-40 | 1.475e-79 | 3.8406e-40
[6 empty] | 0.0916 | 0.048721 | 24746390025.8002 | 24.7464 | 0.04041 | 9.3757 | 13274.1568 | -0.84035 | 9.0085e-18 | 5.0301e-23 | 3.2082e-11 | 1 | 5.0301e-23 | 4.5314e-40 | -3.808e-40 | 1.4501e-79 | 3.808e-40
[6 empty] | 0.0918 | 0.048827 | 24800421445.0705 | 24.8004 | 0.040322 | 9.3962 | 13385.4794 | 0.70105 | 5.4557e-18 | 3.2077e-23 | 2.8987e-11 | 1 | 3.2077e-23 | 1.75e-40 | 1.2269e-40 | 1.5052e-80 | 1.2269e-40
[6 empty] | 0.092 | 0.048934 | 24854452864.3408 | 24.8545 | 0.040234 | 9.4167 | 13497.5102 | 0.96307 | 3.2864e-18 | 2.0395e-23 | 2.6184e-11 | 1 | 2.0395e-23 | 6.7027e-41 | 6.4552e-41 | 4.1669e-81 | 6.4552e-41
[6 empty] | 0.0922 | 0.04904 | 24908484283.6111 | 24.9085 | 0.040147 | 9.4372 | 13610.2522 | 0.80927 | 1.969e-18 | 1.2929e-23 | 2.3647e-11 | 1 | 1.2929e-23 | 2.5458e-41 | 2.0602e-41 | 4.2444e-82 | 2.0602e-41
[6 empty] | 0.0924 | 0.049146 | 24962515702.8814 | 24.9625 | 0.04006 | 9.4576 | 13723.7086 | 0.96408 | 1.1733e-18 | 8.1721e-24 | 2.1352e-11 | 1 | 8.1721e-24 | 9.588e-42 | 9.2436e-42 | 8.5444e-83 | 9.2436e-42
[6 empty] | 0.0926 | 0.049253 | 25016547122.1517 | 25.0165 | 0.039974 | 9.4781 | 13837.8824 | 0.69122 | 6.9528e-19 | 5.1498e-24 | 1.9274e-11 | 1 | 5.1498e-24 | 3.5805e-42 | 2.4749e-42 | 6.1253e-84 | 2.4749e-42
[6 empty] | 0.0928 | 0.049359 | 25070578541.422 | 25.0706 | 0.039887 | 9.4986 | 13952.7766 | -0.85921 | 4.0975e-19 | 3.2355e-24 | 1.7395e-11 | 1 | 3.2355e-24 | 1.3257e-42 | -1.1391e-42 | 1.2975e-84 | 1.1391e-42
[6 empty] | 0.093 | 0.049466 | 25124609960.6923 | 25.1246 | 0.039802 | 9.519 | 14068.3944 | 0.40094 | 2.4013e-19 | 2.0267e-24 | 1.5696e-11 | 1 | 2.0267e-24 | 4.8666e-43 | 1.9512e-43 | 3.8074e-86 | 1.9512e-43
[6 empty] | 0.0932 | 0.049572 | 25178641379.9626 | 25.1786 | 0.039716 | 9.5395 | 14184.7387 | -0.49511 | 1.3993e-19 | 1.2657e-24 | 1.416e-11 | 1 | 1.2657e-24 | 1.7711e-43 | -8.7688e-44 | 7.6891e-87 | 8.7688e-44
[6 empty] | 0.0934 | 0.049678 | 25232672799.2329 | 25.2327 | 0.039631 | 9.56 | 14301.8128 | 0.97639 | 8.1082e-20 | 7.8799e-25 | 1.2771e-11 | 1 | 7.8799e-25 | 6.3892e-44 | 6.2384e-44 | 3.8917e-87 | 6.2384e-44
[6 empty] | 0.0936 | 0.049785 | 25286704218.5032 | 25.2867 | 0.039546 | 9.5804 | 14419.6196 | -0.21889 | 4.6713e-20 | 4.8911e-25 | 1.1516e-11 | 1 | 4.8911e-25 | 2.2848e-44 | -5.001e-45 | 2.501e-89 | 5.001e-45
[6 empty] | 0.0938 | 0.049891 | 25340735637.7735 | 25.3407 | 0.039462 | 9.6009 | 14538.1622 | -0.87164 | 2.6757e-20 | 3.0266e-25 | 1.0382e-11 | 1 | 3.0266e-25 | 8.0983e-45 | -7.0589e-45 | 4.9827e-89 | 7.0589e-45
[6 empty] | 0.094 | 0.049997 | 25394767057.0438 | 25.3948 | 0.039378 | 9.6214 | 14657.4438 | -0.91579 | 1.5237e-20 | 1.8671e-25 | 9.3573e-12 | 1 | 1.8671e-25 | 2.845e-45 | -2.6054e-45 | 6.7879e-90 | 2.6054e-45
[6 empty] | 0.0942 | 0.050104 | 25448798476.3142 | 25.4488 | 0.039295 | 9.6419 | 14777.4675 | -0.49193 | 8.6262e-21 | 1.1483e-25 | 8.432e-12 | 1 | 1.1483e-25 | 9.905e-46 | -4.8726e-46 | 2.3742e-91 | 4.8726e-46
[6 empty] | 0.0944 | 0.05021 | 25502829895.5845 | 25.5028 | 0.039211 | 9.6623 | 14898.2363 | 0.76692 | 4.8546e-21 | 7.0398e-26 | 7.5966e-12 | 1 | 7.0398e-26 | 3.4175e-46 | 2.621e-46 | 6.8695e-92 | 2.621e-46
[6 empty] | 0.0946 | 0.050317 | 25556861314.8548 | 25.5569 | 0.039128 | 9.6828 | 15019.7535 | 0.13054 | 2.7158e-21 | 4.3025e-26 | 6.8424e-12 | 1 | 4.3025e-26 | 1.1685e-46 | 1.5253e-47 | 2.3266e-94 | 1.5253e-47
[6 empty] | 0.0948 | 0.050423 | 25610892734.1251 | 25.6109 | 0.039046 | 9.7033 | 15142.0221 | -0.37503 | 1.5101e-21 | 2.6214e-26 | 6.1617e-12 | 1 | 2.6214e-26 | 3.9586e-47 | -1.4846e-47 | 2.204e-94 | 1.4846e-47
[6 empty] | 0.095 | 0.050529 | 25664924153.3954 | 25.6649 | 0.038964 | 9.7237 | 15265.0453 | -0.11636 | 8.3461e-22 | 1.5921e-26 | 5.5476e-12 | 1 | 1.5921e-26 | 1.3288e-47 | -1.5462e-48 | 2.3906e-96 | 1.5462e-48
[6 empty] | 0.0952 | 0.050636 | 25718955572.6657 | 25.719 | 0.038882 | 9.7442 | 15388.8261 | 0.98097 | 4.5846e-22 | 9.639e-27 | 4.9935e-12 | 1 | 9.639e-27 | 4.4191e-48 | 4.335e-48 | 1.8792e-95 | 4.335e-48
[6 empty] | 0.0954 | 0.050742 | 25772986991.936 | 25.773 | 0.0388 | 9.7647 | 15513.3679 | 0.25071 | 2.5028e-22 | 5.8173e-27 | 4.4938e-12 | 1 | 5.8173e-27 | 1.4559e-48 | 3.6503e-49 | 1.3324e-97 | 3.6503e-49
[6 empty] | 0.0956 | 0.050848 | 25827018411.2063 | 25.827 | 0.038719 | 9.7852 | 15638.6736 | -0.10435 | 1.3578e-22 | 3.4997e-27 | 4.0432e-12 | 1 | 3.4997e-27 | 4.7519e-49 | -4.9586e-50 | 2.4587e-99 | 4.9586e-50
[6 empty] | 0.0958 | 0.050955 | 25881049830.4766 | 25.881 | 0.038638 | 9.8056 | 15764.7466 | 0.29999 | 7.3203e-23 | 2.0987e-27 | 3.6369e-12 | 1 | 2.0987e-27 | 1.5363e-49 | 4.6086e-50 | 2.1239e-99 | 4.6086e-50
[6 empty] | 0.096 | 0.051061 | 25935081249.7469 | 25.9351 | 0.038558 | 9.8261 | 15891.5898 | 0.99626 | 3.9216e-23 | 1.2545e-27 | 3.2708e-12 | 1 | 1.2545e-27 | 4.9195e-50 | 4.9011e-50 | 2.4021e-99 | 4.9011e-50
[6 empty] | 0.0962 | 0.051168 | 25989112669.0172 | 25.9891 | 0.038478 | 9.8466 | 16019.2066 | -0.29143 | 2.0875e-23 | 7.4745e-28 | 2.9409e-12 | 1 | 7.4745e-28 | 1.5603e-50 | -4.5471e-51 | 2.0676e-101 | 4.5471e-51
[6 empty] | 0.0964 | 0.051274 | 26043144088.2875 | 26.0431 | 0.038398 | 9.867 | 16147.6001 | -0.11584 | 1.104e-23 | 4.4391e-28 | 2.6436e-12 | 1 | 4.4391e-28 | 4.9009e-51 | -5.6773e-52 | 3.2231e-103 | 5.6773e-52
[6 empty] | 0.0966 | 0.05138 | 26097175507.5578 | 26.0972 | 0.038318 | 9.8875 | 16276.7734 | -0.24931 | 5.8012e-24 | 2.6278e-28 | 2.3759e-12 | 1 | 2.6278e-28 | 1.5244e-51 | -3.8005e-52 | 1.4444e-103 | 3.8005e-52
[6 empty] | 0.0968 | 0.051487 | 26151206926.8281 | 26.1512 | 0.038239 | 9.908 | 16406.7299 | 0.98597 | 3.0284e-24 | 1.5505e-28 | 2.1348e-12 | 1 | 1.5505e-28 | 4.6956e-52 | 4.6297e-52 | 2.1434e-103 | 4.6297e-52
[6 empty] | 0.097 | 0.051593 | 26205238346.0985 | 26.2052 | 0.03816 | 9.9285 | 16537.4726 | 0.19763 | 1.5705e-24 | 9.1186e-29 | 1.9178e-12 | 1 | 9.1186e-29 | 1.4321e-52 | 2.8302e-53 | 8.0103e-106 | 2.8302e-53
[6 empty] | 0.0972 | 0.051699 | 26259269765.3688 | 26.2593 | 0.038082 | 9.9489 | 16669.0048 | -0.21407 | 8.0905e-25 | 5.3451e-29 | 1.7224e-12 | 1 | 5.3451e-29 | 4.3244e-53 | -9.2572e-54 | 8.5697e-107 | 9.2572e-54
[6 empty] | 0.0974 | 0.051806 | 26313301184.6391 | 26.3133 | 0.038004 | 9.9694 | 16801.3297 | 0.16157 | 4.14e-25 | 3.1228e-29 | 1.5466e-12 | 1 | 3.1228e-29 | 1.2928e-53 | 2.0889e-54 | 4.3634e-108 | 2.0889e-54
[6 empty] | 0.0976 | 0.051912 | 26367332603.9094 | 26.3673 | 0.037926 | 9.9899 | 16934.4506 | 0.97262 | 2.1042e-25 | 1.8184e-29 | 1.3885e-12 | 1 | 1.8184e-29 | 3.8262e-54 | 3.7214e-54 | 1.3849e-107 | 3.7214e-54
[6 empty] | 0.0978 | 0.052019 | 26421364023.1797 | 26.4214 | 0.037848 | 10.0103 | 17068.3706 | -0.16695 | 1.0622e-25 | 1.0553e-29 | 1.2462e-12 | 1 | 1.0553e-29 | 1.1209e-54 | -1.8715e-55 | 3.5024e-110 | 1.8715e-55
[6 empty] | 0.098 | 0.052125 | 26475395442.45 | 26.4754 | 0.037771 | 10.0308 | 17203.0929 | -0.19707 | 5.3251e-26 | 6.1042e-30 | 1.1183e-12 | 1 | 6.1042e-30 | 3.2506e-55 | -6.406e-56 | 4.1036e-111 | 6.406e-56
[6 empty] | 0.0982 | 0.052231 | 26529426861.7203 | 26.5294 | 0.037694 | 10.0513 | 17338.6209 | -0.23882 | 2.6513e-26 | 3.519e-30 | 1.0032e-12 | 1 | 3.519e-30 | 9.3298e-56 | -2.2281e-56 | 4.9645e-112 | 2.2281e-56
[6 empty] | 0.0984 | 0.052338 | 26583458280.9906 | 26.5835 | 0.037617 | 10.0718 | 17474.9578 | 0.9967 | 1.3108e-26 | 2.0218e-30 | 8.9984e-13 | 1 | 2.0218e-30 | 2.6502e-56 | 2.6415e-56 | 6.9773e-112 | 2.6415e-56
[6 empty] | 0.0986 | 0.052444 | 26637489700.2609 | 26.6375 | 0.037541 | 10.0922 | 17612.1068 | 0.3972 | 6.4354e-27 | 1.1577e-30 | 8.0693e-13 | 1 | 1.1577e-30 | 7.45e-57 | 2.9591e-57 | 8.7564e-114 | 2.9591e-57
[6 empty] | 0.0988 | 0.05255 | 26691521119.5312 | 26.6915 | 0.037465 | 10.1127 | 17750.0712 | 0.14229 | 3.137e-27 | 6.6062e-31 | 7.2345e-13 | 1 | 6.6062e-31 | 2.0724e-57 | 2.9488e-58 | 8.6955e-116 | 2.9488e-58
[6 empty] | 0.099 | 0.052657 | 26745552538.8015 | 26.7456 | 0.037389 | 10.1332 | 17888.8543 | 0.64096 | 1.5183e-27 | 3.7569e-31 | 6.4846e-13 | 1 | 3.7569e-31 | 5.7041e-58 | 3.6561e-58 | 1.3367e-115 | 3.6561e-58
[6 empty] | 0.0992 | 0.052763 | 26799583958.0718 | 26.7996 | 0.037314 | 10.1536 | 18028.4592 | 0.87766 | 7.2956e-28 | 2.1293e-31 | 5.8112e-13 | 1 | 2.1293e-31 | 1.5534e-58 | 1.3634e-58 | 1.8588e-116 | 1.3634e-58
[6 empty] | 0.0994 | 0.05287 | 26853615377.3421 | 26.8536 | 0.037239 | 10.1741 | 18168.8894 | -0.90405 | 3.4802e-28 | 1.2026e-31 | 5.2065e-13 | 1 | 1.2026e-31 | 4.1854e-59 | -3.7838e-59 | 1.4317e-117 | 3.7838e-59
[6 empty] | 0.0996 | 0.052976 | 26907646796.6124 | 26.9076 | 0.037164 | 10.1946 | 18310.1481 | 0.85009 | 1.6481e-28 | 6.7689e-32 | 4.6637e-13 | 1 | 6.7689e-32 | 1.1156e-59 | 9.4833e-60 | 8.9933e-119 | 9.4833e-60
[6 empty] | 0.0998 | 0.053082 | 26961678215.8828 | 26.9617 | 0.03709 | 10.215 | 18452.2386 | -0.98655 | 7.7469e-29 | 3.7967e-32 | 4.1766e-13 | 1 | 3.7967e-32 | 2.9413e-60 | -2.9017e-60 | 8.42e-120 | 2.9017e-60
[6 empty] | 0.1 | 0.053189 | 27015709635.1531 | 27.0157 | 0.037015 | 10.2355 | 18595.1641 | -0.14673 | 3.6145e-29 | 2.1222e-32 | 3.7396e-13 | 1 | 2.1222e-32 | 7.6706e-61 | -1.1255e-61 | 1.2668e-122 | 1.1255e-61
[6 empty] | 0.1002 | 0.053295 | 27069741054.4234 | 27.0697 | 0.036942 | 10.256 | 18738.9281 | 0.56633 | 1.6738e-29 | 1.182e-32 | 3.3475e-13 | 1 | 1.182e-32 | 1.9785e-61 | 1.1205e-61 | 1.2555e-122 | 1.1205e-61
[6 empty] | 0.1004 | 0.053402 | 27123772473.6937 | 27.1238 | 0.036868 | 10.2765 | 18883.5337 | 0.48786 | 7.6928e-30 | 6.5609e-33 | 2.9959e-13 | 1 | 6.5609e-33 | 5.0471e-62 | 2.4623e-62 | 6.0629e-124 | 2.4623e-62
[6 empty] | 0.1006 | 0.053508 | 27177803892.964 | 27.1778 | 0.036795 | 10.2969 | 19028.9844 | -0.41487 | 3.5087e-30 | 3.6287e-33 | 2.6806e-13 | 1 | 3.6287e-33 | 1.2732e-62 | -5.2822e-63 | 2.7902e-125 | 5.2822e-63
[6 empty] | 0.1008 | 0.053614 | 27231835312.2343 | 27.2318 | 0.036722 | 10.3174 | 19175.2835 | -0.80042 | 1.5881e-30 | 1.9999e-33 | 2.398e-13 | 1 | 1.9999e-33 | 3.176e-63 | -2.5422e-63 | 6.4627e-126 | 2.5422e-63
[6 empty] | 0.101 | 0.053721 | 27285866731.5046 | 27.2859 | 0.036649 | 10.3379 | 19322.4343 | 0.9904 | 7.1325e-31 | 1.0983e-33 | 2.1447e-13 | 1 | 1.0983e-33 | 7.8336e-64 | 7.7584e-64 | 6.0193e-127 | 7.7584e-64
[6 empty] | 0.1012 | 0.053827 | 27339898150.7749 | 27.3399 | 0.036577 | 10.3583 | 19470.44 | -0.88251 | 3.1784e-31 | 6.0101e-34 | 1.9177e-13 | 1 | 6.0101e-34 | 1.9103e-64 | -1.6858e-64 | 2.8421e-128 | 1.6858e-64
[6 empty] | 0.1014 | 0.053933 | 27393929570.0452 | 27.3939 | 0.036504 | 10.3788 | 19619.3042 | -0.12777 | 1.4053e-31 | 3.2771e-34 | 1.7143e-13 | 1 | 3.2771e-34 | 4.6053e-65 | -5.8843e-66 | 3.4625e-131 | 5.8843e-66
[6 empty] | 0.1016 | 0.05404 | 27447960989.3155 | 27.448 | 0.036433 | 10.3993 | 19769.0301 | 0.80899 | 6.1643e-32 | 1.7804e-34 | 1.5322e-13 | 1 | 1.7804e-34 | 1.0975e-65 | 8.8788e-66 | 7.8833e-131 | 8.8788e-66
[6 empty] | 0.1018 | 0.054146 | 27501992408.5858 | 27.502 | 0.036361 | 10.4198 | 19919.6211 | 0.91191 | 2.6824e-32 | 9.6382e-35 | 1.3692e-13 | 1 | 9.6382e-35 | 2.5854e-66 | 2.3576e-66 | 5.5584e-132 | 2.3576e-66
[6 empty] | 0.102 | 0.054253 | 27556023827.8561 | 27.556 | 0.03629 | 10.4402 | 20071.0805 | 0.46613 | 1.1579e-32 | 5.1986e-35 | 1.2232e-13 | 1 | 5.1986e-35 | 6.0196e-67 | 2.8059e-67 | 7.8732e-134 | 2.8059e-67
[6 empty] | 0.1022 | 0.054359 | 27610055247.1264 | 27.6101 | 0.036219 | 10.4607 | 20223.4118 | -0.86738 | 4.9579e-33 | 2.7939e-35 | 1.0925e-13 | 1 | 2.7939e-35 | 1.3852e-67 | -1.2015e-67 | 1.4435e-134 | 1.2015e-67
[6 empty] | 0.1024 | 0.054465 | 27664086666.3967 | 27.6641 | 0.036148 | 10.4812 | 20376.6182 | 0.313 | 2.1056e-33 | 1.496e-35 | 9.756e-14 | 1 | 1.496e-35 | 3.1499e-68 | 9.8593e-69 | 9.7206e-137 | 9.8593e-69
[6 empty] | 0.1026 | 0.054572 | 27718118085.6671 | 27.7181 | 0.036077 | 10.5016 | 20530.7033 | -0.44877 | 8.8689e-34 | 7.981e-36 | 8.71e-14 | 1 | 7.981e-36 | 7.0783e-69 | -3.1765e-69 | 1.009e-137 | 3.1765e-69
[6 empty] | 0.1028 | 0.054678 | 27772149504.9374 | 27.7721 | 0.036007 | 10.5221 | 20685.6704 | 0.99708 | 3.7047e-34 | 4.2422e-36 | 7.7744e-14 | 1 | 4.2422e-36 | 1.5716e-69 | 1.567e-69 | 2.4556e-138 | 1.567e-69
[6 empty] | 0.103 | 0.054784 | 27826180924.2077 | 27.8262 | 0.035937 | 10.5426 | 20841.5228 | 0.26407 | 1.5346e-34 | 2.2465e-36 | 6.9378e-14 | 1 | 2.2465e-36 | 3.4476e-70 | 9.1041e-71 | 8.2884e-141 | 9.1041e-71
[6 empty] | 0.1032 | 0.054891 | 27880212343.478 | 27.8802 | 0.035868 | 10.5631 | 20998.2641 | -0.071084 | 6.3037e-35 | 1.1853e-36 | 6.1898e-14 | 1 | 1.1853e-36 | 7.4718e-71 | -5.3112e-72 | 2.8209e-143 | 5.3112e-72
[6 empty] | 0.1034 | 0.054997 | 27934243762.7483 | 27.9342 | 0.035798 | 10.5835 | 21155.8976 | 0.46416 | 2.5674e-35 | 6.2304e-37 | 5.5213e-14 | 1 | 6.2304e-37 | 1.5996e-71 | 7.4247e-72 | 5.5127e-143 | 7.4247e-72
[6 empty] | 0.1036 | 0.055104 | 27988275182.0186 | 27.9883 | 0.035729 | 10.604 | 21314.4266 | 0.93542 | 1.0368e-35 | 3.2627e-37 | 4.9239e-14 | 1 | 3.2627e-37 | 3.3827e-72 | 3.1643e-72 | 1.0012e-143 | 3.1643e-72
[6 empty] | 0.1038 | 0.05521 | 28042306601.2889 | 28.0423 | 0.03566 | 10.6245 | 21473.8548 | -0.90824 | 4.1506e-36 | 1.7022e-37 | 4.3901e-14 | 1 | 1.7022e-37 | 7.0653e-73 | -6.417e-73 | 4.1178e-145 | 6.417e-73
[6 empty] | 0.104 | 0.055316 | 28096338020.5592 | 28.0963 | 0.035592 | 10.6449 | 21634.1854 | 0.9485 | 1.6473e-36 | 8.8474e-38 | 3.9134e-14 | 1 | 8.8474e-38 | 1.4574e-73 | 1.3824e-73 | 1.911e-146 | 1.3824e-73
[6 empty] | 0.1042 | 0.055423 | 28150369439.8295 | 28.1504 | 0.035524 | 10.6654 | 21795.422 | -0.76933 | 6.4808e-37 | 4.5811e-38 | 3.4876e-14 | 1 | 4.5811e-38 | 2.9689e-74 | -2.2841e-74 | 5.217e-148 | 2.2841e-74
[6 empty] | 0.1044 | 0.055529 | 28204400859.0998 | 28.2044 | 0.035455 | 10.6859 | 21957.5679 | -0.86588 | 2.5273e-37 | 2.363e-38 | 3.1075e-14 | 1 | 2.363e-38 | 5.972e-75 | -5.171e-75 | 2.6739e-149 | 5.171e-75
[6 empty] | 0.1046 | 0.055635 | 28258432278.3701 | 28.2584 | 0.035388 | 10.7064 | 22120.6267 | -0.67639 | 9.768e-38 | 1.2143e-38 | 2.7682e-14 | 1 | 1.2143e-38 | 1.1861e-75 | -8.0226e-76 | 6.4362e-151 | 8.0226e-76
[6 empty] | 0.1048 | 0.055742 | 28312463697.6404 | 28.3125 | 0.03532 | 10.7268 | 22284.6017 | -0.97682 | 3.7417e-38 | 6.2158e-39 | 2.4654e-14 | 1 | 6.2158e-39 | 2.3258e-76 | -2.2719e-76 | 5.1614e-152 | 2.2719e-76
[6 empty] | 0.105 | 0.055848 | 28366495116.9107 | 28.3665 | 0.035253 | 10.7473 | 22449.4965 | -0.25181 | 1.4204e-38 | 3.1696e-39 | 2.1952e-14 | 1 | 3.1696e-39 | 4.5022e-77 | -1.1337e-77 | 1.2853e-154 | 1.1337e-77
[6 empty] | 0.1052 | 0.055955 | 28420526536.1811 | 28.4205 | 0.035186 | 10.7678 | 22615.3145 | 0.80812 | 5.3434e-39 | 1.6101e-39 | 1.9542e-14 | 1 | 1.6101e-39 | 8.6031e-78 | 6.9524e-78 | 4.8335e-155 | 6.9524e-78
[6 empty] | 0.1054 | 0.056061 | 28474557955.4514 | 28.4746 | 0.035119 | 10.7882 | 22782.0592 | -0.64472 | 1.9917e-39 | 8.1469e-40 | 1.7393e-14 | 1 | 8.1469e-40 | 1.6226e-78 | -1.0461e-78 | 1.0944e-156 | 1.0461e-78
[6 empty] | 0.1056 | 0.056167 | 28528589374.7217 | 28.5286 | 0.035053 | 10.8087 | 22949.7341 | -0.45274 | 7.3559e-40 | 4.1064e-40 | 1.5477e-14 | 1 | 4.1064e-40 | 3.0206e-79 | -1.3675e-79 | 1.8702e-158 | 1.3675e-79
[6 empty] | 0.1058 | 0.056274 | 28582620793.992 | 28.5826 | 0.034986 | 10.8292 | 23118.3427 | 0.5376 | 2.6915e-40 | 2.0617e-40 | 1.3768e-14 | 1 | 2.0617e-40 | 5.5491e-80 | 2.9832e-80 | 8.8994e-160 | 2.9832e-80
[6 empty] | 0.106 | 0.05638 | 28636652213.2623 | 28.6367 | 0.03492 | 10.8497 | 23287.8885 | 0.61928 | 9.7564e-41 | 1.0311e-40 | 1.2246e-14 | 1 | 1.0311e-40 | 1.006e-80 | 6.2298e-81 | 3.881e-161 | 6.2298e-81
[6 empty] | 0.1062 | 0.056486 | 28690683632.5326 | 28.6907 | 0.034855 | 10.8701 | 23458.3749 | -0.17176 | 3.5033e-41 | 5.1363e-41 | 1.089e-14 | 1 | 5.1363e-41 | 1.7994e-81 | -3.0906e-82 | 9.5518e-164 | 3.0906e-82
[6 empty] | 0.1064 | 0.056593 | 28744715051.8029 | 28.7447 | 0.034789 | 10.8906 | 23629.8055 | -0.92625 | 1.2461e-41 | 2.5486e-41 | 9.6813e-15 | 1 | 2.5486e-41 | 3.1757e-82 | -2.9415e-82 | 8.6523e-164 | 2.9415e-82
[6 empty] | 0.1066 | 0.056699 | 28798746471.0732 | 28.7987 | 0.034724 | 10.9111 | 23802.1839 | 0.99974 | 4.3897e-42 | 1.2596e-41 | 8.6052e-15 | 1 | 1.2596e-41 | 5.5291e-83 | 5.5277e-83 | 3.0555e-165 | 5.5277e-83
[6 empty] | 0.1068 | 0.056806 | 28852777890.3435 | 28.8528 | 0.034659 | 10.9315 | 23975.5135 | -0.86821 | 1.5316e-42 | 6.2003e-42 | 7.6469e-15 | 1 | 6.2003e-42 | 9.4963e-84 | -8.2448e-84 | 6.7977e-167 | 8.2448e-84
[6 empty] | 0.107 | 0.056912 | 28906809309.6138 | 28.9068 | 0.034594 | 10.952 | 24149.7979 | -0.43062 | 5.292e-43 | 3.04e-42 | 6.7939e-15 | 1 | 3.04e-42 | 1.6088e-84 | -6.9277e-85 | 4.7993e-169 | 6.9277e-85
[6 empty] | 0.1072 | 0.057018 | 28960840728.8841 | 28.9608 | 0.034529 | 10.9725 | 24325.0405 | 0.23899 | 1.8107e-43 | 1.4845e-42 | 6.0347e-15 | 1 | 1.4845e-42 | 2.688e-85 | 6.4242e-86 | 4.127e-171 | 6.4242e-86
[6 empty] | 0.1074 | 0.057125 | 29014872148.1544 | 29.0149 | 0.034465 | 10.9929 | 24501.2451 | -0.034012 | 6.1345e-44 | 7.2205e-43 | 5.3591e-15 | 1 | 7.2205e-43 | 4.4294e-86 | -1.5065e-87 | 2.2697e-174 | 1.5065e-87
[6 empty] | 0.1076 | 0.057231 | 29068903567.4247 | 29.0689 | 0.034401 | 11.0134 | 24678.415 | -0.9565 | 2.0578e-44 | 3.4977e-43 | 4.7581e-15 | 1 | 3.4977e-43 | 7.1975e-87 | -6.8844e-87 | 4.7395e-173 | 6.8844e-87
[6 empty] | 0.1078 | 0.057337 | 29122934986.695 | 29.1229 | 0.034337 | 11.0339 | 24856.5539 | 0.33622 | 6.8338e-45 | 1.6875e-43 | 4.2236e-15 | 1 | 1.6875e-43 | 1.1532e-87 | 3.8773e-88 | 1.5033e-175 | 3.8773e-88
[6 empty] | 0.108 | 0.057444 | 29176966405.9654 | 29.177 | 0.034274 | 11.0544 | 25035.6654 | -0.37423 | 2.2467e-45 | 8.1084e-44 | 3.7483e-15 | 1 | 8.1084e-44 | 1.8217e-88 | -6.8173e-89 | 4.6476e-177 | 6.8173e-89
[6 empty] | 0.1082 | 0.05755 | 29230997825.2357 | 29.231 | 0.03421 | 11.0748 | 25215.7529 | 0.98551 | 7.3115e-46 | 3.8802e-44 | 3.3258e-15 | 1 | 3.8802e-44 | 2.837e-89 | 2.7959e-89 | 7.8171e-178 | 2.7959e-89
[6 empty] | 0.1084 | 0.057657 | 29285029244.506 | 29.285 | 0.034147 | 11.0953 | 25396.8201 | 0.25242 | 2.3552e-46 | 1.8492e-44 | 2.9502e-15 | 1 | 1.8492e-44 | 4.3553e-90 | 1.0994e-90 | 1.2086e-180 | 1.0994e-90
[6 empty] | 0.1086 | 0.057763 | 29339060663.7763 | 29.3391 | 0.034084 | 11.1158 | 25578.8706 | 0.093168 | 7.5087e-47 | 8.777e-45 | 2.6165e-15 | 1 | 8.777e-45 | 6.5904e-91 | 6.1402e-92 | 3.7702e-183 | 6.1402e-92
[6 empty] | 0.1088 | 0.057869 | 29393092083.0466 | 29.3931 | 0.034022 | 11.1362 | 25761.908 | 0.79456 | 2.3691e-47 | 4.1487e-45 | 2.32e-15 | 1 | 4.1487e-45 | 9.8288e-92 | 7.8096e-92 | 6.0989e-183 | 7.8096e-92
[6 empty] | 0.109 | 0.057976 | 29447123502.3169 | 29.4471 | 0.033959 | 11.1567 | 25945.9358 | 0.39666 | 7.3973e-48 | 1.9528e-45 | 2.0566e-15 | 1 | 1.9528e-45 | 1.4446e-92 | 5.73e-93 | 3.2833e-185 | 5.73e-93
[6 empty] | 0.1092 | 0.058082 | 29501154921.5872 | 29.5012 | 0.033897 | 11.1772 | 26130.9577 | -0.67426 | 2.2855e-48 | 9.1542e-46 | 1.8227e-15 | 1 | 9.1542e-46 | 2.0922e-93 | -1.4107e-93 | 1.99e-186 | 1.4107e-93
[6 empty] | 0.1094 | 0.058189 | 29555186340.8575 | 29.5552 | 0.033835 | 11.1977 | 26316.9772 | 0.074317 | 6.9866e-49 | 4.2733e-46 | 1.6151e-15 | 1 | 4.2733e-46 | 2.9856e-94 | 2.2188e-95 | 4.9229e-190 | 2.2188e-95
[6 empty] | 0.1096 | 0.058295 | 29609217760.1278 | 29.6092 | 0.033773 | 11.2181 | 26503.9981 | 0.99976 | 2.113e-49 | 1.9865e-46 | 1.4308e-15 | 1 | 1.9865e-46 | 4.1975e-95 | 4.1965e-95 | 1.7611e-189 | 4.1965e-95
[6 empty] | 0.1098 | 0.058401 | 29663249179.3981 | 29.6632 | 0.033712 | 11.2386 | 26692.0238 | 0.90127 | 6.3223e-50 | 9.1956e-47 | 1.2673e-15 | 1 | 9.1956e-47 | 5.8137e-96 | 5.2397e-96 | 2.7454e-191 | 5.2397e-96
[6 empty] | 0.11 | 0.058508 | 29717280598.6684 | 29.7173 | 0.03365 | 11.2591 | 26881.0581 | 0.9959 | 1.8712e-50 | 4.2388e-47 | 1.1222e-15 | 1 | 4.2388e-47 | 7.9317e-97 | 7.8991e-97 | 6.2396e-193 | 7.8991e-97
[6 empty] | 0.1102 | 0.058614 | 29771312017.9387 | 29.7713 | 0.033589 | 11.2795 | 27071.1046 | -0.070699 | 5.4781e-51 | 1.9456e-47 | 9.9348e-16 | 1 | 1.9456e-47 | 1.0658e-97 | -7.5354e-99 | 5.6783e-197 | 7.5354e-99
[6 empty] | 0.1104 | 0.05872 | 29825343437.209 | 29.8253 | 0.033529 | 11.3 | 27262.1669 | -0.483 | 1.5862e-51 | 8.8928e-48 | 8.7934e-16 | 1 | 8.8928e-48 | 1.4106e-98 | -6.813e-99 | 4.6417e-197 | 6.813e-99
[6 empty] | 0.1106 | 0.058827 | 29879374856.4793 | 29.8794 | 0.033468 | 11.3205 | 27454.2487 | 0.059395 | 4.5421e-52 | 4.0473e-48 | 7.7814e-16 | 1 | 4.0473e-48 | 1.8383e-99 | 1.0919e-100 | 1.1922e-200 | 1.0919e-100
[6 empty] | 0.1108 | 0.058933 | 29933406275.7497 | 29.9334 | 0.033407 | 11.341 | 27647.3537 | 0.98684 | 1.2862e-52 | 1.8341e-48 | 6.8844e-16 | 1 | 1.8341e-48 | 2.3591e-100 | 2.328e-100 | 5.4198e-200 | 2.328e-100
[6 empty] | 0.111 | 0.05904 | 29987437695.02 | 29.9874 | 0.033347 | 11.3614 | 27841.4854 | 0.6899 | 3.6016e-53 | 8.2761e-49 | 6.0894e-16 | 1 | 8.2761e-49 | 2.9807e-101 | 2.0564e-101 | 4.2287e-202 | 2.0564e-101
[6 empty] | 0.1112 | 0.059146 | 30041469114.2903 | 30.0415 | 0.033287 | 11.3819 | 28036.6476 | 0.91064 | 9.9712e-54 | 3.7183e-49 | 5.385e-16 | 1 | 3.7183e-49 | 3.7076e-102 | 3.3763e-102 | 1.1399e-203 | 3.3763e-102
[6 empty] | 0.1114 | 0.059252 | 30095500533.5606 | 30.0955 | 0.033228 | 11.4024 | 28232.844 | 0.54731 | 2.7293e-54 | 1.6634e-49 | 4.7611e-16 | 1 | 1.6634e-49 | 4.5398e-103 | 2.4847e-103 | 6.1736e-206 | 2.4847e-103
[6 empty] | 0.1116 | 0.059359 | 30149531952.8309 | 30.1495 | 0.033168 | 11.4228 | 28430.0782 | -0.95369 | 7.385e-55 | 7.4088e-50 | 4.2085e-16 | 1 | 7.4088e-50 | 5.4715e-104 | -5.2181e-104 | 2.7228e-207 | 5.2181e-104
[6 empty] | 0.1118 | 0.059465 | 30203563372.1012 | 30.2036 | 0.033109 | 11.4433 | 28628.3539 | 0.78943 | 1.9753e-55 | 3.2856e-50 | 3.7192e-16 | 1 | 3.2856e-50 | 6.4901e-105 | 5.1235e-105 | 2.625e-209 | 5.1235e-105
[6 empty] | 0.112 | 0.059571 | 30257594791.3715 | 30.2576 | 0.03305 | 11.4638 | 28827.6749 | 0.47128 | 5.2223e-56 | 1.4507e-50 | 3.2861e-16 | 1 | 1.4507e-50 | 7.5759e-106 | 3.5704e-106 | 1.2748e-211 | 3.5704e-106
[6 empty] | 0.1122 | 0.059678 | 30311626210.6418 | 30.3116 | 0.032991 | 11.4843 | 29028.0448 | -0.19995 | 1.3645e-56 | 6.3773e-51 | 2.9028e-16 | 1 | 6.3773e-51 | 8.702e-107 | -1.7399e-107 | 3.0274e-214 | 1.7399e-107
[6 empty] | 0.1124 | 0.059784 | 30365657629.9121 | 30.3657 | 0.032932 | 11.5047 | 29229.4673 | 0.15862 | 3.5235e-57 | 2.7911e-51 | 2.5636e-16 | 1 | 2.7911e-51 | 9.8345e-108 | 1.5599e-108 | 2.4334e-216 | 1.5599e-108
[6 empty] | 0.1126 | 0.059891 | 30419689049.1824 | 30.4197 | 0.032873 | 11.5252 | 29431.9462 | 0.99999 | 8.9908e-58 | 1.2162e-51 | 2.2636e-16 | 1 | 1.2162e-51 | 1.0934e-108 | 1.0934e-108 | 1.1956e-216 | 1.0934e-108
[6 empty] | 0.1128 | 0.059997 | 30473720468.4527 | 30.4737 | 0.032815 | 11.5457 | 29635.4851 | -0.79054 | 2.2668e-58 | 5.2759e-52 | 1.9982e-16 | 1 | 5.2759e-52 | 1.1959e-109 | -9.4543e-110 | 8.9385e-219 | 9.4543e-110
[6 empty] | 0.113 | 0.060103 | 30527751887.723 | 30.5278 | 0.032757 | 11.5661 | 29840.0879 | 0.96643 | 5.6466e-59 | 2.2785e-52 | 1.7635e-16 | 1 | 2.2785e-52 | 1.2866e-110 | 1.2434e-110 | 1.546e-220 | 1.2434e-110
[6 empty] | 0.1132 | 0.06021 | 30581783306.9933 | 30.5818 | 0.032699 | 11.5866 | 30045.7582 | -0.35587 | 1.3896e-59 | 9.7965e-53 | 1.5561e-16 | 1 | 9.7965e-53 | 1.3613e-111 | -4.8444e-112 | 2.3468e-223 | 4.8444e-112
[6 empty] | 0.1134 | 0.060316 | 30635814726.2636 | 30.6358 | 0.032642 | 11.6071 | 30252.4999 | -0.82338 | 3.378e-60 | 4.1932e-53 | 1.3728e-16 | 1 | 4.1932e-53 | 1.4165e-112 | -1.1663e-112 | 1.3602e-224 | 1.1663e-112
[6 empty] | 0.1136 | 0.060422 | 30689846145.534 | 30.6898 | 0.032584 | 11.6276 | 30460.3165 | -0.47572 | 8.1112e-61 | 1.7867e-53 | 1.2108e-16 | 1 | 1.7867e-53 | 1.4493e-113 | -6.8944e-114 | 4.7533e-227 | 6.8944e-114
[6 empty] | 0.1138 | 0.060529 | 30743877564.8043 | 30.7439 | 0.032527 | 11.648 | 30669.212 | 0.86969 | 1.9236e-61 | 7.579e-54 | 1.0676e-16 | 1 | 7.579e-54 | 1.4579e-114 | 1.2679e-114 | 1.6076e-228 | 1.2679e-114
[6 empty] | 0.114 | 0.060635 | 30797908984.0746 | 30.7979 | 0.03247 | 11.6685 | 30879.19 | -0.51918 | 4.5053e-62 | 3.2003e-54 | 9.4121e-17 | 1 | 3.2003e-54 | 1.4419e-115 | -7.4858e-116 | 5.6037e-231 | 7.4858e-116
[6 empty] | 0.1142 | 0.060742 | 30851940403.3449 | 30.8519 | 0.032413 | 11.689 | 31090.2544 | 0.90164 | 1.042e-62 | 1.3453e-54 | 8.2958e-17 | 1 | 1.3453e-54 | 1.4017e-116 | 1.2639e-116 | 1.5974e-232 | 1.2639e-116
[6 empty] | 0.1144 | 0.060848 | 30905971822.6152 | 30.906 | 0.032356 | 11.7094 | 31302.4089 | -0.34316 | 2.3796e-63 | 5.629e-55 | 7.3103e-17 | 1 | 5.629e-55 | 1.3395e-117 | -4.5965e-118 | 2.1127e-235 | 4.5965e-118
[6 empty] | 0.1146 | 0.060954 | 30960003241.8855 | 30.96 | 0.0323 | 11.7299 | 31515.6572 | -0.66702 | 5.3652e-64 | 2.3446e-55 | 6.4404e-17 | 1 | 2.3446e-55 | 1.2579e-118 | -8.3907e-119 | 7.0403e-237 | 8.3907e-119
[6 empty] | 0.1148 | 0.061061 | 31014034661.1558 | 31.014 | 0.032243 | 11.7504 | 31730.0033 | -0.012445 | 1.1943e-64 | 9.7209e-56 | 5.6728e-17 | 1 | 9.7209e-56 | 1.1609e-119 | -1.4448e-121 | 2.0873e-242 | 1.4448e-121
[6 empty] | 0.115 | 0.061167 | 31068066080.4261 | 31.0681 | 0.032187 | 11.7708 | 31945.4508 | 0.97228 | 2.6242e-65 | 4.0118e-56 | 4.9956e-17 | 1 | 4.0118e-56 | 1.0528e-120 | 1.0236e-120 | 1.0477e-240 | 1.0236e-120
[6 empty] | 0.1152 | 0.061273 | 31122097499.6964 | 31.1221 | 0.032132 | 11.7913 | 32162.0037 | -0.99982 | 5.6914e-66 | 1.648e-56 | 4.3982e-17 | 1 | 1.648e-56 | 9.3798e-122 | -9.3781e-122 | 8.7948e-243 | 9.3781e-122
[6 empty] | 0.1154 | 0.06138 | 31176128918.9667 | 31.1761 | 0.032076 | 11.8118 | 32379.6656 | 0.61298 | 1.2183e-66 | 6.7387e-57 | 3.8714e-17 | 1 | 6.7387e-57 | 8.2098e-123 | 5.0325e-123 | 2.5326e-245 | 5.0325e-123
[6 empty] | 0.1156 | 0.061486 | 31230160338.237 | 31.2302 | 0.03202 | 11.8323 | 32598.4404 | 0.97464 | 2.5737e-67 | 2.7426e-57 | 3.407e-17 | 1 | 2.7426e-57 | 7.0586e-124 | 6.8796e-124 | 4.7329e-247 | 6.8796e-124
[6 empty] | 0.1158 | 0.061593 | 31284191757.5073 | 31.2842 | 0.031965 | 11.8527 | 32818.332 | 0.96999 | 5.3651e-68 | 1.111e-57 | 2.9976e-17 | 1 | 1.111e-57 | 5.9606e-125 | 5.7818e-125 | 3.3429e-249 | 5.7818e-125
[6 empty] | 0.116 | 0.061699 | 31338223176.7776 | 31.3382 | 0.03191 | 11.8732 | 33039.3441 | 0.65617 | 1.1035e-68 | 4.4795e-58 | 2.6368e-17 | 1 | 4.4795e-58 | 4.9432e-126 | 3.2436e-126 | 1.0521e-251 | 3.2436e-126
[6 empty] | 0.1162 | 0.061805 | 31392254596.0479 | 31.3923 | 0.031855 | 11.8937 | 33261.4807 | -0.99811 | 2.2394e-69 | 1.7976e-58 | 2.319e-17 | 1 | 1.7976e-58 | 4.0256e-127 | -4.018e-127 | 1.6144e-253 | 4.018e-127
[6 empty] | 0.1164 | 0.061912 | 31446286015.3183 | 31.4463 | 0.0318 | 11.9141 | 33484.7455 | 0.98873 | 4.4833e-70 | 7.1797e-59 | 2.039e-17 | 1 | 7.1797e-59 | 3.2188e-128 | 3.1826e-128 | 1.0129e-255 | 3.1826e-128
[6 empty] | 0.1166 | 0.062018 | 31500317434.5886 | 31.5003 | 0.031746 | 11.9346 | 33709.1424 | -0.076672 | 8.8535e-71 | 2.854e-59 | 1.7924e-17 | 1 | 2.854e-59 | 2.5268e-129 | -1.9374e-130 | 3.7534e-260 | 1.9374e-130
[6 empty] | 0.1168 | 0.062124 | 31554348853.8589 | 31.5543 | 0.031691 | 11.9551 | 33934.6752 | -0.67324 | 1.7245e-71 | 1.1291e-59 | 1.5753e-17 | 1 | 1.1291e-59 | 1.9472e-130 | -1.3109e-130 | 1.7185e-260 | 1.3109e-130
[6 empty] | 0.117 | 0.062231 | 31608380273.1292 | 31.6084 | 0.031637 | 11.9756 | 34161.3479 | -0.25766 | 3.3127e-72 | 4.4459e-60 | 1.3841e-17 | 1 | 4.4459e-60 | 1.4728e-131 | -3.7948e-132 | 1.44e-263 | 3.7948e-132
[6 empty] | 0.1172 | 0.062337 | 31662411692.3995 | 31.6624 | 0.031583 | 11.996 | 34389.1642 | 0.97809 | 6.2756e-73 | 1.7422e-60 | 1.2159e-17 | 1 | 1.7422e-60 | 1.0933e-132 | 1.0694e-132 | 1.1435e-264 | 1.0694e-132
[6 empty] | 0.1174 | 0.062444 | 31716443111.6698 | 31.7164 | 0.031529 | 12.0165 | 34618.1281 | -0.83532 | 1.1723e-73 | 6.7942e-61 | 1.0679e-17 | 1 | 6.7942e-61 | 7.9646e-134 | -6.653e-134 | 4.4262e-267 | 6.653e-134
[6 empty] | 0.1176 | 0.06255 | 31770474530.9401 | 31.7705 | 0.031476 | 12.037 | 34848.2434 | 0.98066 | 2.159e-74 | 2.6369e-61 | 9.3774e-18 | 1 | 2.6369e-61 | 5.693e-135 | 5.5829e-135 | 3.1169e-269 | 5.5829e-135
[6 empty] | 0.1178 | 0.062656 | 31824505950.2104 | 31.8245 | 0.031422 | 12.0574 | 35079.5141 | 0.53171 | 3.9202e-75 | 1.0184e-61 | 8.2323e-18 | 1 | 1.0184e-61 | 3.9925e-136 | 2.1228e-136 | 4.5063e-272 | 2.1228e-136
[6 empty] | 0.118 | 0.062763 | 31878537369.4807 | 31.8785 | 0.031369 | 12.0779 | 35311.944 | 0.4905 | 7.0167e-76 | 3.9144e-62 | 7.2255e-18 | 1 | 3.9144e-62 | 2.7466e-137 | 1.3472e-137 | 1.8151e-274 | 1.3472e-137
[6 empty] | 0.1182 | 0.062869 | 31932568788.751 | 31.9326 | 0.031316 | 12.0984 | 35545.5371 | 0.99837 | 1.2379e-76 | 1.4972e-62 | 6.3404e-18 | 1 | 1.4972e-62 | 1.8535e-138 | 1.8504e-138 | 3.4241e-276 | 1.8504e-138
[6 empty] | 0.1184 | 0.062975 | 31986600208.0213 | 31.9866 | 0.031263 | 12.1189 | 35780.2972 | -0.69505 | 2.1525e-77 | 5.6988e-63 | 5.5625e-18 | 1 | 5.6988e-63 | 1.2267e-139 | -8.5261e-140 | 7.2694e-279 | 8.5261e-140
[6 empty] | 0.1186 | 0.063082 | 32040631627.2916 | 32.0406 | 0.03121 | 12.1393 | 36016.2282 | 0.882 | 3.6885e-78 | 2.1585e-63 | 4.879e-18 | 1 | 2.1585e-63 | 7.9616e-141 | 7.0221e-141 | 4.931e-281 | 7.0221e-141
[6 empty] | 0.1188 | 0.063188 | 32094663046.5619 | 32.0947 | 0.031158 | 12.1598 | 36253.3341 | -0.54386 | 6.2283e-79 | 8.1353e-64 | 4.2785e-18 | 1 | 8.1353e-64 | 5.0669e-142 | -2.7557e-142 | 7.5939e-284 | 2.7557e-142
[6 empty] | 0.119 | 0.063295 | 32148694465.8323 | 32.1487 | 0.031105 | 12.1803 | 36491.6189 | -0.86809 | 1.0362e-79 | 3.0511e-64 | 3.7511e-18 | 1 | 3.0511e-64 | 3.1615e-143 | -2.7445e-143 | 7.5323e-286 | 2.7445e-143
[6 empty] | 0.1192 | 0.063401 | 32202725885.1026 | 32.2027 | 0.031053 | 12.2007 | 36731.0863 | -0.33811 | 1.6984e-80 | 1.1386e-64 | 3.288e-18 | 1 | 1.1386e-64 | 1.9339e-144 | -6.5385e-145 | 4.2752e-289 | 6.5385e-145
[6 empty] | 0.1194 | 0.063507 | 32256757304.3729 | 32.2568 | 0.031001 | 12.2212 | 36971.7405 | 0.99975 | 2.7424e-81 | 4.2281e-65 | 2.8814e-18 | 1 | 4.2281e-65 | 1.1595e-145 | 1.1592e-145 | 1.3438e-290 | 1.1592e-145
[6 empty] | 0.1196 | 0.063614 | 32310788723.6432 | 32.3108 | 0.030949 | 12.2417 | 37213.5853 | -0.99677 | 4.3615e-82 | 1.5622e-65 | 2.5245e-18 | 1 | 1.5622e-65 | 6.8138e-147 | -6.7918e-147 | 4.6128e-293 | 6.7918e-147
[6 empty] | 0.1198 | 0.06372 | 32364820142.9135 | 32.3648 | 0.030898 | 12.2622 | 37456.6247 | 0.49204 | 6.832e-83 | 5.7434e-66 | 2.2114e-18 | 1 | 5.7434e-66 | 3.9239e-148 | 1.9307e-148 | 3.7277e-296 | 1.9307e-148
[6 empty] | 0.12 | 0.063827 | 32418851562.1838 | 32.4189 | 0.030846 | 12.2826 | 37700.8627 | 0.96889 | 1.0539e-83 | 2.1009e-66 | 1.9366e-18 | 1 | 2.1009e-66 | 2.2142e-149 | 2.1453e-149 | 4.6024e-298 | 2.1453e-149
[6 empty] | 0.1202 | 0.063933 | 32472882981.4541 | 32.4729 | 0.030795 | 12.3031 | 37946.3032 | 0.79829 | 1.6009e-84 | 7.6466e-67 | 1.6956e-18 | 1 | 7.6466e-67 | 1.2241e-150 | 9.7722e-151 | 9.5495e-301 | 9.7722e-151
[6 empty] | 0.1204 | 0.064039 | 32526914400.7244 | 32.5269 | 0.030744 | 12.3236 | 38192.9501 | -0.62748 | 2.3943e-85 | 2.769e-67 | 1.4843e-18 | 1 | 2.769e-67 | 6.6299e-152 | -4.1601e-152 | 1.7307e-303 | 4.1601e-152
[6 empty] | 0.1206 | 0.064146 | 32580945819.9947 | 32.5809 | 0.030693 | 12.344 | 38440.8076 | 0.34286 | 3.5255e-86 | 9.9765e-68 | 1.299e-18 | 1 | 9.9765e-68 | 3.5172e-153 | 1.2059e-153 | 1.4542e-306 | 1.2059e-153
[6 empty] | 0.1208 | 0.064252 | 32634977239.265 | 32.635 | 0.030642 | 12.3645 | 38689.8795 | -0.9445 | 5.11e-87 | 3.5762e-68 | 1.1366e-18 | 1 | 3.5762e-68 | 1.8275e-154 | -1.726e-154 | 2.9792e-308 | 1.726e-154
[6 empty] | 0.121 | 0.064358 | 32689008658.5353 | 32.689 | 0.030591 | 12.385 | 38940.1699 | -0.19775 | 7.2905e-88 | 1.2754e-68 | 9.9432e-19 | 1 | 1.2754e-68 | 9.2982e-156 | -1.8387e-156 | 0 | 1.8387e-156
[6 empty] | 0.1212 | 0.064465 | 32743040077.8056 | 32.743 | 0.030541 | 12.4055 | 39191.6828 | -0.37512 | 1.0237e-88 | 4.5252e-69 | 8.6963e-19 | 1 | 4.5252e-69 | 4.6325e-157 | -1.7377e-157 | 0 | 1.7377e-157
[6 empty] | 0.1214 | 0.064571 | 32797071497.0759 | 32.7971 | 0.030491 | 12.4259 | 39444.4222 | -0.97465 | 1.4145e-89 | 1.5974e-69 | 7.6041e-19 | 1 | 1.5974e-69 | 2.2596e-158 | -2.2023e-158 | 0 | 2.2023e-158
[6 empty] | 0.1216 | 0.064678 | 32851102916.3462 | 32.8511 | 0.03044 | 12.4464 | 39698.392 | 0.96283 | 1.9233e-90 | 5.6096e-70 | 6.6476e-19 | 1 | 5.6096e-70 | 1.0789e-159 | 1.0388e-159 | 0 | 1.0388e-159
[6 empty] | 0.1218 | 0.064784 | 32905134335.6166 | 32.9051 | 0.03039 | 12.4669 | 39953.5964 | -0.89523 | 2.573e-91 | 1.9598e-70 | 5.8101e-19 | 1 | 1.9598e-70 | 5.0424e-161 | -4.5141e-161 | 0 | 4.5141e-161
[6 empty] | 0.122 | 0.06489 | 32959165754.8869 | 32.9592 | 0.030341 | 12.4873 | 40210.0393 | -0.76107 | 3.3862e-92 | 6.8112e-71 | 5.077e-19 | 1 | 6.8112e-71 | 2.3064e-162 | -1.7554e-162 | 0 | 1.7554e-162
[6 empty] | 0.1222 | 0.064997 | 33013197174.1572 | 33.0132 | 0.030291 | 12.5078 | 40467.7248 | -0.80747 | 4.3837e-93 | 2.355e-71 | 4.4354e-19 | 1 | 2.355e-71 | 1.0324e-163 | -8.3359e-164 | 0 | 8.3359e-164
[6 empty] | 0.1224 | 0.065103 | 33067228593.4275 | 33.0672 | 0.030241 | 12.5283 | 40726.6569 | -0.77089 | 5.5819e-94 | 8.0997e-72 | 3.8741e-19 | 1 | 8.0997e-72 | 4.5212e-165 | -3.4854e-165 | 0 | 3.4854e-165
[6 empty] | 0.1226 | 0.065209 | 33121260012.6978 | 33.1213 | 0.030192 | 12.5487 | 40986.8396 | 0.99267 | 6.99e-95 | 2.7713e-72 | 3.383e-19 | 1 | 2.7713e-72 | 1.9371e-166 | 1.9229e-166 | 0 | 1.9229e-166
[6 empty] | 0.1228 | 0.065316 | 33175291431.9681 | 33.1753 | 0.030143 | 12.5692 | 41248.2771 | -0.69207 | 8.6076e-96 | 9.432e-73 | 2.9535e-19 | 1 | 9.432e-73 | 8.1187e-168 | -5.6187e-168 | 0 | 5.6187e-168
[6 empty] | 0.123 | 0.065422 | 33229322851.2384 | 33.2293 | 0.030094 | 12.5897 | 41510.9734 | -0.9245 | 1.0422e-96 | 3.1933e-73 | 2.578e-19 | 1 | 3.1933e-73 | 3.328e-169 | -3.0768e-169 | 0 | 3.0768e-169
[6 empty] | 0.1232 | 0.065529 | 33283354270.5087 | 33.2834 | 0.030045 | 12.6102 | 41774.9325 | -0.94742 | 1.2406e-97 | 1.0754e-73 | 2.2497e-19 | 1 | 1.0754e-73 | 1.3341e-170 | -1.264e-170 | 0 | 1.264e-170
[6 empty] | 0.1234 | 0.065635 | 33337385689.779 | 33.3374 | 0.029996 | 12.6306 | 42040.1586 | -0.53476 | 1.4518e-98 | 3.6024e-74 | 1.9628e-19 | 1 | 3.6024e-74 | 5.2298e-172 | -2.7967e-172 | 0 | 2.7967e-172
[6 empty] | 0.1236 | 0.065741 | 33391417109.0493 | 33.3914 | 0.029948 | 12.6511 | 42306.6557 | 0.89235 | 1.6698e-99 | 1.2003e-74 | 1.7121e-19 | 1 | 1.2003e-74 | 2.0043e-173 | 1.7886e-173 | 0 | 1.7886e-173
[6 empty] | 0.1238 | 0.065848 | 33445448528.3196 | 33.4454 | 0.029899 | 12.6716 | 42574.4279 | -0.35763 | 1.8876e-100 | 3.9782e-75 | 1.4931e-19 | 1 | 3.9782e-75 | 7.5094e-175 | -2.6856e-175 | 0 | 2.6856e-175
[6 empty] | 0.124 | 0.065954 | 33499479947.5899 | 33.4995 | 0.029851 | 12.692 | 42843.4792 | -0.99684 | 2.097e-101 | 1.3114e-75 | 1.3019e-19 | 1 | 1.3114e-75 | 2.75e-176 | -2.7413e-176 | 0 | 2.7413e-176
[6 empty] | 0.1242 | 0.06606 | 33553511366.8602 | 33.5535 | 0.029803 | 12.7125 | 43113.8139 | -0.97201 | 2.289e-102 | 4.2999e-76 | 1.1348e-19 | 1 | 4.2999e-76 | 9.8425e-178 | -9.567e-178 | 0 | 9.567e-178
[6 empty] | 0.1244 | 0.066167 | 33607542786.1305 | 33.6075 | 0.029755 | 12.733 | 43385.436 | 0.11124 | 2.4548e-103 | 1.4023e-76 | 9.8901e-20 | 1 | 1.4023e-76 | 3.4424e-179 | 3.8292e-180 | 0 | 3.8292e-180
[6 empty] | 0.1246 | 0.066273 | 33661574205.4009 | 33.6616 | 0.029707 | 12.7535 | 43658.3495 | 0.28931 | 2.5863e-104 | 4.5484e-77 | 8.6173e-20 | 1 | 4.5484e-77 | 1.1763e-180 | 3.4033e-181 | 0 | 3.4033e-181
[6 empty] | 0.1248 | 0.06638 | 33715605624.6712 | 33.7156 | 0.02966 | 12.7739 | 43932.5588 | 0.56229 | 2.6765e-105 | 1.4673e-77 | 7.5068e-20 | 1 | 1.4673e-77 | 3.9272e-182 | 2.2082e-182 | 0 | 2.2082e-182
[6 empty] | 0.125 | 0.066486 | 33769637043.9415 | 33.7696 | 0.029612 | 12.7944 | 44208.0678 | -0.34666 | 2.7204e-106 | 4.7077e-78 | 6.5378e-20 | 1 | 4.7077e-78 | 1.2807e-183 | -4.4396e-184 | 0 | 4.4396e-184
[6 empty] | 0.1252 | 0.066592 | 33823668463.2118 | 33.8237 | 0.029565 | 12.8149 | 44484.8806 | -0.0012733 | 2.7153e-107 | 1.5022e-78 | 5.6927e-20 | 1 | 1.5022e-78 | 4.079e-185 | -5.1939e-188 | 0 | 5.1939e-188
[6 empty] | 0.1254 | 0.066699 | 33877699882.4821 | 33.8777 | 0.029518 | 12.8353 | 44763.0016 | 0.99607 | 2.6613e-108 | 4.7671e-79 | 4.9558e-20 | 1 | 4.7671e-79 | 1.2687e-186 | 1.2637e-186 | 0 | 1.2637e-186
[6 empty] | 0.1256 | 0.066805 | 33931731301.7524 | 33.9317 | 0.029471 | 12.8558 | 45042.4347 | -0.99681 | 2.561e-109 | 1.5045e-79 | 4.3133e-20 | 1 | 1.5045e-79 | 3.8531e-188 | -3.8408e-188 | 0 | 3.8408e-188
[6 empty] | 0.1258 | 0.066911 | 33985762721.0227 | 33.9858 | 0.029424 | 12.8763 | 45323.1842 | 0.48204 | 2.4193e-110 | 4.7222e-80 | 3.7532e-20 | 1 | 4.7222e-80 | 1.1424e-189 | 5.5071e-190 | 0 | 5.5071e-190
[6 empty] | 0.126 | 0.067018 | 34039794140.293 | 34.0398 | 0.029377 | 12.8968 | 45605.2541 | 0.92321 | 2.2434e-111 | 1.474e-80 | 3.2652e-20 | 1 | 1.474e-80 | 3.3066e-191 | 3.0527e-191 | 0 | 3.0527e-191
[6 empty] | 0.1262 | 0.067124 | 34093825559.5633 | 34.0938 | 0.029331 | 12.9172 | 45888.6488 | 0.50125 | 2.0417e-112 | 4.5752e-81 | 2.8399e-20 | 1 | 4.5752e-81 | 9.3411e-193 | 4.6823e-193 | 0 | 4.6823e-193
[6 empty] | 0.1264 | 0.067231 | 34147856978.8336 | 34.1479 | 0.029284 | 12.9377 | 46173.3723 | -0.99332 | 1.8235e-113 | 1.4123e-81 | 2.4696e-20 | 1 | 1.4123e-81 | 2.5753e-194 | -2.558e-194 | 0 | 2.558e-194
[6 empty] | 0.1266 | 0.067337 | 34201888398.1039 | 34.2019 | 0.029238 | 12.9582 | 46459.4288 | 0.99844 | 1.5981e-114 | 4.335e-82 | 2.147e-20 | 1 | 4.335e-82 | 6.9276e-196 | 6.9168e-196 | 0 | 6.9168e-196
[6 empty] | 0.1268 | 0.067443 | 34255919817.3742 | 34.2559 | 0.029192 | 12.9786 | 46746.8225 | -0.0060828 | 1.3741e-115 | 1.3232e-82 | 1.8662e-20 | 1 | 1.3232e-82 | 1.8182e-197 | -1.106e-199 | 0 | 1.106e-199
[6 empty] | 0.127 | 0.06755 | 34309951236.6445 | 34.31 | 0.029146 | 12.9991 | 47035.5577 | -0.29309 | 1.159e-116 | 4.0163e-83 | 1.6217e-20 | 1 | 4.0163e-83 | 4.655e-199 | -1.3643e-199 | 0 | 1.3643e-199
[6 empty] | 0.1272 | 0.067656 | 34363982655.9148 | 34.364 | 0.0291 | 13.0196 | 47325.6385 | 0.68658 | 9.5896e-118 | 1.2122e-83 | 1.409e-20 | 1 | 1.2122e-83 | 1.1624e-200 | 7.9809e-201 | 0 | 7.9809e-201
[6 empty] | 0.1274 | 0.067762 | 34418014075.1852 | 34.418 | 0.029055 | 13.0401 | 47617.069 | -0.019234 | 7.7816e-119 | 3.6378e-84 | 1.2238e-20 | 1 | 3.6378e-84 | 2.8308e-202 | -5.4447e-204 | 0 | 5.4447e-204
[6 empty] | 0.1276 | 0.067869 | 34472045494.4555 | 34.472 | 0.029009 | 13.0605 | 47909.8536 | 0.59376 | 6.1923e-120 | 1.0856e-84 | 1.0628e-20 | 1 | 1.0856e-84 | 6.7222e-204 | 3.9914e-204 | 0 | 3.9914e-204
[6 empty] | 0.1278 | 0.067975 | 34526076913.7258 | 34.5261 | 0.028964 | 13.081 | 48203.9965 | -0.5065 | 4.8317e-121 | 3.221e-85 | 9.2278e-21 | 1 | 3.221e-85 | 1.5563e-205 | -7.8828e-206 | 0 | 7.8828e-206
[6 empty] | 0.128 | 0.068082 | 34580108332.9961 | 34.5801 | 0.028918 | 13.1015 | 48499.5018 | -0.32923 | 3.6962e-122 | 9.5028e-86 | 8.0101e-21 | 1 | 9.5028e-86 | 3.5124e-207 | -1.1564e-207 | 0 | 1.1564e-207
[6 empty] | 0.1282 | 0.068188 | 34634139752.2664 | 34.6341 | 0.028873 | 13.1219 | 48796.3739 | 0.94143 | 2.7718e-123 | 2.7875e-86 | 6.9516e-21 | 1 | 2.7875e-86 | 7.7265e-209 | 7.2739e-209 | 0 | 7.2739e-209
[6 empty] | 0.1284 | 0.068294 | 34688171171.5367 | 34.6882 | 0.028828 | 13.1424 | 49094.6169 | -0.85135 | 2.0374e-124 | 8.1299e-87 | 6.0316e-21 | 1 | 8.1299e-87 | 1.6564e-210 | -1.4101e-210 | 0 | 1.4101e-210
[6 empty] | 0.1286 | 0.068401 | 34742202590.807 | 34.7422 | 0.028783 | 13.1629 | 49394.2351 | 0.81697 | 1.4677e-125 | 2.3575e-87 | 5.2322e-21 | 1 | 2.3575e-87 | 3.46e-212 | 2.8267e-212 | 0 | 2.8267e-212
[6 empty] | 0.1288 | 0.068507 | 34796234010.0773 | 34.7962 | 0.028739 | 13.1834 | 49695.2328 | 0.99981 | 1.036e-126 | 6.7966e-88 | 4.5377e-21 | 1 | 6.7966e-88 | 7.0415e-214 | 7.0402e-214 | 0 | 7.0402e-214
[6 empty] | 0.129 | 0.068614 | 34850265429.3476 | 34.8503 | 0.028694 | 13.2038 | 49997.6141 | 0.69101 | 7.1658e-128 | 1.9481e-88 | 3.9346e-21 | 1 | 1.9481e-88 | 1.396e-215 | 9.6463e-216 | 0 | 9.6463e-216
[6 empty] | 0.1292 | 0.06872 | 34904296848.6179 | 34.9043 | 0.02865 | 13.2243 | 50301.3835 | -0.98767 | 4.8555e-129 | 5.5514e-89 | 3.4109e-21 | 1 | 5.5514e-89 | 2.6955e-217 | -2.6623e-217 | 0 | 2.6623e-217
[6 empty] | 0.1294 | 0.068826 | 34958328267.8882 | 34.9583 | 0.028605 | 13.2448 | 50606.5451 | 0.96379 | 3.2228e-130 | 1.5728e-89 | 2.9562e-21 | 1 | 1.5728e-89 | 5.0687e-219 | 4.8852e-219 | 0 | 4.8852e-219
[6 empty] | 0.1296 | 0.068933 | 35012359687.1585 | 35.0124 | 0.028561 | 13.2652 | 50913.1033 | 0.49929 | 2.0951e-131 | 4.4297e-90 | 2.5616e-21 | 1 | 4.4297e-90 | 9.2808e-221 | 4.6338e-221 | 0 | 4.6338e-221
[6 empty] | 0.1298 | 0.069039 | 35066391106.4288 | 35.0664 | 0.028517 | 13.2857 | 51221.0622 | 0.56931 | 1.3338e-132 | 1.2403e-90 | 2.2191e-21 | 1 | 1.2403e-90 | 1.6544e-222 | 9.4185e-223 | 0 | 9.4185e-223
[6 empty] | 0.13 | 0.069145 | 35120422525.6992 | 35.1204 | 0.028473 | 13.3062 | 51530.4263 | 0.86638 | 8.315e-134 | 3.4524e-91 | 1.922e-21 | 1 | 3.4524e-91 | 2.8707e-224 | 2.4871e-224 | 0 | 2.4871e-224
[6 empty] | 0.1302 | 0.069252 | 35174453944.9695 | 35.1745 | 0.02843 | 13.3266 | 51841.1998 | -0.96141 | 5.0748e-135 | 9.5532e-92 | 1.6643e-21 | 1 | 9.5532e-92 | 4.8481e-226 | -4.661e-226 | 0 | 4.661e-226
[6 empty] | 0.1304 | 0.069358 | 35228485364.2398 | 35.2285 | 0.028386 | 13.3471 | 52153.3871 | 0.12223 | 3.032e-136 | 2.6278e-92 | 1.4409e-21 | 1 | 2.6278e-92 | 7.9677e-228 | 9.7386e-229 | 0 | 9.7386e-229
[6 empty] | 0.1306 | 0.069465 | 35282516783.5101 | 35.2825 | 0.028343 | 13.3676 | 52466.9923 | 0.6261 | 1.7731e-137 | 7.1855e-93 | 1.2472e-21 | 1 | 7.1855e-93 | 1.2741e-229 | 7.977e-230 | 0 | 7.977e-230
[6 empty] | 0.1308 | 0.069571 | 35336548202.7804 | 35.3365 | 0.028299 | 13.3881 | 52782.0199 | -0.19064 | 1.0148e-138 | 1.9531e-93 | 1.0792e-21 | 1 | 1.9531e-93 | 1.982e-231 | -3.7785e-232 | 0 | 3.7785e-232
[6 empty] | 0.131 | 0.069677 | 35390579622.0507 | 35.3906 | 0.028256 | 13.4085 | 53098.4742 | -0.60896 | 5.6834e-140 | 5.2771e-94 | 9.3372e-22 | 1 | 5.2771e-94 | 2.9992e-233 | -1.8264e-233 | 0 | 1.8264e-233
[6 empty] | 0.1312 | 0.069784 | 35444611041.321 | 35.4446 | 0.028213 | 13.429 | 53416.3595 | 0.070236 | 3.1143e-141 | 1.4173e-94 | 8.0764e-22 | 1 | 1.4173e-94 | 4.4138e-235 | 3.1001e-236 | 0 | 3.1001e-236
[6 empty] | 0.1314 | 0.06989 | 35498642460.5913 | 35.4986 | 0.02817 | 13.4495 | 53735.6802 | 0.92914 | 1.6694e-142 | 3.7836e-95 | 6.9844e-22 | 1 | 3.7836e-95 | 6.3164e-237 | 5.8688e-237 | 0 | 5.8688e-237
[6 empty] | 0.1316 | 0.069996 | 35552673879.8616 | 35.5527 | 0.028127 | 13.4699 | 54056.4406 | 0.767 | 8.7537e-144 | 1.004e-95 | 6.0386e-22 | 1 | 1.004e-95 | 8.7884e-239 | 6.7408e-239 | 0 | 6.7408e-239
[6 empty] | 0.1318 | 0.070103 | 35606705299.1319 | 35.6067 | 0.028085 | 13.4904 | 54378.645 | -0.77569 | 4.4892e-145 | 2.6479e-96 | 5.2198e-22 | 1 | 2.6479e-96 | 1.1887e-240 | -9.2206e-241 | 0 | 9.2206e-241
[6 empty] | 0.132 | 0.070209 | 35660736718.4022 | 35.6607 | 0.028042 | 13.5109 | 54702.2978 | 0.81724 | 2.2513e-146 | 6.9414e-97 | 4.511e-22 | 1 | 6.9414e-97 | 1.5627e-242 | 1.2771e-242 | 0 | 1.2771e-242
[6 empty] | 0.1322 | 0.070316 | 35714768137.6725 | 35.7148 | 0.028 | 13.5314 | 55027.4034 | -0.61581 | 1.1039e-147 | 1.8086e-97 | 3.8976e-22 | 1 | 1.8086e-97 | 1.9965e-244 | -1.2295e-244 | 0 | 1.2295e-244
[6 empty] | 0.1324 | 0.070422 | 35768799556.9428 | 35.7688 | 0.027957 | 13.5518 | 55353.9662 | -0.73542 | 5.2917e-149 | 4.6836e-98 | 3.3668e-22 | 1 | 4.6836e-98 | 2.4784e-246 | -1.8227e-246 | 0 | 1.8227e-246
[6 empty] | 0.1326 | 0.070528 | 35822830976.2131 | 35.8228 | 0.027915 | 13.5723 | 55681.9905 | 0.45501 | 2.4796e-150 | 1.2055e-98 | 2.9077e-22 | 1 | 1.2055e-98 | 2.9891e-248 | 1.36e-248 | 0 | 1.36e-248
[6 empty] | 0.1328 | 0.070635 | 35876862395.4835 | 35.8769 | 0.027873 | 13.5928 | 56011.4807 | -0.095226 | 1.1356e-151 | 3.0836e-99 | 2.5106e-22 | 1 | 3.0836e-99 | 3.5017e-250 | -3.3345e-251 | 0 | 3.3345e-251
[6 empty] | 0.133 | 0.070741 | 35930893814.7538 | 35.9309 | 0.027831 | 13.6132 | 56342.4413 | 0.92787 | 5.0823e-153 | 7.8392e-100 | 2.1673e-22 | 1 | 7.8392e-100 | 3.9841e-252 | 3.6968e-252 | 0 | 3.6968e-252
[6 empty] | 0.1332 | 0.070847 | 35984925234.0241 | 35.9849 | 0.027789 | 13.6337 | 56674.8765 | 0.57703 | 2.2225e-154 | 1.9806e-100 | 1.8705e-22 | 1 | 1.9806e-100 | 4.4019e-254 | 2.54e-254 | 0 | 2.54e-254
[6 empty] | 0.1334 | 0.070954 | 36038956653.2944 | 36.039 | 0.027748 | 13.6542 | 57008.7909 | 0.99874 | 9.4949e-156 | 4.9732e-101 | 1.614e-22 | 1 | 4.9732e-101 | 4.722e-256 | 4.716e-256 | 0 | 4.716e-256
[6 empty] | 0.1336 | 0.07106 | 36092988072.5647 | 36.093 | 0.027706 | 13.6747 | 57344.1887 | -0.69477 | 3.9624e-157 | 1.241e-101 | 1.3924e-22 | 1 | 1.241e-101 | 4.9172e-258 | -3.4163e-258 | 0 | 3.4163e-258
[6 empty] | 0.1338 | 0.071167 | 36147019491.835 | 36.147 | 0.027665 | 13.6951 | 57681.0746 | 0.99774 | 1.615e-158 | 3.0773e-102 | 1.2009e-22 | 1 | 3.0773e-102 | 4.9699e-260 | 4.9587e-260 | 0 | 4.9587e-260
[6 empty] | 0.134 | 0.071273 | 36201050911.1053 | 36.2011 | 0.027624 | 13.7156 | 58019.4528 | 0.55612 | 6.4283e-160 | 7.5834e-103 | 1.0356e-22 | 1 | 7.5834e-103 | 4.8748e-262 | 2.711e-262 | 0 | 2.711e-262
[6 empty] | 0.1342 | 0.071379 | 36255082330.3756 | 36.2551 | 0.027582 | 13.7361 | 58359.3278 | 0.92181 | 2.4983e-161 | 1.8571e-103 | 8.9276e-23 | 1 | 1.8571e-103 | 4.6395e-264 | 4.2767e-264 | 0 | 4.2767e-264
[6 empty] | 0.1344 | 0.071486 | 36309113749.6459 | 36.3091 | 0.027541 | 13.7565 | 58700.7041 | -0.11512 | 9.4791e-163 | 4.5191e-104 | 7.6949e-23 | 1 | 4.5191e-104 | 4.2837e-266 | -4.9315e-267 | 0 | 4.9315e-267
[6 empty] | 0.1346 | 0.071592 | 36363145168.9162 | 36.3631 | 0.0275 | 13.777 | 59043.586 | 0.53434 | 3.5107e-164 | 1.0928e-104 | 6.6309e-23 | 1 | 1.0928e-104 | 3.8364e-268 | 2.05e-268 | 0 | 2.05e-268
[6 empty] | 0.1348 | 0.071698 | 36417176588.1865 | 36.4172 | 0.02746 | 13.7975 | 59387.9781 | -0.58052 | 1.2691e-165 | 2.6257e-105 | 5.7127e-23 | 1 | 2.6257e-105 | 3.3322e-270 | -1.9344e-270 | 0 | 1.9344e-270
[6 empty] | 0.135 | 0.071805 | 36471208007.4568 | 36.4712 | 0.027419 | 13.818 | 59733.8848 | -0.28391 | 4.4766e-167 | 6.2691e-106 | 4.9206e-23 | 1 | 6.2691e-106 | 2.8064e-272 | -7.9677e-273 | 0 | 7.9677e-273
[6 empty] | 0.1352 | 0.071911 | 36525239426.7271 | 36.5252 | 0.027378 | 13.8384 | 60081.3105 | 0.99997 | 1.5408e-168 | 1.4873e-106 | 4.2374e-23 | 1 | 1.4873e-106 | 2.2916e-274 | 2.2915e-274 | 0 | 2.2915e-274
[6 empty] | 0.1354 | 0.072018 | 36579270845.9974 | 36.5793 | 0.027338 | 13.8589 | 60430.2599 | -0.97492 | 5.1735e-170 | 3.506e-107 | 3.6483e-23 | 1 | 3.506e-107 | 1.8138e-276 | -1.7683e-276 | 0 | 1.7683e-276
[6 empty] | 0.1356 | 0.072124 | 36633302265.2678 | 36.6333 | 0.027298 | 13.8794 | 60780.7372 | -0.40242 | 1.6944e-171 | 8.2117e-108 | 3.1403e-23 | 1 | 8.2117e-108 | 1.3914e-278 | -5.5993e-279 | 0 | 5.5993e-279
[6 empty] | 0.1358 | 0.07223 | 36687333684.5381 | 36.6873 | 0.027257 | 13.8998 | 61132.747 | -0.53593 | 5.4126e-173 | 1.911e-108 | 2.7025e-23 | 1 | 1.911e-108 | 1.0343e-280 | -5.5433e-281 | 0 | 5.5433e-281
[6 empty] | 0.136 | 0.072337 | 36741365103.8084 | 36.7414 | 0.027217 | 13.9203 | 61486.2938 | -0.77552 | 1.686e-174 | 4.4186e-109 | 2.3252e-23 | 1 | 4.4186e-109 | 7.4496e-283 | -5.7774e-283 | 0 | 5.7774e-283
[6 empty] | 0.1362 | 0.072443 | 36795396523.0787 | 36.7954 | 0.027177 | 13.9408 | 61841.3821 | 0.7168 | 5.1204e-176 | 1.0151e-109 | 2.0002e-23 | 1 | 1.0151e-109 | 5.1976e-285 | 3.7256e-285 | 0 | 3.7256e-285
[6 empty] | 0.1364 | 0.072549 | 36849427942.349 | 36.8494 | 0.027137 | 13.9613 | 62198.0164 | 0.74139 | 1.516e-177 | 2.3168e-110 | 1.7202e-23 | 1 | 2.3168e-110 | 3.5123e-287 | 2.604e-287 | 0 | 2.604e-287
[6 empty] | 0.1366 | 0.072656 | 36903459361.6193 | 36.9035 | 0.027098 | 13.9817 | 62556.2013 | 0.76972 | 4.375e-179 | 5.2535e-111 | 1.479e-23 | 1 | 5.2535e-111 | 2.2984e-289 | 1.7691e-289 | 0 | 1.7691e-289
[6 empty] | 0.1368 | 0.072762 | 36957490780.8896 | 36.9575 | 0.027058 | 14.0022 | 62915.9411 | 0.61697 | 1.2304e-180 | 1.1835e-111 | 1.2714e-23 | 1 | 1.1835e-111 | 1.4562e-291 | 8.9847e-292 | 0 | 8.9847e-292
[6 empty] | 0.137 | 0.072869 | 37011522200.1599 | 37.0115 | 0.027019 | 14.0227 | 63277.2405 | -0.6041 | 3.372e-182 | 2.6488e-112 | 1.0927e-23 | 1 | 2.6488e-112 | 8.932e-294 | -5.3958e-294 | 0 | 5.3958e-294
[6 empty] | 0.1372 | 0.072975 | 37065553619.4302 | 37.0656 | 0.026979 | 14.0431 | 63640.104 | -0.80264 | 9.0034e-184 | 5.8896e-113 | 9.3891e-24 | 1 | 5.8896e-113 | 5.3026e-296 | -4.2561e-296 | 0 | 4.2561e-296
[6 empty] | 0.1374 | 0.073081 | 37119585038.7005 | 37.1196 | 0.02694 | 14.0636 | 64004.5362 | -0.80702 | 2.3417e-185 | 1.3009e-113 | 8.0658e-24 | 1 | 1.3009e-113 | 3.0464e-298 | -2.4585e-298 | 0 | 2.4585e-298
[6 empty] | 0.1376 | 0.073188 | 37173616457.9708 | 37.1736 | 0.026901 | 14.0841 | 64370.5415 | -0.58263 | 5.9322e-187 | 2.8546e-114 | 6.9275e-24 | 1 | 2.8546e-114 | 1.6934e-300 | -9.8663e-301 | 0 | 9.8663e-301
[6 empty] | 0.1378 | 0.073294 | 37227647877.2411 | 37.2276 | 0.026862 | 14.1045 | 64738.1245 | 0.569 | 1.4634e-188 | 6.2225e-115 | 5.9485e-24 | 1 | 6.2225e-115 | 9.1062e-303 | 5.1814e-303 | 0 | 5.1814e-303
[6 empty] | 0.138 | 0.0734 | 37281679296.5114 | 37.2817 | 0.026823 | 14.125 | 65107.2898 | 0.83803 | 3.5152e-190 | 1.3474e-115 | 5.1067e-24 | 1 | 1.3474e-115 | 4.7364e-305 | 3.9693e-305 | 0 | 3.9693e-305
[6 empty] | 0.1382 | 0.073507 | 37335710715.7817 | 37.3357 | 0.026784 | 14.1455 | 65478.0419 | 0.86133 | 8.2203e-192 | 2.8982e-116 | 4.3831e-24 | 1 | 2.8982e-116 | 2.3824e-307 | 2.052e-307 | 0 | 2.052e-307
[6 empty] | 0.1384 | 0.073613 | 37389742135.052 | 37.3897 | 0.026745 | 14.166 | 65850.3855 | 0.45119 | 1.8711e-193 | 6.1924e-117 | 3.7612e-24 | 1 | 6.1924e-117 | 0 | 0 | 0 | 0
[6 empty] | 0.1386 | 0.07372 | 37443773554.3223 | 37.4438 | 0.026707 | 14.1864 | 66224.325 | -0.36912 | 4.145e-195 | 1.3143e-117 | 3.2268e-24 | 1 | 1.3143e-117 | 0 | 0 | 0 | 0
[6 empty] | 0.1388 | 0.073826 | 37497804973.5927 | 37.4978 | 0.026668 | 14.2069 | 66599.8651 | -0.96684 | 8.9352e-197 | 2.7706e-118 | 2.7677e-24 | 1 | 2.7706e-118 | 0 | 0 | 0 | 0
[6 empty] | 0.139 | 0.073932 | 37551836392.863 | 37.5518 | 0.02663 | 14.2274 | 66977.0104 | -0.99459 | 1.874e-198 | 5.8016e-119 | 2.3735e-24 | 1 | 5.8016e-119 | 0 | 0 | 0 | 0
[6 empty] | 0.1392 | 0.074039 | 37605867812.1333 | 37.6059 | 0.026592 | 14.2478 | 67355.7655 | 0.088963 | 3.8232e-200 | 1.2066e-119 | 2.0349e-24 | 1 | 1.2066e-119 | 0 | 0 | 0 | 0
[6 empty] | 0.1394 | 0.074145 | 37659899231.4036 | 37.6599 | 0.026553 | 14.2683 | 67736.135 | -0.32008 | 7.5865e-202 | 2.4927e-120 | 1.7442e-24 | 1 | 2.4927e-120 | 0 | 0 | 0 | 0
[6 empty] | 0.1396 | 0.074252 | 37713930650.6739 | 37.7139 | 0.026515 | 14.2888 | 68118.1234 | 0.81923 | 1.464e-203 | 5.1144e-121 | 1.4948e-24 | 1 | 5.1144e-121 | 0 | 0 | 0 | 0
[6 empty] | 0.1398 | 0.074358 | 37767962069.9442 | 37.768 | 0.026477 | 14.3093 | 68501.7354 | 0.58294 | 2.7468e-205 | 1.0422e-121 | 1.2807e-24 | 1 | 1.0422e-121 | 0 | 0 | 0 | 0
[6 empty] | 0.14 | 0.074464 | 37821993489.2145 | 37.822 | 0.02644 | 14.3297 | 68886.9757 | -0.9743 | 5.0101e-207 | 2.1094e-122 | 1.097e-24 | 1 | 2.1094e-122 | 0 | 0 | 0 | 0
[6 empty] | 0.1402 | 0.074571 | 37876024908.4848 | 37.876 | 0.026402 | 14.3502 | 69273.8488 | 0.97366 | 8.8826e-209 | 4.2402e-123 | 9.3949e-25 | 1 | 4.2402e-123 | 0 | 0 | 0 | 0
[6 empty] | 0.1404 | 0.074677 | 37930056327.7551 | 37.9301 | 0.026364 | 14.3707 | 69662.3594 | 0.68452 | 1.5305e-210 | 8.4647e-124 | 8.0441e-25 | 1 | 8.4647e-124 | 0 | 0 | 0 | 0
[6 empty] | 0.1406 | 0.074783 | 37984087747.0254 | 37.9841 | 0.026327 | 14.3911 | 70052.5121 | 0.97555 | 2.5624e-212 | 1.6782e-124 | 6.8859e-25 | 1 | 1.6782e-124 | 0 | 0 | 0 | 0
[6 empty] | 0.1408 | 0.07489 | 38038119166.2957 | 38.0381 | 0.026289 | 14.4116 | 70444.3117 | -0.43466 | 4.1678e-214 | 3.3043e-125 | 5.8932e-25 | 1 | 3.3043e-125 | 0 | 0 | 0 | 0
[6 empty] | 0.141 | 0.074996 | 38092150585.566 | 38.0922 | 0.026252 | 14.4321 | 70837.7627 | 0.93261 | 6.5851e-216 | 6.4608e-126 | 5.0425e-25 | 1 | 6.4608e-126 | 0 | 0 | 0 | 0
[6 empty] | 0.1412 | 0.075103 | 38146182004.8364 | 38.1462 | 0.026215 | 14.4526 | 71232.8698 | 0.4511 | 1.0105e-217 | 1.2545e-126 | 4.3136e-25 | 1 | 1.2545e-126 | 0 | 0 | 0 | 0
[6 empty] | 0.1414 | 0.075209 | 38200213424.1067 | 38.2002 | 0.026178 | 14.473 | 71629.6376 | 0.98462 | 1.5057e-219 | 2.419e-127 | 3.6893e-25 | 1 | 2.419e-127 | 0 | 0 | 0 | 0
[6 empty] | 0.1416 | 0.075315 | 38254244843.377 | 38.2542 | 0.026141 | 14.4935 | 72028.071 | -0.74882 | 2.1782e-221 | 4.6319e-128 | 3.1546e-25 | 1 | 4.6319e-128 | 0 | 0 | 0 | 0
[6 empty] | 0.1418 | 0.075422 | 38308276262.6473 | 38.3083 | 0.026104 | 14.514 | 72428.1745 | 0.92234 | 3.059e-223 | 8.8071e-129 | 2.6969e-25 | 1 | 8.8071e-129 | 0 | 0 | 0 | 0
[6 empty] | 0.142 | 0.075528 | 38362307681.9176 | 38.3623 | 0.026067 | 14.5344 | 72829.9528 | 0.99869 | 4.1694e-225 | 1.6628e-129 | 2.305e-25 | 1 | 1.6628e-129 | 0 | 0 | 0 | 0
[6 empty] | 0.1422 | 0.075634 | 38416339101.1879 | 38.4163 | 0.026031 | 14.5549 | 73233.4106 | 0.18459 | 5.5147e-227 | 3.1174e-130 | 1.9697e-25 | 1 | 3.1174e-130 | 0 | 0 | 0 | 0
[6 empty] | 0.1424 | 0.075741 | 38470370520.4582 | 38.4704 | 0.025994 | 14.5754 | 73638.5527 | -0.30412 | 7.0769e-229 | 5.8032e-131 | 1.6827e-25 | 1 | 5.8032e-131 | 0 | 0 | 0 | 0
[6 empty] | 0.1426 | 0.075847 | 38524401939.7285 | 38.5244 | 0.025958 | 14.5959 | 74045.3838 | -0.95103 | 8.81e-231 | 1.0727e-131 | 1.4373e-25 | 1 | 1.0727e-131 | 0 | 0 | 0 | 0
[6 empty] | 0.1428 | 0.075954 | 38578433358.9988 | 38.5784 | 0.025921 | 14.6163 | 74453.9085 | -0.98074 | 1.0637e-232 | 1.9686e-132 | 1.2273e-25 | 1 | 1.9686e-132 | 0 | 0 | 0 | 0
[6 empty] | 0.143 | 0.07606 | 38632464778.2691 | 38.6325 | 0.025885 | 14.6368 | 74864.1315 | 0.048647 | 1.2455e-234 | 3.5873e-133 | 1.0479e-25 | 1 | 3.5873e-133 | 0 | 0 | 0 | 0
[6 empty] | 0.1432 | 0.076166 | 38686496197.5394 | 38.6865 | 0.025849 | 14.6573 | 75276.0577 | -0.41344 | 1.414e-236 | 6.4902e-134 | 8.9441e-26 | 1 | 6.4902e-134 | 0 | 0 | 0 | 0
[6 empty] | 0.1434 | 0.076273 | 38740527616.8097 | 38.7405 | 0.025813 | 14.6777 | 75689.6918 | 0.5891 | 1.5562e-238 | 1.1658e-134 | 7.6327e-26 | 1 | 1.1658e-134 | 0 | 0 | 0 | 0
[6 empty] | 0.1436 | 0.076379 | 38794559036.08 | 38.7946 | 0.025777 | 14.6982 | 76105.0385 | -0.026478 | 1.66e-240 | 2.0791e-135 | 6.5121e-26 | 1 | 2.0791e-135 | 0 | 0 | 0 | 0
[6 empty] | 0.1438 | 0.076485 | 38848590455.3503 | 38.8486 | 0.025741 | 14.7187 | 76522.1025 | -0.67528 | 1.716e-242 | 3.6812e-136 | 5.5548e-26 | 1 | 3.6812e-136 | 0 | 0 | 0 | 0
[6 empty] | 0.144 | 0.076592 | 38902621874.6207 | 38.9026 | 0.025705 | 14.7391 | 76940.8886 | -0.21134 | 1.7187e-244 | 6.4708e-137 | 4.7372e-26 | 1 | 6.4708e-137 | 0 | 0 | 0 | 0
[6 empty] | 0.1442 | 0.076698 | 38956653293.891 | 38.9567 | 0.02567 | 14.7596 | 77361.4015 | 0.24497 | 1.6676e-246 | 1.1292e-137 | 4.039e-26 | 1 | 1.1292e-137 | 0 | 0 | 0 | 0
[6 empty] | 0.1444 | 0.076805 | 39010684713.1613 | 39.0107 | 0.025634 | 14.7801 | 77783.6462 | -0.85404 | 1.5671e-248 | 1.9563e-138 | 3.443e-26 | 1 | 1.9563e-138 | 0 | 0 | 0 | 0
[6 empty] | 0.1446 | 0.076911 | 39064716132.4316 | 39.0647 | 0.025599 | 14.8006 | 78207.6272 | 0.7769 | 1.4262e-250 | 3.3645e-139 | 2.9343e-26 | 1 | 3.3645e-139 | 0 | 0 | 0 | 0
[6 empty] | 0.1448 | 0.077017 | 39118747551.7019 | 39.1187 | 0.025563 | 14.821 | 78633.3494 | -0.60091 | 1.2567e-252 | 5.7441e-140 | 2.5002e-26 | 1 | 5.7441e-140 | 0 | 0 | 0 | 0
[6 empty] | 0.145 | 0.077124 | 39172778970.9722 | 39.1728 | 0.025528 | 14.8415 | 79060.8176 | -0.41962 | 1.072e-254 | 9.7351e-141 | 2.1298e-26 | 1 | 9.7351e-141 | 0 | 0 | 0 | 0
[6 empty] | 0.1452 | 0.07723 | 39226810390.2425 | 39.2268 | 0.025493 | 14.862 | 79490.0366 | 0.99914 | 8.8504e-257 | 1.6378e-141 | 1.8139e-26 | 1 | 1.6378e-141 | 0 | 0 | 0 | 0
[6 empty] | 0.1454 | 0.077336 | 39280841809.5128 | 39.2808 | 0.025458 | 14.8824 | 79921.0112 | -0.86028 | 7.071e-259 | 2.7351e-142 | 1.5445e-26 | 1 | 2.7351e-142 | 0 | 0 | 0 | 0
[6 empty] | 0.1456 | 0.077443 | 39334873228.7831 | 39.3349 | 0.025423 | 14.9029 | 80353.7462 | -0.96382 | 5.4659e-261 | 4.534e-143 | 1.3149e-26 | 1 | 4.534e-143 | 0 | 0 | 0 | 0
[6 empty] | 0.1458 | 0.077549 | 39388904648.0534 | 39.3889 | 0.025388 | 14.9234 | 80788.2464 | -0.77086 | 4.0873e-263 | 7.4604e-144 | 1.1191e-26 | 1 | 7.4604e-144 | 0 | 0 | 0 | 0
[6 empty] | 0.146 | 0.077656 | 39442936067.3237 | 39.4429 | 0.025353 | 14.9439 | 81224.5167 | 0.9612 | 2.9561e-265 | 1.2185e-144 | 9.5227e-27 | 1 | 1.2185e-144 | 0 | 0 | 0 | 0
[6 empty] | 0.1462 | 0.077762 | 39496967486.594 | 39.497 | 0.025318 | 14.9643 | 81662.5618 | 0.072331 | 2.0674e-267 | 1.9752e-145 | 8.1013e-27 | 1 | 1.9752e-145 | 0 | 0 | 0 | 0
[6 empty] | 0.1464 | 0.077868 | 39550998905.8643 | 39.551 | 0.025284 | 14.9848 | 82102.3866 | 0.074167 | 1.398e-269 | 3.1782e-146 | 6.8906e-27 | 1 | 3.1782e-146 | 0 | 0 | 0 | 0
[6 empty] | 0.1466 | 0.077975 | 39605030325.1347 | 39.605 | 0.025249 | 15.0053 | 82543.9959 | 0.95829 | 9.1375e-272 | 5.0755e-147 | 5.8595e-27 | 1 | 5.0755e-147 | 0 | 0 | 0 | 0
[6 empty] | 0.1468 | 0.078081 | 39659061744.405 | 39.6591 | 0.025215 | 15.0257 | 82987.3947 | -0.74915 | 5.7724e-274 | 8.0447e-148 | 4.9816e-27 | 1 | 8.0447e-148 | 0 | 0 | 0 | 0
[6 empty] | 0.147 | 0.078187 | 39713093163.6753 | 39.7131 | 0.025181 | 15.0462 | 83432.5877 | -0.98217 | 3.5236e-276 | 1.2655e-148 | 4.2342e-27 | 1 | 1.2655e-148 | 0 | 0 | 0 | 0
[6 empty] | 0.1472 | 0.078294 | 39767124582.9456 | 39.7671 | 0.025146 | 15.0667 | 83879.5798 | -0.76688 | 2.078e-278 | 1.9758e-149 | 3.5982e-27 | 1 | 1.9758e-149 | 0 | 0 | 0 | 0
[6 empty] | 0.1474 | 0.0784 | 39821156002.2159 | 39.8212 | 0.025112 | 15.0872 | 84328.3759 | 0.97009 | 1.1838e-280 | 3.0613e-150 | 3.0571e-27 | 1 | 3.0613e-150 | 0 | 0 | 0 | 0
[6 empty] | 0.1476 | 0.078507 | 39875187421.4862 | 39.8752 | 0.025078 | 15.1076 | 84778.9809 | 0.0316 | 6.5125e-283 | 4.7074e-151 | 2.5968e-27 | 1 | 4.7074e-151 | 0 | 0 | 0 | 0
[6 empty] | 0.1478 | 0.078613 | 39929218840.7565 | 39.9292 | 0.025044 | 15.1281 | 85231.3996 | 0.060934 | 3.4595e-285 | 7.1836e-152 | 2.2053e-27 | 1 | 7.1836e-152 | 0 | 0 | 0 | 0
[6 empty] | 0.148 | 0.078719 | 39983250260.0268 | 39.9833 | 0.02501 | 15.1486 | 85685.6369 | 0.94336 | 1.7741e-287 | 1.0879e-152 | 1.8724e-27 | 1 | 1.0879e-152 | 0 | 0 | 0 | 0
[6 empty] | 0.1482 | 0.078826 | 40037281679.2971 | 40.0373 | 0.024977 | 15.169 | 86141.6978 | -0.64627 | 8.7817e-290 | 1.6349e-153 | 1.5894e-27 | 1 | 1.6349e-153 | 0 | 0 | 0 | 0
[6 empty] | 0.1484 | 0.078932 | 40091313098.5674 | 40.0913 | 0.024943 | 15.1895 | 86599.5871 | -0.9964 | 4.1948e-292 | 2.4382e-154 | 1.3488e-27 | 1 | 2.4382e-154 | 0 | 0 | 0 | 0
[6 empty] | 0.1486 | 0.079038 | 40145344517.8377 | 40.1453 | 0.024909 | 15.21 | 87059.3098 | -0.42214 | 1.9333e-294 | 3.6082e-155 | 1.1445e-27 | 1 | 3.6082e-155 | 0 | 0 | 0 | 0
[6 empty] | 0.1488 | 0.079145 | 40199375937.108 | 40.1994 | 0.024876 | 15.2305 | 87520.8706 | 0.63579 | 8.5951e-297 | 5.2985e-156 | 9.7084e-28 | 1 | 5.2985e-156 | 0 | 0 | 0 | 0
[6 empty] | 0.149 | 0.079251 | 40253407356.3783 | 40.2534 | 0.024843 | 15.2509 | 87984.2747 | 0.7839 | 3.6855e-299 | 7.7206e-157 | 8.2338e-28 | 1 | 7.7206e-157 | 0 | 0 | 0 | 0
[6 empty] | 0.1492 | 0.079358 | 40307438775.6486 | 40.3074 | 0.024809 | 15.2714 | 88449.5269 | 0.93109 | 1.5239e-301 | 1.1163e-157 | 6.9816e-28 | 1 | 1.1163e-157 | 0 | 0 | 0 | 0
[6 empty] | 0.1494 | 0.079464 | 40361470194.919 | 40.3615 | 0.024776 | 15.2919 | 88916.6321 | -0.20386 | 6.0746e-304 | 1.6014e-158 | 5.9185e-28 | 1 | 1.6014e-158 | 0 | 0 | 0 | 0
[6 empty] | 0.1496 | 0.07957 | 40415501614.1893 | 40.4155 | 0.024743 | 15.3123 | 89385.5953 | 0.87777 | 2.3341e-306 | 2.2796e-159 | 5.0162e-28 | 1 | 2.2796e-159 | 0 | 0 | 0 | 0
[6 empty] | 0.1498 | 0.079677 | 40469533033.4596 | 40.4695 | 0.02471 | 15.3328 | 89856.4215 | 0.61188 | 0 | 3.2196e-160 | 4.2505e-28 | 1 | 3.2196e-160 | 0 | 0 | 0 | 0
[6 empty] | 0.15 | 0.079783 | 40523564452.7299 | 40.5236 | 0.024677 | 15.3533 | 90329.1155 | 0.85628 | 0 | 4.5116e-161 | 3.6009e-28 | 1 | 4.5116e-161 | 0 | 0 | 0 | 0
[6 empty] | 0.1502 | 0.07989 | 40577595872.0002 | 40.5776 | 0.024644 | 15.3738 | 90803.6824 | -0.74574 | 0 | 6.2726e-162 | 3.0499e-28 | 1 | 6.2726e-162 | 0 | 0 | 0 | 0
[6 empty] | 0.1504 | 0.079996 | 40631627291.2705 | 40.6316 | 0.024611 | 15.3942 | 91280.127 | -0.93997 | 0 | 8.6522e-163 | 2.5826e-28 | 1 | 8.6522e-163 | 0 | 0 | 0 | 0
[6 empty] | 0.1506 | 0.080102 | 40685658710.5408 | 40.6857 | 0.024579 | 15.4147 | 91758.4545 | -0.89731 | 0 | 1.184e-163 | 2.1864e-28 | 1 | 1.184e-163 | 0 | 0 | 0 | 0
[6 empty] | 0.1508 | 0.080209 | 40739690129.8111 | 40.7397 | 0.024546 | 15.4352 | 92238.6698 | 0.99996 | 0 | 1.6075e-164 | 1.8506e-28 | 1 | 1.6075e-164 | 0 | 0 | 0 | 0
[6 empty] | 0.151 | 0.080315 | 40793721549.0814 | 40.7937 | 0.024514 | 15.4556 | 92720.7778 | -0.11741 | 0 | 2.1652e-165 | 1.5661e-28 | 1 | 2.1652e-165 | 0 | 0 | 0 | 0
[6 empty] | 0.1512 | 0.080421 | 40847752968.3517 | 40.8478 | 0.024481 | 15.4761 | 93204.7836 | 0.082745 | 0 | 2.8932e-166 | 1.325e-28 | 1 | 2.8932e-166 | 0 | 0 | 0 | 0
[6 empty] | 0.1514 | 0.080528 | 40901784387.622 | 40.9018 | 0.024449 | 15.4966 | 93690.6922 | 0.81658 | 0 | 3.835e-167 | 1.1207e-28 | 1 | 3.835e-167 | 0 | 0 | 0 | 0
[6 empty] | 0.1516 | 0.080634 | 40955815806.8923 | 40.9558 | 0.024417 | 15.517 | 94178.5085 | -0.085899 | 0 | 5.0428e-168 | 9.4776e-29 | 1 | 5.0428e-168 | 0 | 0 | 0 | 0
[6 empty] | 0.1518 | 0.080741 | 41009847226.1626 | 41.0098 | 0.024384 | 15.5375 | 94668.2376 | -0.43077 | 0 | 6.578e-169 | 8.0132e-29 | 1 | 6.578e-169 | 0 | 0 | 0 | 0
[6 empty] | 0.152 | 0.080847 | 41063878645.4329 | 41.0639 | 0.024352 | 15.558 | 95159.8845 | 0.89708 | 0 | 8.5115e-170 | 6.7735e-29 | 1 | 8.5115e-170 | 0 | 0 | 0 | 0
[6 empty] | 0.1522 | 0.080953 | 41117910064.7032 | 41.1179 | 0.02432 | 15.5785 | 95653.4543 | -0.99305 | 0 | 1.0925e-170 | 5.7243e-29 | 1 | 1.0925e-170 | 0 | 0 | 0 | 0
[6 empty] | 0.1524 | 0.08106 | 41171941483.9736 | 41.1719 | 0.024288 | 15.5989 | 96148.9519 | -0.54702 | 0 | 1.3909e-171 | 4.8366e-29 | 1 | 1.3909e-171 | 0 | 0 | 0 | 0
[6 empty] | 0.1526 | 0.081166 | 41225972903.2439 | 41.226 | 0.024257 | 15.6194 | 96646.3824 | -0.99776 | 0 | 1.7564e-172 | 4.0856e-29 | 1 | 1.7564e-172 | 0 | 0 | 0 | 0
[6 empty] | 0.1528 | 0.081272 | 41280004322.5142 | 41.28 | 0.024225 | 15.6399 | 97145.7509 | 0.99697 | 0 | 2.2001e-173 | 3.4505e-29 | 1 | 2.2001e-173 | 0 | 0 | 0 | 0
[6 empty] | 0.153 | 0.081379 | 41334035741.7845 | 41.334 | 0.024193 | 15.6603 | 97647.0624 | 0.14909 | 0 | 2.7333e-174 | 2.9135e-29 | 1 | 2.7333e-174 | 0 | 0 | 0 | 0
[6 empty] | 0.1532 | 0.081485 | 41388067161.0548 | 41.3881 | 0.024162 | 15.6808 | 98150.322 | 0.68485 | 0 | 3.3682e-175 | 2.4595e-29 | 1 | 3.3682e-175 | 0 | 0 | 0 | 0
[6 empty] | 0.1534 | 0.081592 | 41442098580.3251 | 41.4421 | 0.02413 | 15.7013 | 98655.5348 | -0.16989 | 0 | 4.1165e-176 | 2.0758e-29 | 1 | 4.1165e-176 | 0 | 0 | 0 | 0
[6 empty] | 0.1536 | 0.081698 | 41496129999.5954 | 41.4961 | 0.024099 | 15.7218 | 99162.7057 | 0.99967 | 0 | 4.9897e-177 | 1.7515e-29 | 1 | 4.9897e-177 | 0 | 0 | 0 | 0
[6 empty] | 0.1538 | 0.081804 | 41550161418.8657 | 41.5502 | 0.024067 | 15.7422 | 99671.84 | 0.98547 | 0 | 5.9985e-178 | 1.4776e-29 | 1 | 5.9985e-178 | 0 | 0 | 0 | 0
[6 empty] | 0.154 | 0.081911 | 41604192838.136 | 41.6042 | 0.024036 | 15.7627 | 100182.9426 | -0.69219 | 0 | 7.1517e-179 | 1.2463e-29 | 1 | 7.1517e-179 | 0 | 0 | 0 | 0
[6 empty] | 0.1542 | 0.082017 | 41658224257.4063 | 41.6582 | 0.024005 | 15.7832 | 100696.0187 | 0.98196 | 0 | 8.4562e-180 | 1.0509e-29 | 1 | 8.4562e-180 | 0 | 0 | 0 | 0
[6 empty] | 0.1544 | 0.082123 | 41712255676.6766 | 41.7123 | 0.023974 | 15.8036 | 101211.0734 | 0.99972 | 0 | 9.9157e-181 | 8.8596e-30 | 1 | 9.9157e-181 | 0 | 0 | 0 | 0
[6 empty] | 0.1546 | 0.08223 | 41766287095.9469 | 41.7663 | 0.023943 | 15.8241 | 101728.1118 | -0.26687 | 0 | 1.1531e-181 | 7.4675e-30 | 1 | 1.1531e-181 | 0 | 0 | 0 | 0
[6 empty] | 0.1548 | 0.082336 | 41820318515.2172 | 41.8203 | 0.023912 | 15.8446 | 102247.139 | 0.80433 | 0 | 1.3297e-182 | 6.2927e-30 | 1 | 1.3297e-182 | 0 | 0 | 0 | 0
[6 empty] | 0.155 | 0.082443 | 41874349934.4875 | 41.8743 | 0.023881 | 15.8651 | 102768.1601 | 0.43613 | 0 | 1.5205e-183 | 5.3016e-30 | 1 | 1.5205e-183 | 0 | 0 | 0 | 0
[6 empty] | 0.1552 | 0.082549 | 41928381353.7579 | 41.9284 | 0.02385 | 15.8855 | 103291.1803 | 0.9225 | 0 | 1.7242e-184 | 4.4655e-30 | 1 | 1.7242e-184 | 0 | 0 | 0 | 0
[6 empty] | 0.1554 | 0.082655 | 41982412773.0282 | 41.9824 | 0.023819 | 15.906 | 103816.2047 | -0.71464 | 0 | 1.9387e-185 | 3.7605e-30 | 1 | 1.9387e-185 | 0 | 0 | 0 | 0
[6 empty] | 0.1556 | 0.082762 | 42036444192.2985 | 42.0364 | 0.023789 | 15.9265 | 104343.2384 | -0.99978 | 0 | 2.1616e-186 | 3.1661e-30 | 1 | 2.1616e-186 | 0 | 0 | 0 | 0
[6 empty] | 0.1558 | 0.082868 | 42090475611.5688 | 42.0905 | 0.023758 | 15.9469 | 104872.2865 | -0.28535 | 0 | 2.3898e-187 | 2.6651e-30 | 1 | 2.3898e-187 | 0 | 0 | 0 | 0
[6 empty] | 0.156 | 0.082974 | 42144507030.8391 | 42.1445 | 0.023728 | 15.9674 | 105403.3543 | 0.15016 | 0 | 2.6197e-188 | 2.2428e-30 | 1 | 2.6197e-188 | 0 | 0 | 0 | 0
[6 empty] | 0.1562 | 0.083081 | 42198538450.1094 | 42.1985 | 0.023698 | 15.9879 | 105936.4469 | 0.90394 | 0 | 2.8474e-189 | 1.8871e-30 | 1 | 2.8474e-189 | 0 | 0 | 0 | 0
[6 empty] | 0.1564 | 0.083187 | 42252569869.3797 | 42.2526 | 0.023667 | 16.0084 | 106471.5695 | 0.077013 | 0 | 3.0685e-190 | 1.5874e-30 | 1 | 3.0685e-190 | 0 | 0 | 0 | 0
[6 empty] | 0.1566 | 0.083294 | 42306601288.65 | 42.3066 | 0.023637 | 16.0288 | 107008.7272 | -0.13137 | 0 | 3.2786e-191 | 1.335e-30 | 1 | 3.2786e-191 | 0 | 0 | 0 | 0
[6 empty] | 0.1568 | 0.0834 | 42360632707.9203 | 42.3606 | 0.023607 | 16.0493 | 107547.9252 | -0.96019 | 0 | 3.4731e-192 | 1.1225e-30 | 1 | 3.4731e-192 | 0 | 0 | 0 | 0
[6 empty] | 0.157 | 0.083506 | 42414664127.1906 | 42.4147 | 0.023577 | 16.0698 | 108089.1687 | -0.3877 | 0 | 3.6477e-193 | 9.4359e-31 | 1 | 3.6477e-193 | 0 | 0 | 0 | 0
[6 empty] | 0.1572 | 0.083613 | 42468695546.4609 | 42.4687 | 0.023547 | 16.0902 | 108632.4628 | 0.5642 | 0 | 3.7981e-194 | 7.9304e-31 | 1 | 3.7981e-194 | 0 | 0 | 0 | 0
[6 empty] | 0.1574 | 0.083719 | 42522726965.7312 | 42.5227 | 0.023517 | 16.1107 | 109177.8129 | 0.95058 | 0 | 3.9206e-195 | 6.6636e-31 | 1 | 3.9206e-195 | 0 | 0 | 0 | 0
[6 empty] | 0.1576 | 0.083825 | 42576758385.0015 | 42.5768 | 0.023487 | 16.1312 | 109725.2241 | 0.89678 | 0 | 4.0122e-196 | 5.5979e-31 | 1 | 4.0122e-196 | 0 | 0 | 0 | 0
[6 empty] | 0.1578 | 0.083932 | 42630789804.2719 | 42.6308 | 0.023457 | 16.1517 | 110274.7016 | -0.98768 | 0 | 4.0703e-197 | 4.7016e-31 | 1 | 4.0703e-197 | 0 | 0 | 0 | 0
[6 empty] | 0.158 | 0.084038 | 42684821223.5422 | 42.6848 | 0.023428 | 16.1721 | 110826.2506 | -0.34915 | 0 | 4.0935e-198 | 3.9479e-31 | 1 | 4.0935e-198 | 0 | 0 | 0 | 0
[6 empty] | 0.1582 | 0.084145 | 42738852642.8125 | 42.7389 | 0.023398 | 16.1926 | 111379.8764 | -0.87341 | 0 | 4.081e-199 | 3.3144e-31 | 1 | 4.081e-199 | 0 | 0 | 0 | 0
[6 empty] | 0.1584 | 0.084251 | 42792884062.0828 | 42.7929 | 0.023368 | 16.2131 | 111935.5842 | 0.65035 | 0 | 4.033e-200 | 2.7818e-31 | 1 | 4.033e-200 | 0 | 0 | 0 | 0
[6 empty] | 0.1586 | 0.084357 | 42846915481.3531 | 42.8469 | 0.023339 | 16.2335 | 112493.3793 | -0.64454 | 0 | 3.9507e-201 | 2.3344e-31 | 1 | 3.9507e-201 | 0 | 0 | 0 | 0
[6 empty] | 0.1588 | 0.084464 | 42900946900.6234 | 42.9009 | 0.02331 | 16.254 | 113053.2668 | -0.016395 | 0 | 3.8362e-202 | 1.9584e-31 | 1 | 3.8362e-202 | 0 | 0 | 0 | 0
[6 empty] | 0.159 | 0.08457 | 42954978319.8937 | 42.955 | 0.02328 | 16.2745 | 113615.252 | 0.36742 | 0 | 3.6923e-203 | 1.6427e-31 | 1 | 3.6923e-203 | 0 | 0 | 0 | 0
[6 empty] | 0.1592 | 0.084677 | 43009009739.164 | 43.009 | 0.023251 | 16.2949 | 114179.3403 | 0.97929 | 0 | 3.5224e-204 | 1.3775e-31 | 1 | 3.5224e-204 | 0 | 0 | 0 | 0
[6 empty] | 0.1594 | 0.084783 | 43063041158.4343 | 43.063 | 0.023222 | 16.3154 | 114745.5368 | 0.87471 | 0 | 3.3307e-205 | 1.1549e-31 | 1 | 3.3307e-205 | 0 | 0 | 0 | 0
[6 empty] | 0.1596 | 0.084889 | 43117072577.7046 | 43.1171 | 0.023193 | 16.3359 | 115313.8468 | -0.98243 | 0 | 3.1215e-206 | 9.6808e-32 | 1 | 3.1215e-206 | 0 | 0 | 0 | 0
[6 empty] | 0.1598 | 0.084996 | 43171103996.9749 | 43.1711 | 0.023164 | 16.3564 | 115884.2757 | -0.40549 | 0 | 2.8995e-207 | 8.1128e-32 | 1 | 2.8995e-207 | 0 | 0 | 0 | 0
[6 empty] | 0.16 | 0.085102 | 43225135416.2452 | 43.2251 | 0.023135 | 16.3768 | 116456.8286 | -0.93226 | 0 | 2.6693e-208 | 6.7973e-32 | 1 | 2.6693e-208 | 0 | 0 | 0 | 0
[6 empty] | 0.1602 | 0.085208 | 43279166835.5155 | 43.2792 | 0.023106 | 16.3973 | 117031.511 | 0.82574 | 0 | 2.4355e-209 | 5.6938e-32 | 1 | 2.4355e-209 | 0 | 0 | 0 | 0
[6 empty] | 0.1604 | 0.085315 | 43333198254.7858 | 43.3332 | 0.023077 | 16.4178 | 117608.328 | -0.26142 | 0 | 2.2022e-210 | 4.7684e-32 | 1 | 2.2022e-210 | 0 | 0 | 0 | 0
[6 empty] | 0.1606 | 0.085421 | 43387229674.0562 | 43.3872 | 0.023048 | 16.4382 | 118187.2851 | 0.5968 | 0 | 1.9735e-211 | 3.9925e-32 | 1 | 1.9735e-211 | 0 | 0 | 0 | 0
[6 empty] | 0.1608 | 0.085528 | 43441261093.3265 | 43.4413 | 0.02302 | 16.4587 | 118768.3874 | -0.52032 | 0 | 1.7526e-212 | 3.3422e-32 | 1 | 1.7526e-212 | 0 | 0 | 0 | 0
[6 empty] | 0.161 | 0.085634 | 43495292512.5968 | 43.4953 | 0.022991 | 16.4792 | 119351.6404 | 0.51071 | 0 | 1.5424e-213 | 2.7971e-32 | 1 | 1.5424e-213 | 0 | 0 | 0 | 0
[6 empty] | 0.1612 | 0.08574 | 43549323931.8671 | 43.5493 | 0.022962 | 16.4997 | 119937.0494 | -0.51132 | 0 | 1.3452e-214 | 2.3405e-32 | 1 | 1.3452e-214 | 0 | 0 | 0 | 0
[6 empty] | 0.1614 | 0.085847 | 43603355351.1374 | 43.6034 | 0.022934 | 16.5201 | 120524.6196 | 0.58845 | 0 | 1.1625e-215 | 1.9579e-32 | 1 | 1.1625e-215 | 0 | 0 | 0 | 0
[6 empty] | 0.1616 | 0.085953 | 43657386770.4077 | 43.6574 | 0.022906 | 16.5406 | 121114.3565 | -0.25069 | 0 | 9.9558e-217 | 1.6375e-32 | 1 | 9.9558e-217 | 0 | 0 | 0 | 0
[6 empty] | 0.1618 | 0.086059 | 43711418189.678 | 43.7114 | 0.022877 | 16.5611 | 121706.2654 | 0.86039 | 0 | 8.4484e-218 | 1.3693e-32 | 1 | 8.4484e-218 | 0 | 0 | 0 | 0
[6 empty] | 0.162 | 0.086166 | 43765449608.9483 | 43.7654 | 0.022849 | 16.5815 | 122300.3516 | -0.97812 | 0 | 7.1038e-219 | 1.1447e-32 | 1 | 7.1038e-219 | 0 | 0 | 0 | 0
[6 empty] | 0.1622 | 0.086272 | 43819481028.2186 | 43.8195 | 0.022821 | 16.602 | 122896.6205 | -0.665 | 0 | 5.9186e-220 | 9.5674e-33 | 1 | 5.9186e-220 | 0 | 0 | 0 | 0
[6 empty] | 0.1624 | 0.086379 | 43873512447.4889 | 43.8735 | 0.022793 | 16.6225 | 123495.0774 | -0.75769 | 0 | 4.8859e-221 | 7.9948e-33 | 1 | 4.8859e-221 | 0 | 0 | 0 | 0
[6 empty] | 0.1626 | 0.086485 | 43927543866.7592 | 43.9275 | 0.022765 | 16.643 | 124095.7278 | 0.25089 | 0 | 3.9963e-222 | 6.6791e-33 | 1 | 3.9963e-222 | 0 | 0 | 0 | 0
[6 empty] | 0.1628 | 0.086591 | 43981575286.0295 | 43.9816 | 0.022737 | 16.6634 | 124698.5769 | 0.55654 | 0 | 3.2386e-223 | 5.5788e-33 | 1 | 3.2386e-223 | 0 | 0 | 0 | 0
[6 empty] | 0.163 | 0.086698 | 44035606705.2998 | 44.0356 | 0.022709 | 16.6839 | 125303.6303 | -0.95726 | 0 | 2.6003e-224 | 4.6587e-33 | 1 | 2.6003e-224 | 0 | 0 | 0 | 0
[6 empty] | 0.1632 | 0.086804 | 44089638124.5701 | 44.0896 | 0.022681 | 16.7044 | 125910.8932 | 0.80082 | 0 | 2.0685e-225 | 3.8895e-33 | 1 | 2.0685e-225 | 0 | 0 | 0 | 0
[6 empty] | 0.1634 | 0.08691 | 44143669543.8405 | 44.1437 | 0.022653 | 16.7248 | 126520.3712 | 0.79542 | 0 | 1.6302e-226 | 3.2465e-33 | 1 | 1.6302e-226 | 0 | 0 | 0 | 0
[6 empty] | 0.1636 | 0.087017 | 44197700963.1108 | 44.1977 | 0.022626 | 16.7453 | 127132.0695 | -0.9661 | 0 | 1.2728e-227 | 2.7093e-33 | 1 | 1.2728e-227 | 0 | 0 | 0 | 0
[6 empty] | 0.1638 | 0.087123 | 44251732382.3811 | 44.2517 | 0.022598 | 16.7658 | 127745.9936 | 0.49547 | 0 | 9.8449e-229 | 2.2604e-33 | 1 | 9.8449e-229 | 0 | 0 | 0 | 0
[6 empty] | 0.164 | 0.08723 | 44305763801.6514 | 44.3058 | 0.02257 | 16.7863 | 128362.149 | 0.11493 | 0 | 7.5437e-230 | 1.8855e-33 | 1 | 7.5437e-230 | 0 | 0 | 0 | 0
[6 empty] | 0.1642 | 0.087336 | 44359795220.9217 | 44.3598 | 0.022543 | 16.8067 | 128980.541 | -0.57854 | 0 | 5.7262e-231 | 1.5725e-33 | 1 | 5.7262e-231 | 0 | 0 | 0 | 0
[6 empty] | 0.1644 | 0.087442 | 44413826640.192 | 44.4138 | 0.022516 | 16.8272 | 129601.1751 | -0.90157 | 0 | 4.3058e-232 | 1.3111e-33 | 1 | 4.3058e-232 | 0 | 0 | 0 | 0
[6 empty] | 0.1646 | 0.087549 | 44467858059.4623 | 44.4679 | 0.022488 | 16.8477 | 130224.0567 | -0.92151 | 0 | 3.2072e-233 | 1.0929e-33 | 1 | 3.2072e-233 | 0 | 0 | 0 | 0
[6 empty] | 0.1648 | 0.087655 | 44521889478.7326 | 44.5219 | 0.022461 | 16.8681 | 130849.1912 | 0.93714 | 0 | 2.3664e-234 | 9.1083e-34 | 1 | 2.3664e-234 | 0 | 0 | 0 | 0
[6 empty] | 0.165 | 0.087761 | 44575920898.0029 | 44.5759 | 0.022434 | 16.8886 | 131476.5842 | 0.84239 | 0 | 1.7295e-235 | 7.5893e-34 | 1 | 1.7295e-235 | 0 | 0 | 0 | 0
[6 empty] | 0.1652 | 0.087868 | 44629952317.2732 | 44.63 | 0.022406 | 16.9091 | 132106.241 | 0.71845 | 0 | 1.252e-236 | 6.3222e-34 | 1 | 1.252e-236 | 0 | 0 | 0 | 0
[6 empty] | 0.1654 | 0.087974 | 44683983736.5435 | 44.684 | 0.022379 | 16.9296 | 132738.1672 | -0.32925 | 0 | 8.9771e-238 | 5.2654e-34 | 1 | 8.9771e-238 | 0 | 0 | 0 | 0
[6 empty] | 0.1656 | 0.088081 | 44738015155.8138 | 44.738 | 0.022352 | 16.95 | 133372.3682 | -0.6715 | 0 | 6.3756e-239 | 4.3844e-34 | 1 | 6.3756e-239 | 0 | 0 | 0 | 0
[6 empty] | 0.1658 | 0.088187 | 44792046575.0841 | 44.792 | 0.022325 | 16.9705 | 134008.8495 | 0.91002 | 0 | 4.4847e-240 | 3.6499e-34 | 1 | 4.4847e-240 | 0 | 0 | 0 | 0
[6 empty] | 0.166 | 0.088293 | 44846077994.3544 | 44.8461 | 0.022298 | 16.991 | 134647.6166 | -0.82742 | 0 | 3.1244e-241 | 3.0378e-34 | 1 | 3.1244e-241 | 0 | 0 | 0 | 0
[6 empty] | 0.1662 | 0.0884 | 44900109413.6248 | 44.9001 | 0.022272 | 17.0114 | 135288.6749 | -0.7181 | 0 | 2.1558e-242 | 2.5278e-34 | 1 | 2.1558e-242 | 0 | 0 | 0 | 0
[6 empty] | 0.1664 | 0.088506 | 44954140832.8951 | 44.9541 | 0.022245 | 17.0319 | 135932.03 | 0.99517 | 0 | 1.4731e-243 | 2.103e-34 | 1 | 1.4731e-243 | 0 | 0 | 0 | 0
[6 empty] | 0.1666 | 0.088612 | 45008172252.1654 | 45.0082 | 0.022218 | 17.0524 | 136577.6874 | 0.15774 | 0 | 9.9696e-245 | 1.7492e-34 | 1 | 9.9696e-245 | 0 | 0 | 0 | 0
[6 empty] | 0.1668 | 0.088719 | 45062203671.4357 | 45.0622 | 0.022192 | 17.0728 | 137225.6525 | 0.81658 | 0 | 6.6818e-246 | 1.4545e-34 | 1 | 6.6818e-246 | 0 | 0 | 0 | 0
[6 empty] | 0.167 | 0.088825 | 45116235090.706 | 45.1162 | 0.022165 | 17.0933 | 137875.931 | -0.79815 | 0 | 4.4348e-247 | 1.2093e-34 | 1 | 4.4348e-247 | 0 | 0 | 0 | 0
[6 empty] | 0.1672 | 0.088932 | 45170266509.9763 | 45.1703 | 0.022138 | 17.1138 | 138528.5282 | -0.070157 | 0 | 2.9149e-248 | 1.0052e-34 | 1 | 2.9149e-248 | 0 | 0 | 0 | 0
[6 empty] | 0.1674 | 0.089038 | 45224297929.2466 | 45.2243 | 0.022112 | 17.1343 | 139183.4498 | -0.99954 | 0 | 1.8972e-249 | 8.353e-35 | 1 | 1.8972e-249 | 0 | 0 | 0 | 0
[6 empty] | 0.1676 | 0.089144 | 45278329348.5169 | 45.2783 | 0.022086 | 17.1547 | 139840.7014 | 0.80896 | 0 | 1.2228e-250 | 6.9399e-35 | 1 | 1.2228e-250 | 0 | 0 | 0 | 0
[6 empty] | 0.1678 | 0.089251 | 45332360767.7872 | 45.3324 | 0.022059 | 17.1752 | 140500.2883 | 0.88657 | 0 | 7.8038e-252 | 5.7646e-35 | 1 | 7.8038e-252 | 0 | 0 | 0 | 0
[6 empty] | 0.168 | 0.089357 | 45386392187.0575 | 45.3864 | 0.022033 | 17.1957 | 141162.2163 | -0.89284 | 0 | 4.9316e-253 | 4.7873e-35 | 1 | 4.9316e-253 | 0 | 0 | 0 | 0
[6 empty] | 0.1682 | 0.089463 | 45440423606.3278 | 45.4404 | 0.022007 | 17.2161 | 141826.4907 | 0.59683 | 0 | 3.0859e-254 | 3.9748e-35 | 1 | 3.0859e-254 | 0 | 0 | 0 | 0
[6 empty] | 0.1684 | 0.08957 | 45494455025.5981 | 45.4945 | 0.021981 | 17.2366 | 142493.1174 | 0.030575 | 0 | 1.9119e-255 | 3.2994e-35 | 1 | 1.9119e-255 | 0 | 0 | 0 | 0
[6 empty] | 0.1686 | 0.089676 | 45548486444.8684 | 45.5485 | 0.021955 | 17.2571 | 143162.1016 | -0.20408 | 0 | 1.1729e-256 | 2.7382e-35 | 1 | 1.1729e-256 | 0 | 0 | 0 | 0
[6 empty] | 0.1688 | 0.089783 | 45602517864.1387 | 45.6025 | 0.021929 | 17.2776 | 143833.4492 | -0.91633 | 0 | 7.1237e-258 | 2.272e-35 | 1 | 7.1237e-258 | 0 | 0 | 0 | 0
[6 empty] | 0.169 | 0.089889 | 45656549283.4091 | 45.6565 | 0.021903 | 17.298 | 144507.1656 | 0.25395 | 0 | 4.2838e-259 | 1.8847e-35 | 1 | 4.2838e-259 | 0 | 0 | 0 | 0
[6 empty] | 0.1692 | 0.089995 | 45710580702.6794 | 45.7106 | 0.021877 | 17.3185 | 145183.2564 | -0.78653 | 0 | 2.5504e-260 | 1.5631e-35 | 1 | 2.5504e-260 | 0 | 0 | 0 | 0
[6 empty] | 0.1694 | 0.090102 | 45764612121.9497 | 45.7646 | 0.021851 | 17.339 | 145861.7273 | -0.71177 | 0 | 1.5033e-261 | 1.2961e-35 | 1 | 1.5033e-261 | 0 | 0 | 0 | 0
[6 empty] | 0.1696 | 0.090208 | 45818643541.22 | 45.8186 | 0.021825 | 17.3594 | 146542.5838 | -0.077012 | 0 | 8.7722e-263 | 1.0745e-35 | 1 | 8.7722e-263 | 0 | 0 | 0 | 0
[6 empty] | 0.1698 | 0.090315 | 45872674960.4903 | 45.8727 | 0.021799 | 17.3799 | 147225.8315 | -0.99211 | 0 | 5.0676e-264 | 8.9052e-36 | 1 | 5.0676e-264 | 0 | 0 | 0 | 0
[6 empty] | 0.17 | 0.090421 | 45926706379.7606 | 45.9267 | 0.021774 | 17.4004 | 147911.4761 | -0.79507 | 0 | 2.8981e-265 | 7.3791e-36 | 1 | 2.8981e-265 | 0 | 0 | 0 | 0
[6 empty] | 0.1702 | 0.090527 | 45980737799.0309 | 45.9807 | 0.021748 | 17.4209 | 148599.5232 | 0.77125 | 0 | 1.6407e-266 | 6.1131e-36 | 1 | 1.6407e-266 | 0 | 0 | 0 | 0
[6 empty] | 0.1704 | 0.090634 | 46034769218.3012 | 46.0348 | 0.021723 | 17.4413 | 149289.9785 | 0.99999 | 0 | 9.1949e-268 | 5.0633e-36 | 1 | 9.1949e-268 | 0 | 0 | 0 | 0
[6 empty] | 0.1706 | 0.09074 | 46088800637.5715 | 46.0888 | 0.021697 | 17.4618 | 149982.8475 | -0.14217 | 0 | 5.1009e-269 | 4.1928e-36 | 1 | 5.1009e-269 | 0 | 0 | 0 | 0
[6 empty] | 0.1708 | 0.090846 | 46142832056.8418 | 46.1428 | 0.021672 | 17.4823 | 150678.1359 | 0.90824 | 0 | 2.801e-270 | 3.4712e-36 | 1 | 2.801e-270 | 0 | 0 | 0 | 0
[6 empty] | 0.171 | 0.090953 | 46196863476.1121 | 46.1969 | 0.021646 | 17.5027 | 151375.8493 | 0.98849 | 0 | 1.5225e-271 | 2.8731e-36 | 1 | 1.5225e-271 | 0 | 0 | 0 | 0
[6 empty] | 0.1712 | 0.091059 | 46250894895.3824 | 46.2509 | 0.021621 | 17.5232 | 152075.9935 | -0.83489 | 0 | 8.1913e-273 | 2.3776e-36 | 1 | 8.1913e-273 | 0 | 0 | 0 | 0
[6 empty] | 0.1714 | 0.091166 | 46304926314.6527 | 46.3049 | 0.021596 | 17.5437 | 152778.574 | 0.14775 | 0 | 4.362e-274 | 1.9671e-36 | 1 | 4.362e-274 | 0 | 0 | 0 | 0
[6 empty] | 0.1716 | 0.091272 | 46358957733.923 | 46.359 | 0.021571 | 17.5642 | 153483.5966 | -0.91581 | 0 | 2.2991e-275 | 1.6271e-36 | 1 | 2.2991e-275 | 0 | 0 | 0 | 0
[6 empty] | 0.1718 | 0.091378 | 46412989153.1934 | 46.413 | 0.021546 | 17.5846 | 154191.0669 | 0.98033 | 0 | 1.1994e-276 | 1.3456e-36 | 1 | 1.1994e-276 | 0 | 0 | 0 | 0
[6 empty] | 0.172 | 0.091485 | 46467020572.4637 | 46.467 | 0.021521 | 17.6051 | 154900.9906 | 0.99251 | 0 | 6.1924e-278 | 1.1125e-36 | 1 | 6.1924e-278 | 0 | 0 | 0 | 0
[6 empty] | 0.1722 | 0.091591 | 46521051991.734 | 46.5211 | 0.021496 | 17.6256 | 155613.3734 | -0.8043 | 0 | 3.1642e-279 | 9.1961e-37 | 1 | 3.1642e-279 | 0 | 0 | 0 | 0
[6 empty] | 0.1724 | 0.091697 | 46575083411.0043 | 46.5751 | 0.021471 | 17.646 | 156328.2209 | 0.48033 | 0 | 1.6002e-280 | 7.5999e-37 | 1 | 1.6002e-280 | 0 | 0 | 0 | 0
[6 empty] | 0.1726 | 0.091804 | 46629114830.2746 | 46.6291 | 0.021446 | 17.6665 | 157045.539 | -0.50892 | 0 | 8.0085e-282 | 6.2794e-37 | 1 | 8.0085e-282 | 0 | 0 | 0 | 0
[6 empty] | 0.1728 | 0.09191 | 46683146249.5449 | 46.6831 | 0.021421 | 17.687 | 157765.3333 | 0.7855 | 0 | 3.9665e-283 | 5.1872e-37 | 1 | 3.9665e-283 | 0 | 0 | 0 | 0
[6 empty] | 0.173 | 0.092017 | 46737177668.8152 | 46.7372 | 0.021396 | 17.7075 | 158487.6095 | 0.57566 | 0 | 1.9441e-284 | 4.284e-37 | 1 | 1.9441e-284 | 0 | 0 | 0 | 0
[6 empty] | 0.1732 | 0.092123 | 46791209088.0855 | 46.7912 | 0.021372 | 17.7279 | 159212.3734 | 0.32464 | 0 | 9.4294e-286 | 3.5373e-37 | 1 | 9.4294e-286 | 0 | 0 | 0 | 0
[6 empty] | 0.1734 | 0.092229 | 46845240507.3558 | 46.8452 | 0.021347 | 17.7484 | 159939.6306 | 0.93866 | 0 | 4.5257e-287 | 2.9201e-37 | 1 | 4.5257e-287 | 0 | 0 | 0 | 0
[6 empty] | 0.1736 | 0.092336 | 46899271926.6261 | 46.8993 | 0.021322 | 17.7689 | 160669.387 | 0.85002 | 0 | 2.1494e-288 | 2.41e-37 | 1 | 2.1494e-288 | 0 | 0 | 0 | 0
[6 empty] | 0.1738 | 0.092442 | 46953303345.8964 | 46.9533 | 0.021298 | 17.7893 | 161401.6482 | -0.67863 | 0 | 1.0101e-289 | 1.9886e-37 | 1 | 1.0101e-289 | 0 | 0 | 0 | 0
[6 empty] | 0.174 | 0.092548 | 47007334765.1667 | 47.0073 | 0.021273 | 17.8098 | 162136.42 | -0.89425 | 0 | 4.6971e-291 | 1.6406e-37 | 1 | 4.6971e-291 | 0 | 0 | 0 | 0
[6 empty] | 0.1742 | 0.092655 | 47061366184.437 | 47.0614 | 0.021249 | 17.8303 | 162873.7082 | 0.86679 | 0 | 2.1611e-292 | 1.3531e-37 | 1 | 2.1611e-292 | 0 | 0 | 0 | 0
[6 empty] | 0.1744 | 0.092761 | 47115397603.7074 | 47.1154 | 0.021224 | 17.8507 | 163613.5185 | -0.52849 | 0 | 9.8383e-294 | 1.1158e-37 | 1 | 9.8383e-294 | 0 | 0 | 0 | 0
[6 empty] | 0.1746 | 0.092868 | 47169429022.9777 | 47.1694 | 0.0212 | 17.8712 | 164355.8567 | 0.35743 | 0 | 4.4314e-295 | 9.1987e-38 | 1 | 4.4314e-295 | 0 | 0 | 0 | 0
[6 empty] | 0.1748 | 0.092974 | 47223460442.248 | 47.2235 | 0.021176 | 17.8917 | 165100.7286 | -0.62874 | 0 | 1.9748e-296 | 7.5819e-38 | 1 | 1.9748e-296 | 0 | 0 | 0 | 0
[6 empty] | 0.175 | 0.09308 | 47277491861.5183 | 47.2775 | 0.021152 | 17.9122 | 165848.1399 | -0.38224 | 0 | 8.7068e-298 | 6.2479e-38 | 1 | 8.7068e-298 | 0 | 0 | 0 | 0
[6 empty] | 0.1752 | 0.093187 | 47331523280.7886 | 47.3315 | 0.021128 | 17.9326 | 166598.0965 | -0.47226 | 0 | 3.7979e-299 | 5.1475e-38 | 1 | 3.7979e-299 | 0 | 0 | 0 | 0
[6 empty] | 0.1754 | 0.093293 | 47385554700.0589 | 47.3856 | 0.021103 | 17.9531 | 167350.6041 | -0.92274 | 0 | 1.6389e-300 | 4.2399e-38 | 1 | 1.6389e-300 | 0 | 0 | 0 | 0
[6 empty] | 0.1756 | 0.093399 | 47439586119.3292 | 47.4396 | 0.021079 | 17.9736 | 168105.6686 | -0.77344 | 0 | 6.9967e-302 | 3.4916e-38 | 1 | 6.9967e-302 | 0 | 0 | 0 | 0
[6 empty] | 0.1758 | 0.093506 | 47493617538.5995 | 47.4936 | 0.021055 | 17.994 | 168863.2957 | 0.37181 | 0 | 2.9549e-303 | 2.8747e-38 | 1 | 2.9549e-303 | 0 | 0 | 0 | 0
[6 empty] | 0.176 | 0.093612 | 47547648957.8698 | 47.5476 | 0.021032 | 18.0145 | 169623.4913 | 0.43568 | 0 | 1.2345e-304 | 2.3663e-38 | 1 | 1.2345e-304 | 0 | 0 | 0 | 0
[6 empty] | 0.1762 | 0.093719 | 47601680377.1401 | 47.6017 | 0.021008 | 18.035 | 170386.2612 | -0.88567 | 0 | 5.1019e-306 | 1.9474e-38 | 1 | 5.1019e-306 | 0 | 0 | 0 | 0
[6 empty] | 0.1764 | 0.093825 | 47655711796.4104 | 47.6557 | 0.020984 | 18.0555 | 171151.6112 | -0.75461 | 0 | 2.0857e-307 | 1.6023e-38 | 1 | 2.0857e-307 | 0 | 0 | 0 | 0
[6 empty] | 0.1766 | 0.093931 | 47709743215.6807 | 47.7097 | 0.02096 | 18.0759 | 171919.5471 | -0.78282 | 0 | 0 | 1.318e-38 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1768 | 0.094038 | 47763774634.951 | 47.7638 | 0.020936 | 18.0964 | 172690.0749 | 0.06158 | 0 | 0 | 1.084e-38 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.177 | 0.094144 | 47817806054.2213 | 47.8178 | 0.020913 | 18.1169 | 173463.2003 | -0.22992 | 0 | 0 | 8.9127e-39 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1772 | 0.09425 | 47871837473.4917 | 47.8718 | 0.020889 | 18.1373 | 174238.9292 | -0.012471 | 0 | 0 | 7.3267e-39 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1774 | 0.094357 | 47925868892.762 | 47.9259 | 0.020866 | 18.1578 | 175017.2675 | -0.70977 | 0 | 0 | 6.0216e-39 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1776 | 0.094463 | 47979900312.0323 | 47.9799 | 0.020842 | 18.1783 | 175798.221 | 0.86711 | 0 | 0 | 4.9479e-39 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1778 | 0.09457 | 48033931731.3026 | 48.0339 | 0.020819 | 18.1988 | 176581.7957 | -0.69911 | 0 | 0 | 4.0647e-39 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.178 | 0.094676 | 48087963150.5729 | 48.088 | 0.020795 | 18.2192 | 177367.9973 | 0.02931 | 0 | 0 | 3.3384e-39 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1782 | 0.094782 | 48141994569.8432 | 48.142 | 0.020772 | 18.2397 | 178156.8318 | -0.31843 | 0 | 0 | 2.7413e-39 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1784 | 0.094889 | 48196025989.1135 | 48.196 | 0.020749 | 18.2602 | 178948.3051 | -0.11576 | 0 | 0 | 2.2505e-39 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1786 | 0.094995 | 48250057408.3838 | 48.2501 | 0.020725 | 18.2806 | 179742.4231 | -0.55549 | 0 | 0 | 1.8472e-39 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1788 | 0.095102 | 48304088827.6541 | 48.3041 | 0.020702 | 18.3011 | 180539.1916 | -0.97716 | 0 | 0 | 1.5158e-39 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.179 | 0.095208 | 48358120246.9244 | 48.3581 | 0.020679 | 18.3216 | 181338.6167 | -0.31872 | 0 | 0 | 1.2436e-39 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1792 | 0.095314 | 48412151666.1947 | 48.4122 | 0.020656 | 18.3421 | 182140.7042 | -0.61074 | 0 | 0 | 1.02e-39 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1794 | 0.095421 | 48466183085.465 | 48.4662 | 0.020633 | 18.3625 | 182945.4601 | -0.91884 | 0 | 0 | 8.3648e-40 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1796 | 0.095527 | 48520214504.7353 | 48.5202 | 0.02061 | 18.383 | 183752.8902 | 0.93418 | 0 | 0 | 6.8581e-40 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1798 | 0.095633 | 48574245924.0057 | 48.5742 | 0.020587 | 18.4035 | 184563.0006 | 0.70714 | 0 | 0 | 5.6215e-40 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.18 | 0.09574 | 48628277343.276 | 48.6283 | 0.020564 | 18.4239 | 185375.7972 | 0.090368 | 0 | 0 | 4.6069e-40 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1802 | 0.095846 | 48682308762.5463 | 48.6823 | 0.020541 | 18.4444 | 186191.2859 | 0.98802 | 0 | 0 | 3.7745e-40 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1804 | 0.095953 | 48736340181.8166 | 48.7363 | 0.020519 | 18.4649 | 187009.4727 | 0.043153 | 0 | 0 | 3.0919e-40 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1806 | 0.096059 | 48790371601.0869 | 48.7904 | 0.020496 | 18.4854 | 187830.3635 | 0.77834 | 0 | 0 | 2.5321e-40 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1808 | 0.096165 | 48844403020.3572 | 48.8444 | 0.020473 | 18.5058 | 188653.9644 | 0.9847 | 0 | 0 | 2.0733e-40 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.181 | 0.096272 | 48898434439.6275 | 48.8984 | 0.020451 | 18.5263 | 189480.2814 | -0.99529 | 0 | 0 | 1.6972e-40 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1812 | 0.096378 | 48952465858.8978 | 48.9525 | 0.020428 | 18.5468 | 190309.3203 | -0.90535 | 0 | 0 | 1.389e-40 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1814 | 0.096484 | 49006497278.1681 | 49.0065 | 0.020405 | 18.5672 | 191141.0873 | 0.36824 | 0 | 0 | 1.1365e-40 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1816 | 0.096591 | 49060528697.4384 | 49.0605 | 0.020383 | 18.5877 | 191975.5883 | -0.70722 | 0 | 0 | 9.2976e-41 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1818 | 0.096697 | 49114560116.7087 | 49.1146 | 0.020361 | 18.6082 | 192812.8292 | 0.7116 | 0 | 0 | 7.6043e-41 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.182 | 0.096804 | 49168591535.979 | 49.1686 | 0.020338 | 18.6286 | 193652.8163 | -0.91999 | 0 | 0 | 6.218e-41 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1822 | 0.09691 | 49222622955.2493 | 49.2226 | 0.020316 | 18.6491 | 194495.5553 | -0.36702 | 0 | 0 | 5.0833e-41 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1824 | 0.097016 | 49276654374.5196 | 49.2767 | 0.020294 | 18.6696 | 195341.0525 | -0.032912 | 0 | 0 | 4.1548e-41 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1826 | 0.097123 | 49330685793.7899 | 49.3307 | 0.020271 | 18.6901 | 196189.3137 | -0.064122 | 0 | 0 | 3.3951e-41 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1828 | 0.097229 | 49384717213.0603 | 49.3847 | 0.020249 | 18.7105 | 197040.3452 | -0.27253 | 0 | 0 | 2.7737e-41 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.183 | 0.097335 | 49438748632.3306 | 49.4387 | 0.020227 | 18.731 | 197894.1528 | -0.83139 | 0 | 0 | 2.2656e-41 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1832 | 0.097442 | 49492780051.6009 | 49.4928 | 0.020205 | 18.7515 | 198750.7427 | 0.88897 | 0 | 0 | 1.8501e-41 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1834 | 0.097548 | 49546811470.8712 | 49.5468 | 0.020183 | 18.7719 | 199610.1209 | -0.3175 | 0 | 0 | 1.5105e-41 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1836 | 0.097655 | 49600842890.1415 | 49.6008 | 0.020161 | 18.7924 | 200472.2935 | 0.86897 | 0 | 0 | 1.2329e-41 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1838 | 0.097761 | 49654874309.4118 | 49.6549 | 0.020139 | 18.8129 | 201337.2665 | -0.86911 | 0 | 0 | 1.0061e-41 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.184 | 0.097867 | 49708905728.6821 | 49.7089 | 0.020117 | 18.8334 | 202205.0461 | -0.34607 | 0 | 0 | 8.209e-42 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1842 | 0.097974 | 49762937147.9524 | 49.7629 | 0.020095 | 18.8538 | 203075.6383 | -0.017682 | 0 | 0 | 6.6962e-42 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1844 | 0.09808 | 49816968567.2227 | 49.817 | 0.020073 | 18.8743 | 203949.0493 | -0.065832 | 0 | 0 | 5.4609e-42 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1846 | 0.098186 | 49870999986.493 | 49.871 | 0.020052 | 18.8948 | 204825.2851 | -0.20131 | 0 | 0 | 4.4526e-42 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1848 | 0.098293 | 49925031405.7633 | 49.925 | 0.02003 | 18.9152 | 205704.3518 | -0.70464 | 0 | 0 | 3.6296e-42 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.185 | 0.098399 | 49979062825.0336 | 49.9791 | 0.020008 | 18.9357 | 206586.2556 | 0.9955 | 0 | 0 | 2.9581e-42 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1852 | 0.098506 | 50033094244.3039 | 50.0331 | 0.019987 | 18.9562 | 207471.0025 | 0.28949 | 0 | 0 | 2.4103e-42 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1854 | 0.098612 | 50087125663.5743 | 50.0871 | 0.019965 | 18.9767 | 208358.5987 | 0.92491 | 0 | 0 | 1.9635e-42 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1856 | 0.098718 | 50141157082.8446 | 50.1412 | 0.019944 | 18.9971 | 209249.0504 | 0.1988 | 0 | 0 | 1.5991e-42 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1858 | 0.098825 | 50195188502.1149 | 50.1952 | 0.019922 | 19.0176 | 210142.3636 | 0.96383 | 0 | 0 | 1.3021e-42 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.186 | 0.098931 | 50249219921.3852 | 50.2492 | 0.019901 | 19.0381 | 211038.5445 | -0.8487 | 0 | 0 | 1.06e-42 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1862 | 0.099037 | 50303251340.6555 | 50.3033 | 0.019879 | 19.0585 | 211937.5993 | -0.43876 | 0 | 0 | 8.6277e-43 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1864 | 0.099144 | 50357282759.9258 | 50.3573 | 0.019858 | 19.079 | 212839.5342 | 0.15583 | 0 | 0 | 7.0206e-43 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1866 | 0.09925 | 50411314179.1961 | 50.4113 | 0.019837 | 19.0995 | 213744.3552 | 0.11387 | 0 | 0 | 5.7116e-43 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1868 | 0.099357 | 50465345598.4664 | 50.4653 | 0.019816 | 19.12 | 214652.0686 | -0.3155 | 0 | 0 | 4.6456e-43 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.187 | 0.099463 | 50519377017.7367 | 50.5194 | 0.019794 | 19.1404 | 215562.6806 | -0.69678 | 0 | 0 | 3.7777e-43 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1872 | 0.099569 | 50573408437.007 | 50.5734 | 0.019773 | 19.1609 | 216476.1973 | 0.99361 | 0 | 0 | 3.0713e-43 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1874 | 0.099676 | 50627439856.2773 | 50.6274 | 0.019752 | 19.1814 | 217392.6249 | 0.69362 | 0 | 0 | 2.4964e-43 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1876 | 0.099782 | 50681471275.5476 | 50.6815 | 0.019731 | 19.2018 | 218311.9697 | 0.36662 | 0 | 0 | 2.0287e-43 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1878 | 0.099888 | 50735502694.8179 | 50.7355 | 0.01971 | 19.2223 | 219234.2378 | 0.98646 | 0 | 0 | 1.6483e-43 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.188 | 0.099995 | 50789534114.0882 | 50.7895 | 0.019689 | 19.2428 | 220159.4354 | 0.16534 | 0 | 0 | 1.3389e-43 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1882 | 0.1001 | 50843565533.3586 | 50.8436 | 0.019668 | 19.2632 | 221087.5688 | 0.93113 | 0 | 0 | 1.0873e-43 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1884 | 0.10021 | 50897596952.6289 | 50.8976 | 0.019647 | 19.2837 | 222018.6443 | 0.70336 | 0 | 0 | 8.8281e-44 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1886 | 0.10031 | 50951628371.8992 | 50.9516 | 0.019626 | 19.3042 | 222952.6679 | 0.18938 | 0 | 0 | 7.1662e-44 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1888 | 0.10042 | 51005659791.1695 | 51.0057 | 0.019606 | 19.3247 | 223889.646 | 0.82717 | 0 | 0 | 5.8159e-44 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.189 | 0.10053 | 51059691210.4398 | 51.0597 | 0.019585 | 19.3451 | 224829.5849 | -0.99998 | 0 | 0 | 4.7189e-44 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1892 | 0.10063 | 51113722629.7101 | 51.1137 | 0.019564 | 19.3656 | 225772.4907 | -0.90733 | 0 | 0 | 3.828e-44 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1894 | 0.10074 | 51167754048.9804 | 51.1678 | 0.019544 | 19.3861 | 226718.3697 | 0.76897 | 0 | 0 | 3.1046e-44 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1896 | 0.10085 | 51221785468.2507 | 51.2218 | 0.019523 | 19.4065 | 227667.2283 | 0.70304 | 0 | 0 | 2.5174e-44 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1898 | 0.10095 | 51275816887.521 | 51.2758 | 0.019502 | 19.427 | 228619.0726 | -0.74326 | 0 | 0 | 2.0408e-44 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.19 | 0.10106 | 51329848306.7913 | 51.3298 | 0.019482 | 19.4475 | 229573.909 | -0.86529 | 0 | 0 | 1.654e-44 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1902 | 0.10117 | 51383879726.0616 | 51.3839 | 0.019461 | 19.468 | 230531.7437 | 0.9849 | 0 | 0 | 1.3403e-44 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1904 | 0.10127 | 51437911145.3319 | 51.4379 | 0.019441 | 19.4884 | 231492.583 | 0.9511 | 0 | 0 | 1.0858e-44 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1906 | 0.10138 | 51491942564.6022 | 51.4919 | 0.019421 | 19.5089 | 232456.4333 | -0.59566 | 0 | 0 | 8.7945e-45 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1908 | 0.10148 | 51545973983.8726 | 51.546 | 0.0194 | 19.5294 | 233423.3007 | 0.10476 | 0 | 0 | 7.1216e-45 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.191 | 0.10159 | 51600005403.1429 | 51.6 | 0.01938 | 19.5498 | 234393.1918 | -0.82265 | 0 | 0 | 5.7656e-45 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1912 | 0.1017 | 51654036822.4132 | 51.654 | 0.01936 | 19.5703 | 235366.1127 | -0.93302 | 0 | 0 | 4.6668e-45 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1914 | 0.1018 | 51708068241.6835 | 51.7081 | 0.019339 | 19.5908 | 236342.0698 | 0.12418 | 0 | 0 | 3.7765e-45 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1916 | 0.10191 | 51762099660.9538 | 51.7621 | 0.019319 | 19.6113 | 237321.0694 | -0.86881 | 0 | 0 | 3.0554e-45 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1918 | 0.10202 | 51816131080.2241 | 51.8161 | 0.019299 | 19.6317 | 238303.1179 | 0.73032 | 0 | 0 | 2.4715e-45 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.192 | 0.10212 | 51870162499.4944 | 51.8702 | 0.019279 | 19.6522 | 239288.2216 | -0.51209 | 0 | 0 | 1.9987e-45 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1922 | 0.10223 | 51924193918.7647 | 51.9242 | 0.019259 | 19.6727 | 240276.3868 | 0.91979 | 0 | 0 | 1.616e-45 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1924 | 0.10234 | 51978225338.035 | 51.9782 | 0.019239 | 19.6931 | 241267.62 | -0.3359 | 0 | 0 | 1.3063e-45 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1926 | 0.10244 | 52032256757.3053 | 52.0323 | 0.019219 | 19.7136 | 242261.9274 | 0.93965 | 0 | 0 | 1.0557e-45 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1928 | 0.10255 | 52086288176.5756 | 52.0863 | 0.019199 | 19.7341 | 243259.3156 | -0.40477 | 0 | 0 | 8.5296e-46 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.193 | 0.10265 | 52140319595.8459 | 52.1403 | 0.019179 | 19.7546 | 244259.7907 | 0.85831 | 0 | 0 | 6.8903e-46 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1932 | 0.10276 | 52194351015.1162 | 52.1944 | 0.019159 | 19.775 | 245263.3593 | -0.65116 | 0 | 0 | 5.5648e-46 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1934 | 0.10287 | 52248382434.3865 | 52.2484 | 0.019139 | 19.7955 | 246270.0277 | 0.60489 | 0 | 0 | 4.4933e-46 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1936 | 0.10297 | 52302413853.6568 | 52.3024 | 0.01912 | 19.816 | 247279.8023 | -0.9202 | 0 | 0 | 3.6273e-46 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1938 | 0.10308 | 52356445272.9272 | 52.3564 | 0.0191 | 19.8364 | 248292.6895 | 0.12546 | 0 | 0 | 2.9275e-46 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.194 | 0.10319 | 52410476692.1975 | 52.4105 | 0.01908 | 19.8569 | 249308.6958 | -0.98504 | 0 | 0 | 2.3623e-46 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1942 | 0.10329 | 52464508111.4678 | 52.4645 | 0.019061 | 19.8774 | 250327.8275 | -0.46911 | 0 | 0 | 1.9057e-46 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1944 | 0.1034 | 52518539530.7381 | 52.5185 | 0.019041 | 19.8979 | 251350.0911 | -0.68724 | 0 | 0 | 1.5371e-46 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1946 | 0.10351 | 52572570950.0084 | 52.5726 | 0.019021 | 19.9183 | 252375.4929 | -0.90915 | 0 | 0 | 1.2395e-46 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1948 | 0.10361 | 52626602369.2787 | 52.6266 | 0.019002 | 19.9388 | 253404.0396 | -0.10441 | 0 | 0 | 9.9924e-47 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.195 | 0.10372 | 52680633788.549 | 52.6806 | 0.018982 | 19.9593 | 254435.7374 | -0.97787 | 0 | 0 | 8.0541e-47 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1952 | 0.10382 | 52734665207.8193 | 52.7347 | 0.018963 | 19.9797 | 255470.5929 | 0.48825 | 0 | 0 | 6.4903e-47 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1954 | 0.10393 | 52788696627.0896 | 52.7887 | 0.018943 | 20.0002 | 256508.6124 | -0.70605 | 0 | 0 | 5.229e-47 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1956 | 0.10404 | 52842728046.3599 | 52.8427 | 0.018924 | 20.0207 | 257549.8025 | 0.85972 | 0 | 0 | 4.2119e-47 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1958 | 0.10414 | 52896759465.6302 | 52.8968 | 0.018905 | 20.0411 | 258594.1697 | -0.31805 | 0 | 0 | 3.3919e-47 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.196 | 0.10425 | 52950790884.9005 | 52.9508 | 0.018885 | 20.0616 | 259641.7204 | 0.98829 | 0 | 0 | 2.7309e-47 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1962 | 0.10436 | 53004822304.1708 | 53.0048 | 0.018866 | 20.0821 | 260692.4611 | -0.031127 | 0 | 0 | 2.1982e-47 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1964 | 0.10446 | 53058853723.4412 | 53.0589 | 0.018847 | 20.1026 | 261746.3982 | 0.99935 | 0 | 0 | 1.7691e-47 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1966 | 0.10457 | 53112885142.7115 | 53.1129 | 0.018828 | 20.123 | 262803.5384 | 0.041795 | 0 | 0 | 1.4234e-47 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1968 | 0.10468 | 53166916561.9818 | 53.1669 | 0.018809 | 20.1435 | 263863.888 | 0.99979 | 0 | 0 | 1.145e-47 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.197 | 0.10478 | 53220947981.2521 | 53.2209 | 0.01879 | 20.164 | 264927.4537 | -0.15622 | 0 | 0 | 9.2085e-48 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1972 | 0.10489 | 53274979400.5224 | 53.275 | 0.018771 | 20.1844 | 265994.2419 | 0.93075 | 0 | 0 | 7.4042e-48 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1974 | 0.10499 | 53329010819.7927 | 53.329 | 0.018752 | 20.2049 | 267064.2591 | -0.62822 | 0 | 0 | 5.9521e-48 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1976 | 0.1051 | 53383042239.063 | 53.383 | 0.018733 | 20.2254 | 268137.5119 | 0.47294 | 0 | 0 | 4.7837e-48 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1978 | 0.10521 | 53437073658.3333 | 53.4371 | 0.018714 | 20.2459 | 269214.0068 | -0.99998 | 0 | 0 | 3.8438e-48 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.198 | 0.10531 | 53491105077.6036 | 53.4911 | 0.018695 | 20.2663 | 270293.7504 | -0.57557 | 0 | 0 | 3.0879e-48 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1982 | 0.10542 | 53545136496.8739 | 53.5451 | 0.018676 | 20.2868 | 271376.7492 | -0.23515 | 0 | 0 | 2.4801e-48 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1984 | 0.10553 | 53599167916.1442 | 53.5992 | 0.018657 | 20.3073 | 272463.0097 | -0.82368 | 0 | 0 | 1.9915e-48 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1986 | 0.10563 | 53653199335.4145 | 53.6532 | 0.018638 | 20.3277 | 273552.5386 | 1 | 0 | 0 | 1.5988e-48 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1988 | 0.10574 | 53707230754.6848 | 53.7072 | 0.018619 | 20.3482 | 274645.3424 | 0.89099 | 0 | 0 | 1.2833e-48 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.199 | 0.10585 | 53761262173.9551 | 53.7613 | 0.018601 | 20.3687 | 275741.4277 | -0.69538 | 0 | 0 | 1.0298e-48 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1992 | 0.10595 | 53815293593.2254 | 53.8153 | 0.018582 | 20.3892 | 276840.801 | -0.55208 | 0 | 0 | 8.2618e-49 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1994 | 0.10606 | 53869325012.4958 | 53.8693 | 0.018563 | 20.4096 | 277943.469 | 0.52596 | 0 | 0 | 6.6268e-49 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1996 | 0.10616 | 53923356431.7661 | 53.9234 | 0.018545 | 20.4301 | 279049.4383 | 0.63072 | 0 | 0 | 5.3142e-49 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.1998 | 0.10627 | 53977387851.0364 | 53.9774 | 0.018526 | 20.4506 | 280158.7154 | -0.82903 | 0 | 0 | 4.2607e-49 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2 | 0.10638 | 54031419270.3067 | 54.0314 | 0.018508 | 20.471 | 281271.307 | -0.99212 | 0 | 0 | 3.4153e-49 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2002 | 0.10648 | 54085450689.577 | 54.0855 | 0.018489 | 20.4915 | 282387.2196 | 0.86701 | 0 | 0 | 2.737e-49 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2004 | 0.10659 | 54139482108.8473 | 54.1395 | 0.018471 | 20.512 | 283506.46 | 0.21411 | 0 | 0 | 2.1929e-49 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2006 | 0.1067 | 54193513528.1176 | 54.1935 | 0.018452 | 20.5325 | 284629.0347 | 0.72453 | 0 | 0 | 1.7566e-49 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2008 | 0.1068 | 54247544947.3879 | 54.2475 | 0.018434 | 20.5529 | 285754.9504 | 0.89379 | 0 | 0 | 1.4068e-49 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.201 | 0.10691 | 54301576366.6582 | 54.3016 | 0.018416 | 20.5734 | 286884.2137 | 0.32008 | 0 | 0 | 1.1264e-49 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2012 | 0.10702 | 54355607785.9285 | 54.3556 | 0.018397 | 20.5939 | 288016.8313 | 0.92136 | 0 | 0 | 9.017e-50 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2014 | 0.10712 | 54409639205.1988 | 54.4096 | 0.018379 | 20.6143 | 289152.8098 | 0.63808 | 0 | 0 | 7.2166e-50 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2016 | 0.10723 | 54463670624.4691 | 54.4637 | 0.018361 | 20.6348 | 290292.1559 | 0.35231 | 0 | 0 | 5.7744e-50 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2018 | 0.10733 | 54517702043.7394 | 54.5177 | 0.018343 | 20.6553 | 291434.8762 | 0.92438 | 0 | 0 | 4.6194e-50 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.202 | 0.10744 | 54571733463.0098 | 54.5717 | 0.018325 | 20.6758 | 292580.9775 | -0.98226 | 0 | 0 | 3.6946e-50 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2022 | 0.10755 | 54625764882.2801 | 54.6258 | 0.018306 | 20.6962 | 293730.4664 | -0.86651 | 0 | 0 | 2.9543e-50 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2024 | 0.10765 | 54679796301.5504 | 54.6798 | 0.018288 | 20.7167 | 294883.3497 | 0.82312 | 0 | 0 | 2.3618e-50 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2026 | 0.10776 | 54733827720.8207 | 54.7338 | 0.01827 | 20.7372 | 296039.6339 | 0.91072 | 0 | 0 | 1.8877e-50 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2028 | 0.10787 | 54787859140.091 | 54.7879 | 0.018252 | 20.7576 | 297199.3258 | -0.99983 | 0 | 0 | 1.5085e-50 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.203 | 0.10797 | 54841890559.3613 | 54.8419 | 0.018234 | 20.7781 | 298362.4322 | -0.74145 | 0 | 0 | 1.2051e-50 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2032 | 0.10808 | 54895921978.6316 | 54.8959 | 0.018216 | 20.7986 | 299528.9597 | -0.1607 | 0 | 0 | 9.6259e-51 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2034 | 0.10819 | 54949953397.9019 | 54.95 | 0.018198 | 20.819 | 300698.9151 | -0.99201 | 0 | 0 | 7.6869e-51 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2036 | 0.10829 | 55003984817.1722 | 55.004 | 0.018181 | 20.8395 | 301872.305 | 0.1211 | 0 | 0 | 6.1371e-51 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2038 | 0.1084 | 55058016236.4425 | 55.058 | 0.018163 | 20.86 | 303049.1363 | -0.98322 | 0 | 0 | 4.8987e-51 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.204 | 0.10851 | 55112047655.7128 | 55.112 | 0.018145 | 20.8805 | 304229.4157 | -0.71368 | 0 | 0 | 3.9093e-51 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2042 | 0.10861 | 55166079074.9831 | 55.1661 | 0.018127 | 20.9009 | 305413.1498 | 0.14793 | 0 | 0 | 3.1191e-51 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2044 | 0.10872 | 55220110494.2534 | 55.2201 | 0.018109 | 20.9214 | 306600.3456 | -0.17686 | 0 | 0 | 2.488e-51 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2046 | 0.10882 | 55274141913.5238 | 55.2741 | 0.018092 | 20.9419 | 307791.0097 | 0.1764 | 0 | 0 | 1.9842e-51 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2048 | 0.10893 | 55328173332.7941 | 55.3282 | 0.018074 | 20.9623 | 308985.1489 | -0.15601 | 0 | 0 | 1.5821e-51 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.205 | 0.10904 | 55382204752.0644 | 55.3822 | 0.018056 | 20.9828 | 310182.7699 | 0.73856 | 0 | 0 | 1.2611e-51 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2052 | 0.10914 | 55436236171.3347 | 55.4362 | 0.018039 | 21.0033 | 311383.8797 | 0.96064 | 0 | 0 | 1.0051e-51 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2054 | 0.10925 | 55490267590.605 | 55.4903 | 0.018021 | 21.0238 | 312588.4849 | 0.085791 | 0 | 0 | 8.0086e-52 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2056 | 0.10936 | 55544299009.8753 | 55.5443 | 0.018004 | 21.0442 | 313796.5923 | 0.96867 | 0 | 0 | 6.3798e-52 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2058 | 0.10946 | 55598330429.1456 | 55.5983 | 0.017986 | 21.0647 | 315008.2088 | 0.70577 | 0 | 0 | 5.0812e-52 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.206 | 0.10957 | 55652361848.4159 | 55.6524 | 0.017969 | 21.0852 | 316223.3412 | -0.11924 | 0 | 0 | 4.046e-52 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2062 | 0.10968 | 55706393267.6862 | 55.7064 | 0.017951 | 21.1056 | 317441.9962 | 0.16266 | 0 | 0 | 3.221e-52 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2064 | 0.10978 | 55760424686.9565 | 55.7604 | 0.017934 | 21.1261 | 318664.1808 | -0.058342 | 0 | 0 | 2.5636e-52 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2066 | 0.10989 | 55814456106.2268 | 55.8145 | 0.017917 | 21.1466 | 319889.9017 | 0.42718 | 0 | 0 | 2.04e-52 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2068 | 0.10999 | 55868487525.4971 | 55.8685 | 0.017899 | 21.1671 | 321119.1657 | -0.9741 | 0 | 0 | 1.6229e-52 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.207 | 0.1101 | 55922518944.7674 | 55.9225 | 0.017882 | 21.1875 | 322351.9798 | -0.46989 | 0 | 0 | 1.2909e-52 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2072 | 0.11021 | 55976550364.0377 | 55.9766 | 0.017865 | 21.208 | 323588.3507 | -0.94442 | 0 | 0 | 1.0265e-52 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2074 | 0.11031 | 56030581783.3081 | 56.0306 | 0.017847 | 21.2285 | 324828.2853 | 0.23903 | 0 | 0 | 8.1612e-53 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2076 | 0.11042 | 56084613202.5784 | 56.0846 | 0.01783 | 21.2489 | 326071.7905 | -0.31848 | 0 | 0 | 6.487e-53 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2078 | 0.11053 | 56138644621.8487 | 56.1386 | 0.017813 | 21.2694 | 327318.8732 | 0.43835 | 0 | 0 | 5.1551e-53 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.208 | 0.11063 | 56192676041.119 | 56.1927 | 0.017796 | 21.2899 | 328569.5401 | 0.14021 | 0 | 0 | 4.0958e-53 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2082 | 0.11074 | 56246707460.3893 | 56.2467 | 0.017779 | 21.3104 | 329823.7983 | 0.58268 | 0 | 0 | 3.2534e-53 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2084 | 0.11085 | 56300738879.6596 | 56.3007 | 0.017762 | 21.3308 | 331081.6546 | 0.96366 | 0 | 0 | 2.5837e-53 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2086 | 0.11095 | 56354770298.9299 | 56.3548 | 0.017745 | 21.3513 | 332343.1158 | 0.37298 | 0 | 0 | 2.0514e-53 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2088 | 0.11106 | 56408801718.2002 | 56.4088 | 0.017728 | 21.3718 | 333608.1889 | 0.56999 | 0 | 0 | 1.6284e-53 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.209 | 0.11116 | 56462833137.4705 | 56.4628 | 0.017711 | 21.3922 | 334876.8809 | 0.89922 | 0 | 0 | 1.2924e-53 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2092 | 0.11127 | 56516864556.7408 | 56.5169 | 0.017694 | 21.4127 | 336149.1985 | -0.91087 | 0 | 0 | 1.0254e-53 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2094 | 0.11138 | 56570895976.0111 | 56.5709 | 0.017677 | 21.4332 | 337425.1488 | -0.63014 | 0 | 0 | 8.1346e-54 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2096 | 0.11148 | 56624927395.2814 | 56.6249 | 0.01766 | 21.4537 | 338704.7386 | -0.27625 | 0 | 0 | 6.4516e-54 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2098 | 0.11159 | 56678958814.5517 | 56.679 | 0.017643 | 21.4741 | 339987.9749 | -0.98462 | 0 | 0 | 5.1157e-54 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.21 | 0.1117 | 56732990233.822 | 56.733 | 0.017626 | 21.4946 | 341274.8647 | -0.55069 | 0 | 0 | 4.0554e-54 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2102 | 0.1118 | 56787021653.0923 | 56.787 | 0.01761 | 21.5151 | 342565.4148 | -0.061197 | 0 | 0 | 3.2143e-54 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2104 | 0.11191 | 56841053072.3627 | 56.8411 | 0.017593 | 21.5355 | 343859.6323 | -0.17895 | 0 | 0 | 2.547e-54 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2106 | 0.11202 | 56895084491.633 | 56.8951 | 0.017576 | 21.556 | 345157.5241 | -0.23198 | 0 | 0 | 2.0178e-54 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2108 | 0.11212 | 56949115910.9033 | 56.9491 | 0.01756 | 21.5765 | 346459.0972 | -0.92755 | 0 | 0 | 1.5982e-54 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.211 | 0.11223 | 57003147330.1736 | 57.0031 | 0.017543 | 21.5969 | 347764.3585 | 0.43804 | 0 | 0 | 1.2656e-54 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2112 | 0.11233 | 57057178749.4439 | 57.0572 | 0.017526 | 21.6174 | 349073.315 | -0.99956 | 0 | 0 | 1.002e-54 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2114 | 0.11244 | 57111210168.7142 | 57.1112 | 0.01751 | 21.6379 | 350385.9738 | -0.87892 | 0 | 0 | 7.9307e-55 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2116 | 0.11255 | 57165241587.9845 | 57.1652 | 0.017493 | 21.6584 | 351702.3417 | 0.89757 | 0 | 0 | 6.276e-55 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2118 | 0.11265 | 57219273007.2548 | 57.2193 | 0.017477 | 21.6788 | 353022.4259 | 0.98744 | 0 | 0 | 4.9654e-55 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.212 | 0.11276 | 57273304426.5251 | 57.2733 | 0.01746 | 21.6993 | 354346.2333 | -0.21358 | 0 | 0 | 3.9276e-55 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2122 | 0.11287 | 57327335845.7954 | 57.3273 | 0.017444 | 21.7198 | 355673.7709 | 1 | 0 | 0 | 3.1061e-55 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2124 | 0.11297 | 57381367265.0657 | 57.3814 | 0.017427 | 21.7402 | 357005.0458 | 0.7242 | 0 | 0 | 2.4558e-55 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2126 | 0.11308 | 57435398684.336 | 57.4354 | 0.017411 | 21.7607 | 358340.0649 | -0.60689 | 0 | 0 | 1.9413e-55 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2128 | 0.11319 | 57489430103.6063 | 57.4894 | 0.017395 | 21.7812 | 359678.8354 | -0.89308 | 0 | 0 | 1.5342e-55 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.213 | 0.11329 | 57543461522.8766 | 57.5435 | 0.017378 | 21.8017 | 361021.3642 | 0.82415 | 0 | 0 | 1.2122e-55 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2132 | 0.1134 | 57597492942.147 | 57.5975 | 0.017362 | 21.8221 | 362367.6584 | -0.66229 | 0 | 0 | 9.5757e-56 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2134 | 0.1135 | 57651524361.4173 | 57.6515 | 0.017346 | 21.8426 | 363717.7251 | 0.09414 | 0 | 0 | 7.5627e-56 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2136 | 0.11361 | 57705555780.6876 | 57.7056 | 0.017329 | 21.8631 | 365071.5713 | -0.27108 | 0 | 0 | 5.9715e-56 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2138 | 0.11372 | 57759587199.9579 | 57.7596 | 0.017313 | 21.8835 | 366429.2041 | 0.18911 | 0 | 0 | 4.7141e-56 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.214 | 0.11382 | 57813618619.2282 | 57.8136 | 0.017297 | 21.904 | 367790.6306 | -0.96543 | 0 | 0 | 3.7206e-56 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2142 | 0.11393 | 57867650038.4985 | 57.8677 | 0.017281 | 21.9245 | 369155.8578 | -0.058502 | 0 | 0 | 2.9358e-56 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2144 | 0.11404 | 57921681457.7688 | 57.9217 | 0.017265 | 21.945 | 370524.8928 | -0.68741 | 0 | 0 | 2.3161e-56 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2146 | 0.11414 | 57975712877.0391 | 57.9757 | 0.017249 | 21.9654 | 371897.7427 | 0.7061 | 0 | 0 | 1.8268e-56 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2148 | 0.11425 | 58029744296.3094 | 58.0297 | 0.017233 | 21.9859 | 373274.4147 | 0.12923 | 0 | 0 | 1.4405e-56 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.215 | 0.11436 | 58083775715.5797 | 58.0838 | 0.017217 | 22.0064 | 374654.9158 | 0.93643 | 0 | 0 | 1.1357e-56 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2152 | 0.11446 | 58137807134.85 | 58.1378 | 0.017201 | 22.0268 | 376039.2531 | -0.10709 | 0 | 0 | 8.9513e-57 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2154 | 0.11457 | 58191838554.1203 | 58.1918 | 0.017185 | 22.0473 | 377427.4338 | 0.29166 | 0 | 0 | 7.0539e-57 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2156 | 0.11467 | 58245869973.3906 | 58.2459 | 0.017169 | 22.0678 | 378819.465 | 0.0096767 | 0 | 0 | 5.5575e-57 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2158 | 0.11478 | 58299901392.6609 | 58.2999 | 0.017153 | 22.0883 | 380215.3538 | 0.85799 | 0 | 0 | 4.3775e-57 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.216 | 0.11489 | 58353932811.9313 | 58.3539 | 0.017137 | 22.1087 | 381615.1074 | -0.35733 | 0 | 0 | 3.4473e-57 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2162 | 0.11499 | 58407964231.2016 | 58.408 | 0.017121 | 22.1292 | 383018.7328 | 0.85824 | 0 | 0 | 2.7142e-57 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2164 | 0.1151 | 58461995650.4719 | 58.462 | 0.017105 | 22.1497 | 384426.2374 | 0.81965 | 0 | 0 | 2.1365e-57 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2166 | 0.11521 | 58516027069.7422 | 58.516 | 0.017089 | 22.1701 | 385837.6282 | -0.14462 | 0 | 0 | 1.6814e-57 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2168 | 0.11531 | 58570058489.0125 | 58.5701 | 0.017074 | 22.1906 | 387252.9123 | 0.989 | 0 | 0 | 1.3229e-57 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.217 | 0.11542 | 58624089908.2828 | 58.6241 | 0.017058 | 22.2111 | 388672.0971 | 0.57057 | 0 | 0 | 1.0407e-57 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2172 | 0.11553 | 58678121327.5531 | 58.6781 | 0.017042 | 22.2316 | 390095.1895 | -0.52966 | 0 | 0 | 8.1845e-58 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2174 | 0.11563 | 58732152746.8234 | 58.7322 | 0.017026 | 22.252 | 391522.1969 | -0.95874 | 0 | 0 | 6.4354e-58 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2176 | 0.11574 | 58786184166.0937 | 58.7862 | 0.017011 | 22.2725 | 392953.1265 | 0.34683 | 0 | 0 | 5.0589e-58 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2178 | 0.11585 | 58840215585.364 | 58.8402 | 0.016995 | 22.293 | 394387.9853 | -0.93319 | 0 | 0 | 3.976e-58 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.218 | 0.11595 | 58894247004.6343 | 58.8942 | 0.01698 | 22.3134 | 395826.7807 | -0.95124 | 0 | 0 | 3.1242e-58 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2182 | 0.11606 | 58948278423.9046 | 58.9483 | 0.016964 | 22.3339 | 397269.5198 | 0.48743 | 0 | 0 | 2.4544e-58 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2184 | 0.11616 | 59002309843.1749 | 59.0023 | 0.016948 | 22.3544 | 398716.2099 | -0.86659 | 0 | 0 | 1.9277e-58 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2186 | 0.11627 | 59056341262.4453 | 59.0563 | 0.016933 | 22.3748 | 400166.8582 | -0.27712 | 0 | 0 | 1.5137e-58 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2188 | 0.11638 | 59110372681.7156 | 59.1104 | 0.016918 | 22.3953 | 401621.4719 | 0.33075 | 0 | 0 | 1.1884e-58 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.219 | 0.11648 | 59164404100.9859 | 59.1644 | 0.016902 | 22.4158 | 403080.0582 | 0.94063 | 0 | 0 | 9.3276e-59 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2192 | 0.11659 | 59218435520.2562 | 59.2184 | 0.016887 | 22.4363 | 404542.6245 | -0.19027 | 0 | 0 | 7.3196e-59 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2194 | 0.1167 | 59272466939.5265 | 59.2725 | 0.016871 | 22.4567 | 406009.1779 | 0.69021 | 0 | 0 | 5.7426e-59 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2196 | 0.1168 | 59326498358.7968 | 59.3265 | 0.016856 | 22.4772 | 407479.7258 | 0.46115 | 0 | 0 | 4.5043e-59 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2198 | 0.11691 | 59380529778.0671 | 59.3805 | 0.016841 | 22.4977 | 408954.2753 | 0.61571 | 0 | 0 | 3.5323e-59 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.22 | 0.11702 | 59434561197.3374 | 59.4346 | 0.016825 | 22.5181 | 410432.8338 | 0.45142 | 0 | 0 | 2.7694e-59 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2202 | 0.11712 | 59488592616.6077 | 59.4886 | 0.01681 | 22.5386 | 411915.4085 | 0.66341 | 0 | 0 | 2.1708e-59 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2204 | 0.11723 | 59542624035.878 | 59.5426 | 0.016795 | 22.5591 | 413402.0068 | -0.10027 | 0 | 0 | 1.7012e-59 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2206 | 0.11733 | 59596655455.1483 | 59.5967 | 0.016779 | 22.5796 | 414892.6359 | 0.98769 | 0 | 0 | 1.3329e-59 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2208 | 0.11744 | 59650686874.4186 | 59.6507 | 0.016764 | 22.6 | 416387.3031 | 0.63097 | 0 | 0 | 1.0441e-59 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.221 | 0.11755 | 59704718293.6889 | 59.7047 | 0.016749 | 22.6205 | 417886.0157 | -0.75504 | 0 | 0 | 8.1773e-60 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2212 | 0.11765 | 59758749712.9592 | 59.7587 | 0.016734 | 22.641 | 419388.781 | -0.9327 | 0 | 0 | 6.4028e-60 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2214 | 0.11776 | 59812781132.2296 | 59.8128 | 0.016719 | 22.6614 | 420895.6064 | -0.71783 | 0 | 0 | 5.0122e-60 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2216 | 0.11787 | 59866812551.4999 | 59.8668 | 0.016704 | 22.6819 | 422406.4992 | 0.55419 | 0 | 0 | 3.9228e-60 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2218 | 0.11797 | 59920843970.7702 | 59.9208 | 0.016689 | 22.7024 | 423921.4666 | 0.96546 | 0 | 0 | 3.0694e-60 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.222 | 0.11808 | 59974875390.0405 | 59.9749 | 0.016674 | 22.7229 | 425440.5162 | -0.17323 | 0 | 0 | 2.4012e-60 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2222 | 0.11819 | 60028906809.3108 | 60.0289 | 0.016659 | 22.7433 | 426963.6551 | 0.65003 | 0 | 0 | 1.878e-60 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2224 | 0.11829 | 60082938228.5811 | 60.0829 | 0.016644 | 22.7638 | 428490.8908 | 0.2821 | 0 | 0 | 1.4685e-60 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2226 | 0.1184 | 60136969647.8514 | 60.137 | 0.016629 | 22.7843 | 430022.2305 | 0.89038 | 0 | 0 | 1.1481e-60 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2228 | 0.1185 | 60191001067.1217 | 60.191 | 0.016614 | 22.8047 | 431557.6818 | -0.30564 | 0 | 0 | 8.9734e-61 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.223 | 0.11861 | 60245032486.392 | 60.245 | 0.016599 | 22.8252 | 433097.2519 | -0.47968 | 0 | 0 | 7.0121e-61 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2232 | 0.11872 | 60299063905.6623 | 60.2991 | 0.016584 | 22.8457 | 434640.9482 | 0.99462 | 0 | 0 | 5.4782e-61 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2234 | 0.11882 | 60353095324.9326 | 60.3531 | 0.016569 | 22.8662 | 436188.7781 | -0.64372 | 0 | 0 | 4.279e-61 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2236 | 0.11893 | 60407126744.2029 | 60.4071 | 0.016554 | 22.8866 | 437740.749 | -0.662 | 0 | 0 | 3.3415e-61 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2238 | 0.11904 | 60461158163.4732 | 60.4612 | 0.01654 | 22.9071 | 439296.8683 | 0.98319 | 0 | 0 | 2.6088e-61 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.224 | 0.11914 | 60515189582.7435 | 60.5152 | 0.016525 | 22.9276 | 440857.1434 | -0.61153 | 0 | 0 | 2.0364e-61 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2242 | 0.11925 | 60569221002.0139 | 60.5692 | 0.01651 | 22.948 | 442421.5817 | -0.55067 | 0 | 0 | 1.5892e-61 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2244 | 0.11936 | 60623252421.2842 | 60.6233 | 0.016495 | 22.9685 | 443990.1907 | 0.99944 | 0 | 0 | 1.2399e-61 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2246 | 0.11946 | 60677283840.5545 | 60.6773 | 0.016481 | 22.989 | 445562.9777 | -0.37693 | 0 | 0 | 9.6719e-62 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2248 | 0.11957 | 60731315259.8248 | 60.7313 | 0.016466 | 23.0095 | 447139.9502 | -0.27584 | 0 | 0 | 7.5429e-62 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.225 | 0.11967 | 60785346679.0951 | 60.7853 | 0.016451 | 23.0299 | 448721.1156 | 0.94039 | 0 | 0 | 5.8812e-62 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2252 | 0.11978 | 60839378098.3654 | 60.8394 | 0.016437 | 23.0504 | 450306.4814 | -0.08518 | 0 | 0 | 4.5846e-62 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2254 | 0.11989 | 60893409517.6357 | 60.8934 | 0.016422 | 23.0709 | 451896.055 | -0.013015 | 0 | 0 | 3.5731e-62 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2256 | 0.11999 | 60947440936.906 | 60.9474 | 0.016408 | 23.0913 | 453489.8439 | 0.84918 | 0 | 0 | 2.7841e-62 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2258 | 0.1201 | 61001472356.1763 | 61.0015 | 0.016393 | 23.1118 | 455087.8556 | 0.044572 | 0 | 0 | 2.1688e-62 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.226 | 0.12021 | 61055503775.4466 | 61.0555 | 0.016379 | 23.1323 | 456690.0974 | 0.014985 | 0 | 0 | 1.6892e-62 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2262 | 0.12031 | 61109535194.7169 | 61.1095 | 0.016364 | 23.1527 | 458296.577 | 0.89602 | 0 | 0 | 1.3153e-62 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2264 | 0.12042 | 61163566613.9872 | 61.1636 | 0.01635 | 23.1732 | 459907.3017 | -0.19711 | 0 | 0 | 1.024e-62 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2266 | 0.12053 | 61217598033.2575 | 61.2176 | 0.016335 | 23.1937 | 461522.2791 | -0.38682 | 0 | 0 | 7.9698e-63 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2268 | 0.12063 | 61271629452.5278 | 61.2716 | 0.016321 | 23.2142 | 463141.5166 | 0.98969 | 0 | 0 | 6.2017e-63 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.227 | 0.12074 | 61325660871.7982 | 61.3257 | 0.016306 | 23.2346 | 464765.0219 | -0.8502 | 0 | 0 | 4.8248e-63 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2272 | 0.12084 | 61379692291.0685 | 61.3797 | 0.016292 | 23.2551 | 466392.8023 | -0.99293 | 0 | 0 | 3.7528e-63 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2274 | 0.12095 | 61433723710.3388 | 61.4337 | 0.016278 | 23.2756 | 468024.8654 | 0.11304 | 0 | 0 | 2.9183e-63 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2276 | 0.12106 | 61487755129.6091 | 61.4878 | 0.016263 | 23.296 | 469661.2187 | -0.50529 | 0 | 0 | 2.2689e-63 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2278 | 0.12116 | 61541786548.8794 | 61.5418 | 0.016249 | 23.3165 | 471301.8698 | 0.20847 | 0 | 0 | 1.7636e-63 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.228 | 0.12127 | 61595817968.1497 | 61.5958 | 0.016235 | 23.337 | 472946.8262 | -0.85635 | 0 | 0 | 1.3705e-63 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2282 | 0.12138 | 61649849387.42 | 61.6498 | 0.016221 | 23.3575 | 474596.0955 | 0.88896 | 0 | 0 | 1.0648e-63 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2284 | 0.12148 | 61703880806.6903 | 61.7039 | 0.016206 | 23.3779 | 476249.6851 | -0.016795 | 0 | 0 | 8.2713e-64 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2286 | 0.12159 | 61757912225.9606 | 61.7579 | 0.016192 | 23.3984 | 477907.6026 | 0.73561 | 0 | 0 | 6.4235e-64 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2288 | 0.1217 | 61811943645.2309 | 61.8119 | 0.016178 | 23.4189 | 479569.8557 | -0.45825 | 0 | 0 | 4.9874e-64 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.229 | 0.1218 | 61865975064.5012 | 61.866 | 0.016164 | 23.4393 | 481236.4518 | 0.88007 | 0 | 0 | 3.8715e-64 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2292 | 0.12191 | 61920006483.7715 | 61.92 | 0.01615 | 23.4598 | 482907.3986 | 0.64078 | 0 | 0 | 3.0047e-64 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2294 | 0.12202 | 61974037903.0418 | 61.974 | 0.016136 | 23.4803 | 484582.7036 | -0.99923 | 0 | 0 | 2.3314e-64 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2296 | 0.12212 | 62028069322.3121 | 62.0281 | 0.016122 | 23.5008 | 486262.3745 | 0.43538 | 0 | 0 | 1.8086e-64 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2298 | 0.12223 | 62082100741.5825 | 62.0821 | 0.016108 | 23.5212 | 487946.4188 | 0.56554 | 0 | 0 | 1.4027e-64 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.23 | 0.12233 | 62136132160.8528 | 62.1361 | 0.016094 | 23.5417 | 489634.8441 | -0.91294 | 0 | 0 | 1.0876e-64 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2302 | 0.12244 | 62190163580.1231 | 62.1902 | 0.01608 | 23.5622 | 491327.6581 | 0.99646 | 0 | 0 | 8.4317e-65 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2304 | 0.12255 | 62244194999.3934 | 62.2442 | 0.016066 | 23.5826 | 493024.8683 | 0.78626 | 0 | 0 | 6.535e-65 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2306 | 0.12265 | 62298226418.6637 | 62.2982 | 0.016052 | 23.6031 | 494726.4824 | 0.8947 | 0 | 0 | 5.0639e-65 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2308 | 0.12276 | 62352257837.934 | 62.3523 | 0.016038 | 23.6236 | 496432.508 | -0.94853 | 0 | 0 | 3.9231e-65 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.231 | 0.12287 | 62406289257.2043 | 62.4063 | 0.016024 | 23.6441 | 498142.9528 | -0.45705 | 0 | 0 | 3.0386e-65 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2312 | 0.12297 | 62460320676.4746 | 62.4603 | 0.01601 | 23.6645 | 499857.8243 | -0.79117 | 0 | 0 | 2.353e-65 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2314 | 0.12308 | 62514352095.7449 | 62.5144 | 0.015996 | 23.685 | 501577.1303 | 0.057869 | 0 | 0 | 1.8217e-65 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2316 | 0.12319 | 62568383515.0152 | 62.5684 | 0.015983 | 23.7055 | 503300.8784 | -0.86451 | 0 | 0 | 1.4101e-65 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2318 | 0.12329 | 62622414934.2855 | 62.6224 | 0.015969 | 23.7259 | 505029.0763 | -0.6611 | 0 | 0 | 1.0912e-65 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.232 | 0.1234 | 62676446353.5558 | 62.6764 | 0.015955 | 23.7464 | 506761.7316 | -0.79284 | 0 | 0 | 8.4424e-66 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2322 | 0.1235 | 62730477772.8261 | 62.7305 | 0.015941 | 23.7669 | 508498.852 | 0.67067 | 0 | 0 | 6.5304e-66 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2324 | 0.12361 | 62784509192.0965 | 62.7845 | 0.015927 | 23.7873 | 510240.4451 | 0.95074 | 0 | 0 | 5.0502e-66 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2326 | 0.12372 | 62838540611.3668 | 62.8385 | 0.015914 | 23.8078 | 511986.5188 | 0.94386 | 0 | 0 | 3.9047e-66 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2328 | 0.12382 | 62892572030.6371 | 62.8926 | 0.0159 | 23.8283 | 513737.0807 | -0.93658 | 0 | 0 | 3.0184e-66 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.233 | 0.12393 | 62946603449.9074 | 62.9466 | 0.015886 | 23.8488 | 515492.1384 | 0.74215 | 0 | 0 | 2.3327e-66 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2332 | 0.12404 | 63000634869.1777 | 63.0006 | 0.015873 | 23.8692 | 517251.6997 | 0.89378 | 0 | 0 | 1.8024e-66 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2334 | 0.12414 | 63054666288.448 | 63.0547 | 0.015859 | 23.8897 | 519015.7723 | -0.38642 | 0 | 0 | 1.3923e-66 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2336 | 0.12425 | 63108697707.7183 | 63.1087 | 0.015846 | 23.9102 | 520784.364 | 0.49841 | 0 | 0 | 1.0753e-66 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2338 | 0.12436 | 63162729126.9886 | 63.1627 | 0.015832 | 23.9306 | 522557.4824 | -0.67309 | 0 | 0 | 8.3031e-67 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.234 | 0.12446 | 63216760546.2589 | 63.2168 | 0.015819 | 23.9511 | 524335.1353 | -0.24724 | 0 | 0 | 6.4098e-67 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2342 | 0.12457 | 63270791965.5292 | 63.2708 | 0.015805 | 23.9716 | 526117.3304 | 0.91757 | 0 | 0 | 4.9471e-67 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2344 | 0.12467 | 63324823384.7995 | 63.3248 | 0.015792 | 23.9921 | 527904.0755 | -0.33427 | 0 | 0 | 3.8174e-67 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2346 | 0.12478 | 63378854804.0698 | 63.3789 | 0.015778 | 24.0125 | 529695.3784 | -0.80512 | 0 | 0 | 2.945e-67 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2348 | 0.12489 | 63432886223.3401 | 63.4329 | 0.015765 | 24.033 | 531491.2467 | 0.18567 | 0 | 0 | 2.2715e-67 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.235 | 0.12499 | 63486917642.6104 | 63.4869 | 0.015751 | 24.0535 | 533291.6884 | 0.12199 | 0 | 0 | 1.7516e-67 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2352 | 0.1251 | 63540949061.8808 | 63.5409 | 0.015738 | 24.0739 | 535096.7111 | 0.95534 | 0 | 0 | 1.3504e-67 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2354 | 0.12521 | 63594980481.1511 | 63.595 | 0.015725 | 24.0944 | 536906.3226 | 0.93794 | 0 | 0 | 1.0409e-67 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2356 | 0.12531 | 63649011900.4214 | 63.649 | 0.015711 | 24.1149 | 538720.5307 | 0.28838 | 0 | 0 | 8.0209e-68 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2358 | 0.12542 | 63703043319.6917 | 63.703 | 0.015698 | 24.1354 | 540539.3433 | -0.12264 | 0 | 0 | 6.1797e-68 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.236 | 0.12553 | 63757074738.962 | 63.7571 | 0.015685 | 24.1558 | 542362.768 | -0.98924 | 0 | 0 | 4.76e-68 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2362 | 0.12563 | 63811106158.2323 | 63.8111 | 0.015671 | 24.1763 | 544190.8129 | -0.87326 | 0 | 0 | 3.6657e-68 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2364 | 0.12574 | 63865137577.5026 | 63.8651 | 0.015658 | 24.1968 | 546023.4855 | 0.81666 | 0 | 0 | 2.8223e-68 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2366 | 0.12584 | 63919168996.7729 | 63.9192 | 0.015645 | 24.2172 | 547860.7938 | -0.99579 | 0 | 0 | 2.1725e-68 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2368 | 0.12595 | 63973200416.0432 | 63.9732 | 0.015632 | 24.2377 | 549702.7456 | -0.47985 | 0 | 0 | 1.672e-68 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.237 | 0.12606 | 64027231835.3135 | 64.0272 | 0.015618 | 24.2582 | 551549.3487 | -0.91432 | 0 | 0 | 1.2864e-68 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2372 | 0.12616 | 64081263254.5838 | 64.0813 | 0.015605 | 24.2787 | 553400.6111 | 0.28557 | 0 | 0 | 9.8959e-69 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2374 | 0.12627 | 64135294673.8541 | 64.1353 | 0.015592 | 24.2991 | 555256.5404 | -0.86318 | 0 | 0 | 7.6107e-69 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2376 | 0.12638 | 64189326093.1245 | 64.1893 | 0.015579 | 24.3196 | 557117.1446 | -0.25726 | 0 | 0 | 5.852e-69 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2378 | 0.12648 | 64243357512.3947 | 64.2434 | 0.015566 | 24.3401 | 558982.4315 | -0.88162 | 0 | 0 | 4.4986e-69 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.238 | 0.12659 | 64297388931.6651 | 64.2974 | 0.015553 | 24.3605 | 560852.4091 | 0.34231 | 0 | 0 | 3.4575e-69 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2382 | 0.1267 | 64351420350.9354 | 64.3514 | 0.01554 | 24.381 | 562727.0851 | -0.93354 | 0 | 0 | 2.6568e-69 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2384 | 0.1268 | 64405451770.2057 | 64.4055 | 0.015527 | 24.4015 | 564606.4675 | -0.47432 | 0 | 0 | 2.041e-69 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2386 | 0.12691 | 64459483189.476 | 64.4595 | 0.015514 | 24.422 | 566490.5641 | -0.97641 | 0 | 0 | 1.5676e-69 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2388 | 0.12701 | 64513514608.7463 | 64.5135 | 0.015501 | 24.4424 | 568379.3829 | 0.5904 | 0 | 0 | 1.2037e-69 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.239 | 0.12712 | 64567546028.0166 | 64.5675 | 0.015488 | 24.4629 | 570272.9317 | -0.99421 | 0 | 0 | 9.2414e-70 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2392 | 0.12723 | 64621577447.2869 | 64.6216 | 0.015475 | 24.4834 | 572171.2185 | -0.6429 | 0 | 0 | 7.0932e-70 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2394 | 0.12733 | 64675608866.5572 | 64.6756 | 0.015462 | 24.5038 | 574074.2511 | -0.995 | 0 | 0 | 5.4432e-70 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2396 | 0.12744 | 64729640285.8275 | 64.7296 | 0.015449 | 24.5243 | 575982.0376 | 0.59001 | 0 | 0 | 4.1761e-70 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2398 | 0.12755 | 64783671705.0978 | 64.7837 | 0.015436 | 24.5448 | 577894.5857 | -0.96691 | 0 | 0 | 3.2032e-70 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.24 | 0.12765 | 64837703124.3681 | 64.8377 | 0.015423 | 24.5652 | 579811.9036 | -0.35838 | 0 | 0 | 2.4565e-70 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2402 | 0.12776 | 64891734543.6384 | 64.8917 | 0.01541 | 24.5857 | 581733.999 | -0.79912 | 0 | 0 | 1.8834e-70 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2404 | 0.12787 | 64945765962.9087 | 64.9458 | 0.015397 | 24.6062 | 583660.8799 | -0.15816 | 0 | 0 | 1.4437e-70 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2406 | 0.12797 | 64999797382.179 | 64.9998 | 0.015385 | 24.6267 | 585592.5544 | -0.24372 | 0 | 0 | 1.1064e-70 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2408 | 0.12808 | 65053828801.4494 | 65.0538 | 0.015372 | 24.6471 | 587529.0303 | 0.84611 | 0 | 0 | 8.4769e-71 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.241 | 0.12818 | 65107860220.7197 | 65.1079 | 0.015359 | 24.6676 | 589470.3156 | 0.71015 | 0 | 0 | 6.4935e-71 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2412 | 0.12829 | 65161891639.99 | 65.1619 | 0.015346 | 24.6881 | 591416.4184 | -0.7802 | 0 | 0 | 4.9731e-71 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2414 | 0.1284 | 65215923059.2603 | 65.2159 | 0.015334 | 24.7085 | 593367.3465 | 0.78077 | 0 | 0 | 3.8078e-71 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2416 | 0.1285 | 65269954478.5306 | 65.27 | 0.015321 | 24.729 | 595323.108 | -0.71383 | 0 | 0 | 2.9149e-71 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2418 | 0.12861 | 65323985897.8009 | 65.324 | 0.015308 | 24.7495 | 597283.7108 | -0.86442 | 0 | 0 | 2.2309e-71 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.242 | 0.12872 | 65378017317.0712 | 65.378 | 0.015296 | 24.77 | 599249.1631 | 0.14032 | 0 | 0 | 1.707e-71 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2422 | 0.12882 | 65432048736.3415 | 65.432 | 0.015283 | 24.7904 | 601219.4726 | 0.38039 | 0 | 0 | 1.3059e-71 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2424 | 0.12893 | 65486080155.6118 | 65.4861 | 0.01527 | 24.8109 | 603194.6476 | 0.47643 | 0 | 0 | 9.988e-72 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2426 | 0.12904 | 65540111574.8821 | 65.5401 | 0.015258 | 24.8314 | 605174.696 | -0.34142 | 0 | 0 | 7.6375e-72 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2428 | 0.12914 | 65594142994.1524 | 65.5941 | 0.015245 | 24.8518 | 607159.6258 | 0.2068 | 0 | 0 | 5.8388e-72 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.243 | 0.12925 | 65648174413.4227 | 65.6482 | 0.015233 | 24.8723 | 609149.4451 | 0.83207 | 0 | 0 | 4.4628e-72 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2432 | 0.12936 | 65702205832.693 | 65.7022 | 0.01522 | 24.8928 | 611144.1619 | -0.70915 | 0 | 0 | 3.4103e-72 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2434 | 0.12946 | 65756237251.9634 | 65.7562 | 0.015208 | 24.9133 | 613143.7843 | -0.70602 | 0 | 0 | 2.6054e-72 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2436 | 0.12957 | 65810268671.2337 | 65.8103 | 0.015195 | 24.9337 | 615148.3202 | -0.55138 | 0 | 0 | 1.9901e-72 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2438 | 0.12967 | 65864300090.504 | 65.8643 | 0.015183 | 24.9542 | 617157.7779 | -0.98475 | 0 | 0 | 1.5197e-72 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.244 | 0.12978 | 65918331509.7743 | 65.9183 | 0.01517 | 24.9747 | 619172.1653 | 0.89975 | 0 | 0 | 1.1603e-72 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2442 | 0.12989 | 65972362929.0446 | 65.9724 | 0.015158 | 24.9951 | 621191.4904 | -0.96449 | 0 | 0 | 8.8567e-73 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2444 | 0.12999 | 66026394348.3149 | 66.0264 | 0.015145 | 25.0156 | 623215.7615 | -0.68359 | 0 | 0 | 6.7589e-73 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2446 | 0.1301 | 66080425767.5852 | 66.0804 | 0.015133 | 25.0361 | 625244.9866 | -0.83956 | 0 | 0 | 5.1569e-73 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2448 | 0.13021 | 66134457186.8555 | 66.1345 | 0.015121 | 25.0566 | 627279.1737 | -0.5482 | 0 | 0 | 3.9337e-73 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.245 | 0.13031 | 66188488606.1258 | 66.1885 | 0.015108 | 25.077 | 629318.331 | 0.74722 | 0 | 0 | 3e-73 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2452 | 0.13042 | 66242520025.3961 | 66.2425 | 0.015096 | 25.0975 | 631362.4666 | 0.19608 | 0 | 0 | 2.2874e-73 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2454 | 0.13053 | 66296551444.6664 | 66.2966 | 0.015084 | 25.118 | 633411.5886 | -0.56975 | 0 | 0 | 1.7437e-73 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2456 | 0.13063 | 66350582863.9367 | 66.3506 | 0.015071 | 25.1384 | 635465.705 | -0.1208 | 0 | 0 | 1.3289e-73 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2458 | 0.13074 | 66404614283.207 | 66.4046 | 0.015059 | 25.1589 | 637524.8242 | 0.99728 | 0 | 0 | 1.0126e-73 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.246 | 0.13084 | 66458645702.4773 | 66.4586 | 0.015047 | 25.1794 | 639588.9541 | -0.99956 | 0 | 0 | 7.714e-74 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2462 | 0.13095 | 66512677121.7476 | 66.5127 | 0.015035 | 25.1999 | 641658.1029 | 0.42576 | 0 | 0 | 5.8752e-74 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2464 | 0.13106 | 66566708541.018 | 66.5667 | 0.015023 | 25.2203 | 643732.2787 | 0.9186 | 0 | 0 | 4.4737e-74 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2466 | 0.13116 | 66620739960.2883 | 66.6207 | 0.01501 | 25.2408 | 645811.4898 | 0.59821 | 0 | 0 | 3.4058e-74 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2468 | 0.13127 | 66674771379.5586 | 66.6748 | 0.014998 | 25.2613 | 647895.7443 | -0.90088 | 0 | 0 | 2.5922e-74 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.247 | 0.13138 | 66728802798.8289 | 66.7288 | 0.014986 | 25.2817 | 649985.0503 | 0.82766 | 0 | 0 | 1.9726e-74 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2472 | 0.13148 | 66782834218.0992 | 66.7828 | 0.014974 | 25.3022 | 652079.416 | -0.88666 | 0 | 0 | 1.5007e-74 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2474 | 0.13159 | 66836865637.3695 | 66.8369 | 0.014962 | 25.3227 | 654178.8496 | -0.9327 | 0 | 0 | 1.1415e-74 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2476 | 0.1317 | 66890897056.6398 | 66.8909 | 0.01495 | 25.3431 | 656283.3592 | -0.99993 | 0 | 0 | 8.6802e-75 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2478 | 0.1318 | 66944928475.9101 | 66.9449 | 0.014938 | 25.3636 | 658392.9532 | -0.026044 | 0 | 0 | 6.5994e-75 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.248 | 0.13191 | 66998959895.1804 | 66.999 | 0.014926 | 25.3841 | 660507.6396 | 0.40831 | 0 | 0 | 5.0163e-75 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2482 | 0.13201 | 67052991314.4507 | 67.053 | 0.014914 | 25.4046 | 662627.4267 | 0.35914 | 0 | 0 | 3.8121e-75 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2484 | 0.13212 | 67107022733.721 | 67.107 | 0.014902 | 25.425 | 664752.3227 | -0.72566 | 0 | 0 | 2.8963e-75 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2486 | 0.13223 | 67161054152.9913 | 67.1611 | 0.01489 | 25.4455 | 666882.3357 | -0.73472 | 0 | 0 | 2.2001e-75 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2488 | 0.13233 | 67215085572.2616 | 67.2151 | 0.014878 | 25.466 | 669017.4742 | 0.31393 | 0 | 0 | 1.6708e-75 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.249 | 0.13244 | 67269116991.532 | 67.2691 | 0.014866 | 25.4864 | 671157.7461 | 0.50386 | 0 | 0 | 1.2686e-75 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2492 | 0.13255 | 67323148410.8023 | 67.3231 | 0.014854 | 25.5069 | 673303.1599 | -0.23193 | 0 | 0 | 9.6301e-76 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2494 | 0.13265 | 67377179830.0726 | 67.3772 | 0.014842 | 25.5274 | 675453.7238 | -0.9295 | 0 | 0 | 7.3086e-76 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2496 | 0.13276 | 67431211249.3429 | 67.4312 | 0.01483 | 25.5479 | 677609.4459 | -0.56751 | 0 | 0 | 5.5455e-76 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2498 | 0.13287 | 67485242668.6132 | 67.4852 | 0.014818 | 25.5683 | 679770.3346 | -0.90463 | 0 | 0 | 4.2069e-76 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.25 | 0.13297 | 67539274087.8835 | 67.5393 | 0.014806 | 25.5888 | 681936.3982 | -0.36691 | 0 | 0 | 3.1906e-76 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2502 | 0.13308 | 67593305507.1538 | 67.5933 | 0.014794 | 25.6093 | 684107.6448 | 0.70456 | 0 | 0 | 2.4193e-76 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2504 | 0.13318 | 67647336926.4241 | 67.6473 | 0.014783 | 25.6297 | 686284.0828 | -0.09596 | 0 | 0 | 1.8341e-76 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2506 | 0.13329 | 67701368345.6944 | 67.7014 | 0.014771 | 25.6502 | 688465.7205 | -0.99477 | 0 | 0 | 1.3901e-76 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2508 | 0.1334 | 67755399764.9647 | 67.7554 | 0.014759 | 25.6707 | 690652.5661 | -0.98108 | 0 | 0 | 1.0534e-76 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.251 | 0.1335 | 67809431184.235 | 67.8094 | 0.014747 | 25.6912 | 692844.628 | -0.83923 | 0 | 0 | 7.9802e-77 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2512 | 0.13361 | 67863462603.5053 | 67.8635 | 0.014735 | 25.7116 | 695041.9145 | 0.73964 | 0 | 0 | 6.0444e-77 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2514 | 0.13372 | 67917494022.7756 | 67.9175 | 0.014724 | 25.7321 | 697244.4338 | -0.53933 | 0 | 0 | 4.5772e-77 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2516 | 0.13382 | 67971525442.0459 | 67.9715 | 0.014712 | 25.7526 | 699452.1944 | 0.97545 | 0 | 0 | 3.4654e-77 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.2518 | 0.13393 | 68025556861.3162 | 68.0256 | 0.0147 | 25.773 | 701665.2044 | 0.019936 | 0 | 0 | 2.623e-77 | 1 | 0 | 0 | 0 | 0 | 0
[6 empty] | 0.252 | 0.13404 | 68079588280.5865 | 68.0796 | 0.014689 | 25.7935 | 703883.4723 | -0.27976 | 0 | 0 | 1.985e-77 | 1 | 0 | 0 | 0 | 0 | 0
Note on C8: Depends on the microscope lens. 1.6 mm : Fujiyoshi's Jeol 1.9 mm : Hitachi H8000 2.0 mm : Philips CM12 3.6 mm : Philips CM100 BioTwin
Note on C10: Jong and van Dyck, Ultramicr. 49 (1993) 66-80 ai(therm. el. source) = 0.7 mrad ai(field emission source) = 0.1 mrad FEG is 10^3 to 10^4 times brighter But if you use a FEG correctly, then you can reach much smaller opening angles, ai(field emission source) = 0.001 mrad Note by Wim Hagen, EMBL, on July 14, 2013: Dear Henning, I was using your ctf-simulation excel sheet with John to look at some settings. For the opening angle of the source you reference de Jong and van Dyck. This paper deals with ultimate resolution and therefore they use a high magnification with a condensed beam to get maximum beam current, this increases the half opening angle of the source as seen from the specimen. Peter Tiemeijer states in his Titan Condenser Manual: I=B⋅πα2 ⋅(π/4)D2 ⋅Vrel where I = beam current on the specimen B = brightness of the FEG α = half-coherence angle D = diameter of the illuminated area on the specimen Vrel = the relativized high tension I measured gun brightness on the EMBL Krios. If I then calculate with a starting dose rate of 20 e/A2*sec at 300KV, and adjust dose for the KV used (linear energy transfer), I get this: KV 60 120 200 300 Gun brightness [10^7 A/m^2*sr] 0.59 1.39 2.03 4.15 dose rate [e/A^2*sec] 7.80 12.27 16.43 20.00 Semi angle alpha [mrad] 0.01033 0.00580 0.00416 0.00252 These much smaller semi-angles have quite a dramatic effect on the envelope, maybe this can be mentioned in the sheet for low dose users. Best, Wim Hagen
Note on C14: Should be similar to Cs, depending on the microscope lens.
Note on C15: Reimer, S. 219 : ∆E(therm. el. source) = 1.0 … 2.0 eV ∆E(field emission gun) = 0.2 … 0.7 eV