## Home
Toolkit Manufacturing (v. April 2022)
Please select the device to continue:
| Implants | (Recommended to prepare under yellow light or dark)
V-shape implants
Windows
Numbered windows
Holder
Remounting stage
Mold
Molds for V-shape implants
Note: All the mask and 3D printing desings can be found in the lab server

## Z01 MR Summary (hidden)
| MicroResist mr-DWL Process Parameters |  |  |  |  | Home
MICRORESIST modified SU8 previously used for MLA150 405nm exposure |  | http://microresist.de/en/product/negative-photoresists-2
|  | a) Enter targeted thickness [µm]: |  |  | Note: Thickness from 3 um to 100 um
|  | LSM200 & LSM250 | 50
b) Available product: choose from the rows marked with green
c) Use parameters from table
Wafer Prep
|  |  | mr-DWL-5 | mr-DWL-40
Spin Coat
|  | Spin speed |  |  |  | Velocity profile
|  | [rpm] | NA | 1590.6685
Note: Planarisation by putting wafer in glass Petri dish with matching cover, which creates solvent saturated atmosphere, and allows resist to flow/planarize without forming dry crust on top, and reduces edge beads.
Planarisation
Planarisation time
|  | Step time @ 30°C/RT [min] | 60 | 60
Soft Bake |  | Softbake time |  |  |  | Note: Due to a different solvent formulation, MicroResist mr-DWL allows for short Bake times and quicker ramps compared to Gersteltec. Ramp 8°C/min
|  | Step time @ 50°C [min] | 2 | 5
|  | Step time @ 90°C [min] | 4 | 6.875
Note: recommended dose is for general use on Si substrates. For high aspect ratio, decrease dose with up to 15%. For stress sensitive large areas, increase dose with up to 15%. For float glass substrates, increase dose with 40%.
Expose |  | Exposure dose
|  | [mJ/cm2] | NA | 450
Post Exposure Bake |  | PEB time |  |  |  | Note: Place a single sheet of cleanroom paper under the substrate for better heat uniformity. Ramp 8°C/min
|  | Step time @ 50°C [min] | 2 | 5
|  | Step time @ 85°C [min] | NA | 6.875
Develop | [5 empty] | Note: development upside down is highly recommended. Use two beakers for proper rinsing of substrate. Time in PGMEA should be minimized.
Develop time
|  | Develop time [min] | NA | 4.125
(Optional) Hard Bake
Hard bake time
|  | Step time @ 135°C [min] | 30 | 30 |  | Note: 135°C
Note: MLA150: laser emitting @ 405 nm! For exposure of mr-DWL layers on the MLA150, an edge bead removal is mandatory!
Troubleshooting guidelines

## V-shape Implants
V-Shape implants | Home
|  | Main consumables for the fabrication |  | Main equipment for the fabrication
Process Summary
|  | PDMS | https://www.dow.com/en-us/pdp.sylgard-184-silicone-elastomer-kit.01064291z.html#pdp-sample-options | Vacuum Pump | https://swissvacuum.com/vacuum-division
|  |  |  | Vacuum desiccator | https://www.sigmaaldrich.com/CH/en/product/aldrich/448931
PDMS molding |  | Silane | https://www.sigmaaldrich.com/CH/en/product/aldrich/448931
|  |  |  | Oven | https://shop.llg-labware.com/info466_Universal_drying_oven_LLG-uniiOVENi_lang_UK.htm?UID=55005feb15d8000000000000
|  | Ostemer | https://www.ostemers.com/products/ostemer-litho/
Silanization |  |  |  | Plasma Cleaner | https://harrickplasma.com/plasma-cleaners/basic-plasma-cleaner/
|  |  |  | UV Light | https://www.xyzprinting.com/en/product/uv-curing-chamber
Ostemer deposition
Detailed process flow
Release & cutting
|  | Silanization of the master mold wafer (once only) | • 6 drops of silane (2 drops in each small container, prefarably glass) for 6 hours. Make sure the containers are not close to the wafer.
Examination |  | PDMS molding | • Mix elastomer and curing agent at a ratio of 1:10 wt. Usually, 60 gram of total PDMS is enough for 4 inch wafer in a 35 mm petri dish. • Degas in the degassing chamber in the hood. The vacuum pump is so strong, the PDMS can bloat and spil. Adjust the vacuum accordingly. • Pour it onto wafer in the petri dish. • Degas again for 2 hours. Make sure no single bubble is left. • Cure it at 65°C in the oven min. for min 5 hours. • Cut the PDMS into small molds using razor blades. Refer to alignment markers at the edges
Cutting
|  | Silanization of the PDMS molds | • Place the PDMS molds into silanization chamber. • Put 6 drops (total) of silane onto the far sides. • Wait in full vacuum for 5 hours.
|  | Ostemer deposition | • Place the PDMS molds into silanization chamber. • Put 6 drops of silane onto the far sides. • Wait in full vacuum for 1h 37 mins.(this has to be tested with new chambers). • Warm Ostemer Part A at 48°C overnight. Make sure no crystal left. • Warm part B and and plastic mixing cup up to 48°C for 30 mins. • Mix part A and part B (1.86:1 wt\wt) thoroughly under yellow light. • Degas the mixture in the vacuum chamber for 5 mins. • Plasma treat (1 min) the glass slides at the same time of degassing. • Pour 200 μl drop of mixture onto each silanized PDMS mold. Or make sure you covered half of the PDMS mold with ostomer mixture. • Mechanically sandwich the mold between two glass slides using two clips. • UV light exposure for 10 mins. Rotate several times to make sure it is well polymerized.
|  | Releasing & cutting | • Relese the glass slides starting from connection side of the V implants. This increases the throughput by leaving minimal residue behind. • UV light exposure for another 10 mins to make sure about homogeneous polymerization. • Cut each implant using needles under the microscope.

## Molds for V-shape implants
Molds: V-Shape implants | Home
Summary
HMDS treatment
Coating | Detailed process flow
| Wafer prep !(critical)! | • HMDS is critical for resist bonding firmly onto the silicon wafer. • There are 3 free HMDS equipment in 1) Zone 4, 2) Zone 2 and 3) Zone 13. If they are not available use ACS 200 (recipe #8000). After HMDS, do not lose time. It is time sensitive.
Exposure
Development
| Coating | • EVG 150 automatic coater and developer in Zone 6 • Polymer: AZ10XT-60. • Recipe (8 μm coating): C4_D_10XT_8u_EC • wait min 12 mins before next step. The polymer needs rehydration.
Etching
Metrology
UFT removal | Exposure | • MJB4 exposure tool in Zone 13. • Exposure amount: Between 440-520 mJ\cm2 exposure works well. • Make sure no dust is on the mask and wafer
Metrology
A baking step can be applied before etching. For instance, the wafers can be wait at 85C in the oven in Zone 14 for overnight. The purpose is to dehydrate all possible water from the polymer so that no possible interaction between water moleules and etching chemicals. Even without this, the etching works quite well.
| Development | • EVG 150 automatic coater and developer in Zone 6 • Recipe: D4_N_10XT_8u_PUD • Clean the wafer with water and dry afterwards. • Suggestion: wait the wafer to dehydrate at 85 C in the oven in Zone 14 overnight.
Optional
| Etching & Metrology | • AMS200 in Zone 2. • Recipe: SOI_acc_++++ • Target etching height is 325-360 μm. • The recipe gives 3.5 to 4 μm\min etching rate. • First etch the wafer for 60 min. • Then, measure the etching height via DekTak mechanical profiler in Zone 4. Calculate the etching ratia\min for this case (it can change even from day to day). • Do one more etching that close to your target height. Then, measure again. • Do final etching for fine-tuning.. • Measure the height.
| UFT resist removal | • Warm up the UFT resist removal setup in Zone 2. The optimum temperature is 70 C. • Follow the all the steps. (18 steps) • Do inspection under the microscope and make sure no photoresist is left.
| Metrology | • Measure the heights and take notes behind the wafer.

## Numbered Windows
| Numbered Windows Fabrication |  |  | Home
[6 empty] | Main consumables for the fabrication |  | Main equipment for the fabrication
|  |  |  | Note: SU8 window thickness from 15 um to 20 is acceptable |  | Wafer (100/P/SS/01-100) Sigert Wafer, Germany | Order through EPFL CMi website | Vacuum Pump | https://swissvacuum.com/vacuum-division
[8 empty] | Vacuum desiccator | https://www.sigmaaldrich.com/CH/en/product/aldrich/448931
|  |  |  | Note: Recommended exposure tool: MJB4 |  | SU8 (3010 or 3025) Provided by CMi | for SU8 details
[8 empty] | Oven | https://shop.llg-labware.com/info466_Universal_drying_oven_LLG-uniiOVENi_lang_UK.htm?UID=55005feb15d8000000000000
[6 empty] | Dextran for Sacrificial layer | to order \ details
[8 empty] | Plasma Cleaner | https://harrickplasma.com/plasma-cleaners/basic-plasma-cleaner/
Wafer Prep | [7 empty] | UV Light | https://www.xyzprinting.com/en/product/uv-curing-chamber
Detailed process flow
Sacrificial layer |  | Wafer prep | • Dehydration is recommended (if available). In Zone 13, on ceramic hot plates. These are cheap hot plates and enough godd for dehyration. • Plasma cleaning and activation for better polymer bonding and uniform coating. (Z11, Tepla300, Program 5). It takes roughly 12 mins. Do not wait more than 10 mins between wafer prep and sacrificial layer because the prep is time dependant.
Spin Coat (SU8)
Soft Bake |  | Sacrificial layer | • Preparation (in the lab): The dextran is purchased as powder. Dextran and mili Q water are mixed as 20% (wt/vol). The mixture is waited at 90 C in a water bath for 2 hours to dissolve dextran thoroughly. • Pour it onto wafer covering min 60-70% of the surface. • Spin coat the wafer at 1000 RPM for 15 sec. Use WS-650-23 Laurell Technologies (currently replaced with sawatech coater) in Zone 13. This parameter gives you ~1 um sacrificial layer thickness. • Bake the wafer at 150°C for 2 mins on the hotplate next to the spin coater. |  | reference for sacrificial layer
Exposure |  |  |  |  | Spin coating parameters
Post Exposure Bake
|  | Spincoating of SU8 | • Enter the spincoating parameters into Sawatech in Z13 Segment 1: Time: 5s / Speed: 500rpm Segment 2: Time: 5s / Speed: 500rpm Segment 3: Time: 10s / Speed: 1500rpm Segment 4: Time: 40s / Speed: 1500rpm Segment 5: Time: 1s / Speed: 2500rpm Segment 6: Time: 1s / Speed: 1500rpm Segment 7: Time: 5s / Speed: 1500rpm Segment 8: Time: (15)s / Speed: 0rpm Segment 9 -> 24: Time: 0s / Speed: 0rpm • Pour SU8 covering 60% of the silicon wafer and run the machine. No relaxation time is needed.
Develop
Spincoating & Soft Bake (AZ40XT)
Exposure
Development
|  | Soft Bake | • Warm up the hotplate Sawatech to 95°C. Make sure the temperature is stable. • Bake the wafer for 600 secs for 3010 and 552 secs for 3025. • Wait the wafer at room temperature for 5 mins.
Metal Evaporation
|  | Exposure | • Warm up MJB4 in Zone 13. It takes 4-5 minutes. • Apply 13 seconds of UV light. This is higher than Cmi suggestion.
Removal
|  | Post Exposure Bake | • Warm up Sawatech hotplates to 95°C and 65°C. Make sure the temperature is stable. • Bake the wafer first at 65°C for 1 min. Then, 95°C for 4 mins. (this is for SU8 3010) • Relaxation for 10 mins at room temperature.
Release
|  | Development | • Prepare 2 cups of PGMEA bath. • Development 3 mins in each bath. • Rinse in IPA for 1 mins. • Dry using air gun.
|  | Spincoating & Soft Bake (AZ40XT) | • Spincoat: AZ40XT using ACS200 automatic processor. 20 μm thickness. Use the special recipe (number:??) for our process in this machine.If possible, wait the wafer in a dark box with wet tissues for hydration. This improves the light absorption in the next step.
|  | Exposure | • Warm up MJB4 in Zone 13. It takes 4-5 minutes. • Use the 2nd mask and mask aligner for matching first fabrication and second one. Alignment markers will help you.
|  | Development (AZ40XT) | • Development: remove exposed polymer from the waferT using ACS200. Check the edges and other parts under the microscope.
|  | Metal coating of numbers | • Physical vapor deposition: Ti (2 nm) and Au (10 nm) gold deposition using metal evaporation (EVA760). Ti makes proper bonding between SU8 and Au. Indeed, Ti and Au makes metallic alloy and then Au continues to grow up to 10 nm.
|  | Removal | • Remove positive photoresist using Remover 1165. The timing is important. It may take 1-2 days. Watch the solution.
|  | Release | • Release: Use DI water to release numbered windows. You can wash all the wafer or just put a drop onto a couple of windows needed. Gently handle the windows to avoid scratching.

## MR Runcard (hidden)
Back to Z01 MR Summary
Step N° | Description | Equipment | Program / Parameters | Remarks
X | PHOTOLITHOGRAPHY
X.1 | Wafer preparation | * See below | * | *
X.2 | Coating | Z1 / Sawatec LSM200 | Segment 1: Time: 5s / Speed: 500rpm Segment 2: Time: 5s / Speed: 500rpm Segment 3: Time: ((X-500)/100)s / Speed: Xrpm Segment 4: Time: 40s / Speed: Xrpm Segment 5: Time: 1s / Speed: X+1000rpm Segment 6: Time: 1s / Speed: Xrpm Segment 7: Time: 5s / Speed: Xrpm Segment 8: Time: (X/100)s / Speed: 0rpm Segment 9 -> 24: Time: 0s / Speed: 0rpm | Use the "CMISTART" recipe. X is the spinning speed to achieve the desired thickness. Use the "Z01 MR Summary" sheet to determine X.
X.3 | Relaxation time | Z1 / Sawatec HP401Z Z1 / Hotplate | 30 minutes @ 30°C | Make sure hotplate is at 30°C standby.
X.4 | Softbake | Z1 / Sawatec HP401Z Z1 / Hotplate | Segment 1: Time: 10 s / Temp.: 30°C Base Temp.: 30°C / Vaccum : ON / N2: ON Segment 2: Time: 600 s / Temp.: 50°C End segment: OFF / Vaccum : OFF / N2: ON Segment 3: Time: X s / Temp.: 50°C End segment: OFF / Vaccum : OFF / N2: ON Segment 4: Time: 600 s / Temp.: 90°C End segment: ON / Vaccum : OFF / N2: ON Segment 5: Time: Y s / Temp.: 90°C End segment: ON / Vaccum : OFF / N2: ON Segment 6: Time: 1800 s / Temp.: 30°C End segment: ON / Vaccum : OFF / N2: ON Segment 7 -> 24: NO INFLUENCE | X & Y are the softbake step times. Use the "Z01 MR Summary" sheet to determine X & Y. Ramp rule of thumb: 8°C/min
X.5 | Edge Bead Removal (EBR) | Z1 / Sawatec LSM200 | Segment 1: Time: 300s / Speed: 500rpm | Manual EBR with syringe filled with PGMEA
X.6 | Exposure | Z5 / MLA150 | Process: Mask-less Lithography Mode: High quality Dose : X Defocus : Y | X is the optimal exposure dose. Y is the optimal defocus. Use the "Z101 MR Summary" sheet to determine X. Check the parameters file on the MLA150 computer. Values are only a guide to find specific exposure dose and may be different for each process.
X.7 | Post-exposure bake | Z1 / Hotplate | PEB @ 50°C for X minutes PEB @ 85°C for Y minutes | X & Y are the PEB bake step times. Use the "Z01 MRSummary" sheet to determine X & Y. Ramp rule of thumb: 8°C/min
X.8 | Relaxation delay | Wafer storage box | From one hour to overnight, depending on thickness
X.9 | Development | Z1 / Plade Solvent Wet Bench | PGMEA development in two separate baths | Development time: Use the "Z01 MR Summary" sheet to determine the development time. The necessary time might be longer - if a reaction is visible within isopropanol (white residues appear), return to PGMEA for 1 min. Repeat until isopropanol does not react with dissolved SU8 residues.
X.10 | Rinse | Z1/ Plade Solvent Wet Bench | Isopropanol | 1 min
X.11 | Inspection | Z1 / uScope | Check for residues, resolution and alignment
* Necessary preparation depends on the substrate; see table below.
Si and any other material not listed below
Step N° | Description | Equipment | Program / Parameters | Remarks
X.1 | Wafer preparation | Z2 / Tepla Gigabatch | Prog. PR_strip_high_7min | 7 min
Metallic layers (in case the oxydation of the metallic layer may be a problem)
Step N° | Description | Equipment | Program / Parameters | Remarks
X.1 | Wafer preparation | Z1 / Hotplate | 20 min / 150°C | Only dehydration No HMDS !!!
Pyrex or Float glass
Step N° | Description | Equipment | Program / Parameters | Remarks
X.1 | Wafer preparation | Z2 / Ultrafab Piranha Z2 / Tepla Gigabatch | Prog. PR_strip_high_7min | Highly recommended 7 min

## Remounting stage
| Remounting Stage |  |  | Home |  | Main consumables for the fabrication
[6 empty] | Wafer (100/P/SS/01-100) Sigert Wafer, Germany | Order through EPFL CMi website
Substrate Prep |  | Detailed process flow |  |  |  | SU8 (3010 or 3025) Provided by CMi | for SU8 details
|  | Substrate prep | • Plasma cleaning and activation for better polymer bonding and uniform coating. (Z11, Tepla300, Program 5). It takes roughly 12 mins. Do not wait more than 10 mins between wafer prep and sacrificial layer because the prep is time dependant. |  |  | PAA for Sacrificial layer | to order \ details
Sacrificial layer
3D printing
|  | Sacrificial layer | • Preparation (in the lab): The PAA purchased as a 25 % (w/v) solution in water was neutralized with a saturated solution of NaOH until reaching a pH of 7.5 with a pH indicator band test. Then, dilute the solution to 10% PAA concentration using miliQ water. • Pour it onto wafer covering min 60-70% of the surface. • Spin coat the wafer at ~1500-2000 RPM for 15 sec. Use WS-650-23 Laurell Technologies in Zone 13. This parameter gives you ~1 um sacrificial layer thickness. • Bake the wafer at 150°C for 2 mins on the hotplate next to the spin coater. |  | A tip for sacrificial layer prep
for starting point: use 1 ml PAA (25% aqeous), 340 μl NaOH (10 M aqeous) and dilute with mili Q water until 10% aqeous PAA. This will result in ph of 7.3.
Development
Release
|  | 3D printing | • Use IP-S as material for printing (provided by CMi) • The design can be found in lab server. Ready to use
|  | Development | • Wait in PGMEA to remove uncured polymer for 20 mins. • Rinse with IPA thoroughly.
|  | Release | • Apply a couple of water drops. It relases in a couple of minutes by itself. You can see a borderline while PAA layer is dissolved under the microscope.

## Spin graph
Home
Reference for the film thickness
| x | y
| 0 | 0
| 5 | 500
5 | 10 | 500
15 | 25 | 2000
40 | 65 | 2000
1 | 66 | 3000
1 | 67
5 | 72
20 | 92

## Results mr-DWL-40 (hidden)
mr-DWL-05 |  |  |  | Home
Spin speed [rpm] | Film Thickness [um] |  | Film Thickness [um] | Typical exposure dose [mJ/cm2] | Softbake Time | PEB Time | Dev time
1000 | 80 |  | 10 | 200 | 3 | 2 | 3
2000 | 39 |  | 4 | 100 | 2 | 1 | 1
3000 | 27
4000 | 19
5000 | 16
| Target thickness [um] | Spin speed [rpm] | Exposure dose [mJ/cm2] | Softbake time [min] | PEB time [min] | Dev time [min]
Zone 13 | 50 | 1590.6685 | 450 | 6.875 | 6.875 | 4.125

## Results mr-DWL-5 (hidden)
mr-DWL-05 |  |  |  | Home
Spin speed [rpm] | Film Thickness [um] |  | Film Thickness [um] | Typical exposure dose [mJ/cm2] | Softbake Time | PEB Time | Dev time
1000 | 13 |  | 10 | 200 | 3 | 2 | 3
2000 | 7.5 |  | 4 | 100 | 2 | 1 | 1
3000 | 5
4000 | 4
5000 | 3.5
| Target thickness [um] | Spin speed [rpm] | Exposure dose [mJ/cm2] | Softbake time [min] | PEB time [min] | Dev time [min]
Zone 13 | 50 | 199.6164 | 1038.25 | 9.668 | 18.7695 | 9.383

## Results Su8 3005 (hidden)
MCC 3005 |  |  |  | Home
Spin speed [rpm] | Film Thickness [um] |  | Film Thickness [um] | Typical exposure dose [mJ/cm2] | Softbake Time | PEB Time | Dev time
1000 | 10 |  | 10 | 200 | 3 | 2 | 3
2000 | 7.5 |  | 4 | 100 | 2 | 1 | 1
3000 | 6
4000 | 5
5000 | 4.5
| Target thickness [um] | Spin speed [rpm] | Exposure dose [mJ/cm2] | Softbake time [min] | PEB time [min] | Dev time [min]
Zone 13 | 0 | #DIV/0! | 33 | 1.333 | 0.3333 | -0.333

## Results Su8 3010 (hidden)
MCC 3010 |  |  |  | Home
Spin speed [rpm] | Film Thickness [um] |  | Film Thickness [um] | Typical exposure dose [mJ/cm2] | Softbake Time | PEB Time | Dev time
1000 | 18.5 |  | 15 | 200 | 10 | 4 | 6
2000 | 13 |  | 8 | 125 | 5 | 2 | 4
3000 | 10.7
4000 | 9.5
5000 | 8.7
| Target thickness [um] | Spin speed [rpm] | Exposure dose [mJ/cm2] | Softbake time [min] | PEB time [min] | Dev time [min]
Zone 13 | 0 | #DIV/0! | 39.5 | -0.714 | -0.2857 | 1.7145

## Results Su8 3025 (hidden)
MCC 3025 |  |  |  | Home
Spin speed [rpm] | Film Thickness [um] |  | Film Thickness [um] | Typical exposure dose [mJ/cm2] | Softbake Time | PEB Time | Dev time
1000 | 68 |  | 50 | 250 | 15 | 5 | 8
2000 | 39 |  | 20 | 150 | 10 | 3 | 5
3000 | 27
4000 | 20
5000 | 17
| Target thickness [um] | Spin speed [rpm] | Exposure dose [mJ/cm2] | Softbake time [min] | PEB time [min] | Dev time [min]
Zone 13 | 0 | #DIV/0! | 83.3 | 6.667 | 1.667 | 3

## Results Su8 3035 (hidden)
MCC 3035 |  |  |  | Home
Spin speed [rpm] | Film Thickness [um] |  | Film Thickness [um] | Typical exposure dose [mJ/cm2] | Softbake Time | PEB Time | Dev time
1000 | 87.5 |  | 80 | 250 | 30 | 5 | 12
2000 | 50 |  | 30 | 150 | 10 | 3 | 6
3000 | 33.5
4000 | 25
5000 | 21
| Target thickness [um] | Spin speed [rpm] | Exposure dose [mJ/cm2] | Softbake time [min] | PEB time [min] | Dev time [min]
Zone 13 | 0 | #DIV/0! | 90 | -2 | 1.8 | 2.4

## Results Su8 3050 (hidden)
MCC 3050 |  |  |  | Home
Spin speed [rpm] | Film Thickness [um] |  | Film Thickness [um] | Typical exposure dose [mJ/cm2] | Softbake Time | PEB Time | Dev time
1000 | 115 |  | 100 | 250 | 45 | 5 | 15
2000 | 67.5 |  | 40 | 150 | 15 | 3 | 7
3000 | 47.5
4000 | 34
5000 | 29
| Target thickness [um] | Spin speed [rpm] | Exposure dose [mJ/cm2] | Softbake time [min] | PEB time [min] | Dev time [min]
Zone 13 | 0 | #DIV/0! | 83.3 | -5 | 1.667 | 1.667

## Results GM 1040 (hidden)
GM 1040 |  |  |  | Home
Spin speed [rpm] | Film Thickness [um] |  | Film Thickness [um] | Typical exposure dose [mJ/cm2]
2000 | 1.45 |  | 0.94 | 29
4000 | 0.94 |  | 1.45 | 39
| Target thickness [um] | Spin speed [rpm] | Exposure dose [mJ/cm2]
Zone 01 | https://www.sigmaaldrich.com/CH/en/product/aldrich/448931 | #VALUE! | #VALUE!
Zone 13 | https://www.sigmaaldrich.com/CH/en/product/aldrich/448931 | #VALUE! | #VALUE!

## Results GM 1050 (hidden)
GM 1050 |  |  |  | Home
Spin speed [rpm] | Film Thickness [um] |  | Film Thickness [um] | Typical exposure dose [mJ/cm2]
2000 | 3.72 |  | 2.3 | 45
4000 | 2.3 |  | 3.72 | 55
| Target thickness [um] | Spin speed [rpm] | Exposure dose [mJ/cm2]
Zone 01 | https://www.sigmaaldrich.com/CH/en/product/aldrich/448931 | #VALUE! | #VALUE!
Zone 13 | https://www.sigmaaldrich.com/CH/en/product/aldrich/448931 | #VALUE! | #VALUE!

## Results GM 1060 (hidden)
GM 1060 |  |  |  | Home
Spin speed [rpm] | Film Thickness [um] |  | Film Thickness [um] | Typical exposure dose [mJ/cm2]
2000 | 10.8 |  | 5.75 | 73.14
4000 | 5.75 |  | 10.8 | 89.7
| Target thickness [um] | Spin speed [rpm] | Exposure dose [mJ/cm2]
Zone 01 | https://www.sigmaaldrich.com/CH/en/product/aldrich/448931 | #VALUE! | #VALUE!
Zone 13 | https://www.sigmaaldrich.com/CH/en/product/aldrich/448931 | #VALUE! | #VALUE!

## Results GM 1070 (hidden)
GM 1070 |  |  |  | Home
Spin speed [rpm] | Film Thickness [um] |  | Film Thickness [um] | Typical exposure dose [mJ/cm2]
2000 | 38.5 |  | 18.4 | 121.9
4000 | 18.4 |  | 38.5 | 162.84
| Target thickness [um] | Spin speed [rpm] | Exposure dose [mJ/cm2]
Zone 01 | https://www.sigmaaldrich.com/CH/en/product/aldrich/448931 | #VALUE! | #VALUE!
Zone 13 | https://www.sigmaaldrich.com/CH/en/product/aldrich/448931 | #VALUE! | #VALUE!

## Results GM 1075 (hidden)
GM 1075 |  |  |  | Home
Spin speed [rpm] | Film Thickness [um] |  | Film Thickness [um] | Typical exposure dose [mJ/cm2]
2000 | 94.2 |  | 43 | 147
4000 | 43 |  | 94.2 | 174
| Target thickness [um] | Spin speed [rpm] | Exposure dose [mJ/cm2]
Zone 01 | https://www.sigmaaldrich.com/CH/en/product/aldrich/448931 | #VALUE! | #VALUE!
Zone 13 | https://www.sigmaaldrich.com/CH/en/product/aldrich/448931 | #VALUE! | #VALUE!