## Runcard
Step N° | Description | Equipment | Program / Parameters | Remarks
X | PHOTOLITHOGRAPHY - Mask X
X.1 | Wafer preparation | - | Depends on the substrate; see tables below. | Available options: - dehydration by hotplate or oven - oxygen plasma - pirahna cleaning Make sure to allow dose to
X.2 | Primer deposition | Z6 / Plade divers | Spinning with BLE spinner: 10 s @ 500 RPM, nitrogen off 30 s @ 3000 RPM, nitrogen on (manual) | Dilute VM 651 in IPA: 1:200. Labelled beaker is available. Use pipette to measure 1 ml of VM651, and add to 200 ml of IPA. Leave primer puddle on wafer for 30 s before start of spinning. When finished, clean beaker properly with acetone. Dispose of the waste immediately.
X.3 | Coating | Z1/Sawatec Coater | Segment 1: Time: 5s / Speed: 500rpm Segment 2: Time: 5s / Speed: 500rpm Segment 3: Time: ((X-500)/100)s / Speed: Xrpm Segment 4: Time: 40s / Speed: Xrpm Segment 5: Time: 1s / Speed: X+1000rpm Segment 6: Time: 1s / Speed: Xrpm Segment 7: Time: 5s / Speed: Xrpm Segment 8: Time: (X/100)s / Speed: 0rpm Segment 9 -> 24 : Time: 0s / Speed: 0rpm | Reserve Sawatec machine with PI configuration in advance. Check the "PI processing" label to be present on the tool. X is the speed needed to get the desired thickness. Use "Results" sheets to determine X. Use dedicated tools for transferring wafers to hotplate.
X.4A | PR soft bake (slow version) | Z1/Sawatec Hotplate | Segment 1: Time: 10s / Temp.: 30°C Base Temp.: 30°C / Vaccum: ON / N2: ON / Pin: Down Segment 2: Time: 1800s / Temp.: 75°C End segment: OFF / Vaccum: OFF / N2: ON / Pin: Down Segment 3: Time: 2700s / Temp.: 120°C End segment: OFF / Vaccum: OFF / N2: ON / Pin: Down Segment 4: Time: 300s / Temp.: 120°C End segment: OFF / Vaccum: OFF / N2: ON / Pin: Down Segment 5: Time: 1200s / Temp.: 30°C End segment: ON / Vaccum: OFF / N2: ON / Pin: Down Segment 6 -> 24: NO INFLUENCE | Use the "CMI 2610" or "CMI 2611" recipe. This recipe can be used on the Dataplate as well. DO NOT OPEN the cover as long as the Dataplate hotplate is above 30°C, due to the high concentration of the evaporated NMP solvent. Clean the inside area of the Sawatec cover before and after the process with a cleanroom wipe.
X.4B | PR soft bake (fast version) | Z1/Suss hotplates | Step 1: Time: 180s / Temp.: 65°C Step 2: Time: 180s / Temp: 105°C | Set two hotplates to their repective process temperature and allow to stabilize. After use, put the set point of the hotplates back to room temperature.
X.5 | PR curing | Z2/Heraeus oven | N2 ON above 200°C Pr=1 n=4 r=0 P0=50 T1=0.30 P1=200 T2=1.00 P2=200 T3=0.25 P3=300 T4=1.00 P4=300 | This program takes 5 hours for the wafers to cool below 50°C. Make sure nitrogen flow is on.
Si and any other material, except mentioned below
Step N° | Description | Equipment | Program / Parameters | Remarks
X.1 | Wafer preparation | Z2/Tepla | Prog. 5 | 7 min
Metallic layer (in case oxydation of the metallic layer may be a problem)
Step N° | Description | Equipment | Program / Parameters | Remarks
X.1 | Wafer preparation | Z6/RC8 entrance or Z2/Heraeus oven | 10 min | Only dehydration
Pyrex or Float glass
Step N° | Description | Equipment | Program / Parameters | Remarks
X.1 | Wafer preparation | Z2/Ultrafab piranha or Z2/Tepla | Prog. 5 | Highly recommended 7 min

## Results PI2610
PI2610 on Si |  | PI2610 on PI
Spin speed [rpm] | Film Thickness [um] | Spin speed [rpm] | Film Thickness [um]
1000 | 5.11 | 1273 | 2.968
2000 | 2.53 | 2485 | 1.115
3000 | 1.66 | 1000 | 3.916
4000 | 1.19 | 2000 | 1.55
5000 | 0.98 | 3000 | 0.819
|  | 5000 | 0.151
Desired thickness [um] | Spin speed [rpm] | Desired thickness [um] | Spin speed [rpm]
1 | 4849.8902 | 1 | 2665.9804

## Results PI2611
PI2611 on Si |  | PI2611 on PI
Spin speed [rpm] | Film Thickness [um] | Spin speed [rpm] | Film Thickness [um]
2000 | 6.87 | 2287 | 4.737
3000 | 4.56 | 4429 | 2.385
4000 | 3.34 | 1500 | 8.02
5000 | 2.63 | 2500 | 4.392
|  | 3500 | 2.837
|  | 5000 | 1.638
Desired thickness [um] | Spin speed [rpm] | Desired thickness [um] | Spin speed [rpm]
16 | 897.8867 | 14 | 902.6037

## Dataplate setup
Recipe #00
Step | Mode | Command | Value | Explanation
01. | L | 8 | 135 | Ramp with 135°C/hr | Mode: - L: set target - U: wait until Command: - 1: Set plate temperature - 5: Wait (min:sec) - 8: Ramp (°C/hr)
02. | U | 1 | 75 | Set temp to 75°C
03. | L | 8 | 68 | Ramp with 68°C/hr
04. | U | 1 | 120 | Set temp to 120°C
05. | U | 5 | 500 | Wait 5 min
06. | L | 1 | 35 | Set temp to 30°C
Programming instructions
Command | Operation | Value | Save | Explanation
Power on
Edit | Clear |  | Edit | Repeat Clear until 00 on display, to clear previous recipe | Ent: add a new step in set target mode Ent - Ent: add a new step in wait until mode
Edit | Plate temp | 30 | Set | Set hotplate temp to 30°C, before loading wafers
Edit |  |  |  | Enter Edit menu to set up new recipe
Ent | Ramp °/hr | 135 | Set | Set ramp to 135°C per hour
Ent - Ent | Plate temp | 75 | Set | Wait until plate temp is at 75°C
Ent | Ramp °/hr | 68 | Set | Set ramp to 68°C per hour
Ent - Ent | Plate temp | 120 | Set | Wait until plate temp is at 120°C
Ent - Ent | Timer (m:s) | 500 | Set | Wait 5 minutes
Ent | Plate temp | 30 | Set | Set plate temp to 30°C
Edit |  |  |  | Exit Edit mode
Run |  |  |  | Load wafers and run the recipe